## Instructions
Sandia Primary Standards Lab - Excel Calibration Interval Calculator (RP-1 Methods)
v0.1
uncertainty@sandia.gov
This spreadsheet provides calculation of interval optimization methods in NCSLI Recommended Practice 1. Green cells are user inputs, blue cells are calculated output values. Gray cells, often hidden, are intermediate calculated values or plot data. No Excel Macros are used, however the calculations make use of Excel's LAMBDA function which is only available in Office 365 or 2024.
Spreadsheet Tabs
A1 Simple | Simple Response Method. Not recommended in practice, but included for reference and to "discourage its reinvention".
A2 Incremental | Incremental Response Method. Not recommended in practice, but included for reference and to "discourage its reinvention".
A3 Test Interval | Test whether the assigned interval has been achieving the reliability target based on historical pass/fail values at a single interval.
S2 Binomial Method | Fit historical reliability to a reliability model to find the optimal interval. Requires historical pass/fail data at various intervals.
V1 Reliability Target | Variables method with a reliability target. Uses historical data of a single asset to predict drift and set interval to achieve a reliability.
V1 Uncertainty Target | Variables method with a uncertainty target. Uses historical data of a single asset to predict drift and set interval to achieve an allowable uncertianty growth.
Function Reference | Lists the LAMBDA functions used in Excel's Name Manager for readability. Hidden by default.
About
Sandia National Laboratories is a multimission laboratory managed and operated by National Technology & Engineering Solutions of Sandia, LLC, a wholly owned subsidiary of Honeywell International Inc., for the U.S. Department of Energy’s National Nuclear Security Administration under contract DE-NA0003525.
SAND2026-16029O

## A1 Simple
Method A1 and A2 are not recommended due to inability to hit the reliability target. It is included in this spreadsheet (and RP-1) to "discourage their reinvention".
Reliability Target | 90% | Average Reliability | 84.21%
Increase coefficient a | 0.1
Decrease coefficient b | 0.5759
Enter "1" for in-tolerance, "0" for out of tolerance.
Pass/Fail | Interval
Initial | 365
1 | 401.5
1 | 441.65
0 | 187.3027
1 | 206.033
1 | 226.6363
1 | 249.2999
1 | 274.2299
0 | 116.3002
1 | 127.9303
1 | 140.7233
1 | 154.7956
1 | 170.2752
1 | 187.3027
1 | 206.033
1 | 226.6363
1 | 249.2999
1 | 274.2299
1 | 301.6529
0 | 127.9303

## A2 Incremental
Method A1 and A2 are not recommended due to inability to hit the reliability target. It is included in this spreadsheet (and RP-1) to "discourage their reinvention".
Reliability Target | 90% |  | Average Reliability | 78.95%
Initial Interval | 365
Enter "1" for in-tolerance, "0" for out of tolerance.
Pass/Fail | Delta | New Interval
1 | 1 | 402
0 | 0.5 | 221
1 | 0.25 | 227
1 | 0.25 | 233
1 | 0.25 | 239
1 | 0.25 | 245
0 | 0.125 | 217
0 | 0.125 | 193
1 | 0.0625 | 194
1 | 0.0625 | 195
1 | 0.0625 | 196
1 | 0.0625 | 197
1 | 0.0625 | 198
1 | 0.0625 | 199
1 | 0.0625 | 200
1 | 0.0625 | 201
1 | 0.0625 | 202
0 | 0.03125 | 196
1 | 0.015625 | 196

## A3 Test Interval
May include calibration cycles for similar items (other items of same model number, for example)
Enter "1" for in-tolerance, "0" for out of tolerance. Can use "NA", or other designators to exclude items that were damaged, lost, etc.
Asset | Workorder | Start Date | End Date | Interval Length | IT(1)/OOT(0) | Parameters
|  | 2010-01-01 | 2011-01-01 | 365 | 1 | Interval (days) to Test | 365
|  | 2011-01-01 | 2012-01-01 | 365 | 1 | Tolerance (days) | 30
|  | 2012-01-01 | 2013-01-01 | 366 | 1 | Reliability Target | 95%
|  | 2013-01-01 | 2014-01-04 | 368 | 1 | Change confidence | 50%
|  | 2014-01-01 | 2016-01-01 | 730 | 1 | Maximum change factor | 2
|  | 2016-01-01 | 2018-01-01 | 731 | 1
|  |  |  | 0 | 1
|  |  |  | 0 | 1 | Total calibrations | 4
|  |  |  | 0 | 1 | Passing calibrations | 4
|  |  |  | 0 | 1 | Observed Reliability | 100%
|  |  |  | 0 | 1
|  |  |  | 0 | 1 | True reliability (lower) | 84.09%
|  |  |  | 0 | 1 | True reliability (upper) | 100%
|  |  |  | 0 | 1 | Rejection Confidence | 18.55%
|  |  |  | 0 | 1
|  |  |  | 0 | 1 | Suggested interval | 365
|  |  |  | 0 | 1

## S2 Binomial
Reliability Target | 80% |  | Results |  | Model | Exponential | Mixed Exponential |  | Weibull |  | Random Walk |  | Restricted Walk |  |  | Gamma | Mortality Drift |  | Warranty |  | Drift |  |  | Lognormal
|  |  | Best Fit Model | Mixed Exponential | Model Equation | R = exp(-t*a) | R = (1+at)^(-b) |  | R=exp(-(t*a)^b) |  | R=2Φ( 1/SQRT(a+bt) ) - 1 |  | R=2Φ( 1/SQRT(a+b(1-exp(-c*t)) ) - 1 |  |  | R=exp(-a*t)*(a+a*t+(at)^2/2 + (at)^3/6) | R=exp(-(a*t+b*t^2)) |  | 1/(1+exp(a*(t-b))) |  | R=Φ(a+c*t)+Φ(b-c*t)-1 |  |  | 1-Φ(ln(a*t)/b)
Enter historical reliability data for multiple interval lengths. |  |  | Suggested Interval | 22.9324 | Initial Guess | 0.0001 | 0.03 | 0.4 | 0.003 | 0.5 | 0.5 | 0.01 | 0.5 | 100 | 0.0001 | 0.02 | 0.004 | 1e-06 | 0.005 | 100 | 5 | 0.5 | 0.001 | 0.00713 | 2
Interval Length | Reliability | Calibrations
21 | 0.999 | 4 |  |  | Calculated Interval | 53.1775 | 22.9324 |  | 20.9731 |  | 10.9691 |  | 12.8985 |  |  | 105.5579 | 56.7728 |  | 0 |  | 0 |  |  | 31.4612
42 | 0.8333 | 6
63 | 0.6429 | 14 |  |  | Fit Parameters (MLE) | 0.0041962 | 0.03525 | 0.37666 | 0.0033559 | 0.5652 | 0.49587 | 0.010302 | 0.40117 | 155.9825 | 0.0001033 | 0.021759 | 0.0038842 | 8.1434e-07 | 0.0045958 | 210.1295 | 11.215 | 0.051466 | 0.001 | 0.0059048 | 2
84 | 0.6154 | 13 |  |  | F | 0.62591 | 0.3874 |  | 0.42185 |  | 0.47902 |  | 0.83486 |  |  | 2.9796 | 0.80365 |  | 0.86379 |  | 2.0321 |  |  | 0.37917
140 | 0.5455 | 22 |  |  | F Critical | 2.0842 | 2.1725 |  | 2.1725 |  | 2.1725 |  | 2.2874 |  |  | 2.0842 | 2.1725 |  | 2.1725 |  | 2.2874 |  |  | 2.1725
189 | 0.4082 | 49 |  |  | Rejection Conf. | 26.6296 | 11.4052 |  | 13.6688 |  | 17.7091 |  | 47.266 |  |  | 99.3645 | 43.1094 |  | 47.6356 |  | 92.1361 |  |  | 10.8866
269.5 | 0.5 | 18 |  |  | Accept? | TRUE | TRUE |  | TRUE |  | TRUE |  | TRUE |  |  | FALSE | TRUE |  | TRUE |  | TRUE |  |  | TRUE
346.5 | 0.3333 | 6 |  |  | Figure of Merit | 116.3289 | 410.7762 |  | 262.622 |  | 244.5154 |  | 83.094 |  |  | 28.1091 | 71.8573 |  | 19.9695 |  | 11.8665 |  |  | 50.4733
|  |  | Max Observed Interval | 346.5 | tr | 383.546 | 2006.4786 |  | 691.6059 |  | 1464.2501 |  | 991.0225 |  |  | 253.4656 | 383.5147 |  | 511.7755 |  | 893.0877 |  |  | 911.6339
[5 empty] | NG | 7 | 7 |  | 7 |  | 7 |  | 7 |  |  | 7 | 7 |  | 2 |  | 2 |  |  | 1
Plot Line Target Reliability
|  |  | 0 | 80% | t | R(t) | R(t) |  | R(t) |  | R(t) |  | R(t) |  |  | R(t) | R(t) |  | R(t) |  | R(t) |  |  | R(t)
|  |  | 346.5 | 80% | 0 | 1 | 1 |  | 1 |  | 0.84442 |  | 0.88563 |  |  | 1 | 1 |  | 0.72426 |  | 0.52052 |  |  | 1
[5 empty] | 18.2368 | 0.92633 | 0.82945 |  | 0.81368 |  | 0.77347 |  | 0.76976 |  |  | 0.99925 | 0.93136 |  | 0.70721 |  | 0.51325 |  |  | 0.86742
|  |  | Interval | NG | 36.4737 | 0.85809 | 0.73244 |  | 0.73707 |  | 0.71588 |  | 0.68566 |  |  | 0.99116 | 0.86697 |  | 0.68956 |  | 0.50598 |  |  | 0.77867
|  |  | 53.1775 | 7 | 54.7105 | 0.79487 | 0.66716 |  | 0.68137 |  | 0.66871 |  | 0.62319 |  |  | 0.96706 | 0.80658 |  | 0.67134 |  | 0.49871 |  |  | 0.71395
|  |  | 22.9324 | 7 | 72.9474 | 0.73631 | 0.61911 |  | 0.63674 |  | 0.62941 |  | 0.57485 |  |  | 0.92294 | 0.75 |  | 0.65259 |  | 0.49143 |  |  | 0.66317
|  |  | 20.9731 | 7 | 91.1842 | 0.68207 | 0.58169 |  | 0.59925 |  | 0.59612 |  | 0.53615 |  |  | 0.85999 | 0.69701 |  | 0.63336 |  | 0.48416 |  |  | 0.62155
|  |  | 10.9691 | 7 | 109.4211 | 0.63182 | 0.5514 |  | 0.56685 |  | 0.5675 |  | 0.50431 |  |  | 0.78272 | 0.64742 |  | 0.61369 |  | 0.47689 |  |  | 0.58644
|  |  | 12.8985 | 7 | 127.6579 | 0.58527 | 0.52618 |  | 0.53831 |  | 0.54258 |  | 0.47755 |  |  | 0.69691 | 0.60102 |  | 0.59364 |  | 0.46963 |  |  | 0.55619
|  |  | 105.5579 | 7 | 145.8947 | 0.54215 | 0.50472 |  | 0.51279 |  | 0.52063 |  | 0.45464 |  |  | 0.60821 | 0.55765 |  | 0.57327 |  | 0.46238 |  |  | 0.52972
|  |  | 56.7728 | 7 | 164.1316 | 0.50221 | 0.48615 |  | 0.48975 |  | 0.50112 |  | 0.43477 |  |  | 0.52134 | 0.51713 |  | 0.55265 |  | 0.45515 |  |  | 0.50625
|  |  | 0 | 2 | 182.3684 | 0.46522 | 0.46986 |  | 0.46876 |  | 0.48363 |  | 0.41731 |  |  | 0.43973 | 0.47929 |  | 0.53185 |  | 0.44793 |  |  | 0.48524
|  |  | 0 | 2 | 200.6053 | 0.43094 | 0.45539 |  | 0.44951 |  | 0.46783 |  | 0.40182 |  |  | 0.3656 | 0.44398 |  | 0.51094 |  | 0.44072 |  |  | 0.46626
|  |  | 31.4612 | 1 | 218.8421 | 0.39919 | 0.44244 |  | 0.43175 |  | 0.45348 |  | 0.38795 |  |  | 0.30008 | 0.41105 |  | 0.48999 |  | 0.43354 |  |  | 0.449
[5 empty] | 237.0789 | 0.36979 | 0.43073 |  | 0.41529 |  | 0.44037 |  | 0.37545 |  |  | 0.24348 | 0.38036 |  | 0.46908 |  | 0.42637 |  |  | 0.43321
[5 empty] | 255.3158 | 0.34254 | 0.42009 |  | 0.39997 |  | 0.42833 |  | 0.3641 |  |  | 0.19551 | 0.35177 |  | 0.44827 |  | 0.41924 |  |  | 0.41869
[5 empty] | 273.5526 | 0.31731 | 0.41034 |  | 0.38566 |  | 0.41722 |  | 0.35373 |  |  | 0.15553 | 0.32515 |  | 0.42764 |  | 0.41212 |  |  | 0.40526
[5 empty] | 291.7895 | 0.29393 | 0.40138 |  | 0.37225 |  | 0.40693 |  | 0.34422 |  |  | 0.12268 | 0.30038 |  | 0.40726 |  | 0.40504 |  |  | 0.3928
[5 empty] | 310.0263 | 0.27228 | 0.39309 |  | 0.35964 |  | 0.39736 |  | 0.33546 |  |  | 0.096022 | 0.27735 |  | 0.3872 |  | 0.39799 |  |  | 0.3812
[5 empty] | 328.2632 | 0.25222 | 0.38539 |  | 0.34777 |  | 0.38843 |  | 0.32734 |  |  | 0.074631 | 0.25594 |  | 0.36751 |  | 0.39097 |  |  | 0.37036
[5 empty] | 346.5 | 0.23364 | 0.37822 |  | 0.33656 |  | 0.38008 |  | 0.3198 |  |  | 0.057634 | 0.23606 |  | 0.34825 |  | 0.38398 |  |  | 0.36019

## V1 Reliability Target
Enter historical as-found and as-left values for one item
Calibration Date | As-Found Value | As-Left Value | Resubmission Time t | ∆(t) | Parameters |  | Results
2003-03-29 | 10.173 | 10.073 | 104 | 0.251 | Fit Order | 2 | Suggested Interval | 171.2683
2003-07-11 | 10.324 | 10.048 | 86 | 0.11 | Initial Value | 10.03 | Fit Uncertainty | 0.048882
2003-10-05 | 10.158 | 9.993 | 135 | 0.299 | Initial Uncertainty | 0.28 | Projected Result | 10.5479
2004-02-17 | 10.292 | 10.126 | 70 | 0.1 | Initial Uncertainty Degf | 3.45 | Projected Uncertainty | 0.28423
2004-04-27 | 10.226 | 10.024 | 173 | 0.615 | Reliability Target | 90% | Projected Deg. F. | 3.6606
2004-10-17 | 10.639 | 10.208 | 167 | 0.403 | Lower Tolerance | 9
2005-04-02 | 10.611 | 10.451 |  |  | Upper Tolerance | 11 | Fit Parameters | 0.00015741 | 1.6738e-05
[7 empty] | Syx / degf_fit | 0.070829 | 4
[5 empty] | Plot Setup |  | Fit Covariance | 1.0101e-06 | -6.6068e-09
[5 empty] | Target Search Max | 1730 |  | -6.6068e-09 | 4.5495e-11
[5 empty] | Reliability Plot Max | 432.5
[5 empty] | Data Plot Max | 190.3
[5 empty] | t | R(t) | Target Plot Line |  |  | t | y(t) | u(t) | y-u | y+u
[5 empty] | 0 | 0.97001 | 0 | 0.9 |  | 0 | 0 | 0 | 0 | 0
[5 empty] | 22.7632 | 0.97014 | 432.5 | 0.9 |  | 10.0158 | 0.0032556 | 0.0094077 | -0.0061521 | 0.012663
[5 empty] | 45.5263 | 0.97008 |  |  |  | 20.0316 | 0.0098694 | 0.017505 | -0.0076352 | 0.027374
[5 empty] | 68.2895 | 0.96898 | Interval Plot Line |  |  | 30.0474 | 0.019841 | 0.024301 | -0.0044597 | 0.044142
[5 empty] | 91.0526 | 0.96586 | 171.2683 | 0 |  | 40.0632 | 0.033172 | 0.029811 | 0.0033604 | 0.062983
[5 empty] | 113.8158 | 0.95927 | 171.2683 | 1.1 |  | 50.0789 | 0.04986 | 0.034054 | 0.015805 | 0.083914
[5 empty] | 136.5789 | 0.94651 |  |  |  | 60.0947 | 0.069906 | 0.037058 | 0.032848 | 0.10696
[5 empty] | 159.3421 | 0.92172 |  |  |  | 70.1105 | 0.093311 | 0.038863 | 0.054448 | 0.13217
[5 empty] | 182.1053 | 0.87068 |  |  |  | 80.1263 | 0.12007 | 0.039529 | 0.080545 | 0.1596
[5 empty] | 204.8684 | 0.76423 |  |  |  | 90.1421 | 0.15019 | 0.039149 | 0.11105 | 0.18934
[5 empty] | 227.6316 | 0.58001 |  |  |  | 100.1579 | 0.18367 | 0.037873 | 0.1458 | 0.22155
[5 empty] | 250.3947 | 0.36748 |  |  |  | 110.1737 | 0.22051 | 0.035951 | 0.18456 | 0.25646
[5 empty] | 273.1579 | 0.2068 |  |  |  | 120.1895 | 0.26071 | 0.033803 | 0.2269 | 0.29451
[5 empty] | 295.9211 | 0.11557 |  |  |  | 130.2053 | 0.30426 | 0.03212 | 0.27214 | 0.33638
[5 empty] | 318.6842 | 0.070518 |  |  |  | 140.2211 | 0.35117 | 0.031893 | 0.31928 | 0.38306
[5 empty] | 341.4474 | 0.04878 |  |  |  | 150.2368 | 0.40144 | 0.034173 | 0.36727 | 0.43561
[5 empty] | 364.2105 | 0.037704 |  |  |  | 160.2526 | 0.45507 | 0.039538 | 0.41553 | 0.49461
[5 empty] | 386.9737 | 0.031499 |  |  |  | 170.2684 | 0.51206 | 0.047898 | 0.46416 | 0.55995
[5 empty] | 409.7368 | 0.02763 |  |  |  | 180.2842 | 0.5724 | 0.058841 | 0.51356 | 0.63124
[5 empty] | 432.5 | 0.024958 |  |  |  | 190.3 | 0.6361 | 0.071973 | 0.56413 | 0.70808

## V1 Uncertainty Target
Enter historical as-found and as-left values for one item
Calibration Date | As-Found Value | As-Left Value | Resubmission Time t | ∆(t) | Parameters |  | Results
2003-03-29 | 10.173 | 10.073 | 104 | 0.251 | Fit Order | 2 | Suggested Interval | 328.7755
2003-07-11 | 10.324 | 10.048 | 86 | 0.11 | Initial Value | 10.03 | Fit Uncertainty | 0.41371
2003-10-05 | 10.158 | 9.993 | 135 | 0.299 | Initial Uncertainty | 0.28 | Projected Result | 11.891
2004-02-17 | 10.292 | 10.126 | 70 | 0.1 | Initial uncertainty degf | 3 | Projected Uncertainty | 0.49955
2004-04-27 | 10.226 | 10.024 | 173 | 0.615 | Uncertainty Target | 0.5 | Projected Deg. F. | 6.6448
2004-10-17 | 10.639 | 10.208 | 167 | 0.403
2005-04-02 | 10.611 | 10.451 |  |  | Plot Setup
[5 empty] | Search Max | 1730 | Fit Parameters | 0.00015741 | 1.6738e-05
[5 empty] | Uncertainty Plot Max | 432.5 | Syx / degf_fit | 0.070829 | 4
[5 empty] | Data Plot Max | 190.3 | Fit Covariance | 1.0101e-06 | -6.6068e-09
[8 empty] | -6.6068e-09 | 4.5495e-11
[5 empty] | t | u(t) | Target Plot Line |  |  | t | y(t) | u(t) | y-u | y+u
[5 empty] | 0 | 0.28 | 0 | 0.5 |  | 0 | 0 | 0 | 0 | 0
[5 empty] | 22.7632 | 0.28068 | 432.5 | 0.5 |  | 10.0158 | 0.0032556 | 0.0094077 | -0.0061521 | 0.012663
[5 empty] | 45.5263 | 0.28185 |  |  |  | 20.0316 | 0.0098694 | 0.017505 | -0.0076352 | 0.027374
[5 empty] | 68.2895 | 0.28265 | Interval Plot Line |  |  | 30.0474 | 0.019841 | 0.024301 | -0.0044597 | 0.044142
[5 empty] | 91.0526 | 0.28271 | 328.7755 | 0 |  | 40.0632 | 0.033172 | 0.029811 | 0.0033604 | 0.062983
[5 empty] | 113.8158 | 0.2822 | 328.7755 | 0.88893 |  | 50.0789 | 0.04986 | 0.034054 | 0.015805 | 0.083914
[5 empty] | 136.5789 | 0.28179 |  |  |  | 60.0947 | 0.069906 | 0.037058 | 0.032848 | 0.10696
[5 empty] | 159.3421 | 0.28269 |  |  |  | 70.1105 | 0.093311 | 0.038863 | 0.054448 | 0.13217
[5 empty] | 182.1053 | 0.28658 |  |  |  | 80.1263 | 0.12007 | 0.039529 | 0.080545 | 0.1596
[5 empty] | 204.8684 | 0.29549 |  |  |  | 90.1421 | 0.15019 | 0.039149 | 0.11105 | 0.18934
[5 empty] | 227.6316 | 0.3115 |  |  |  | 100.1579 | 0.18367 | 0.037873 | 0.1458 | 0.22155
[5 empty] | 250.3947 | 0.33634 |  |  |  | 110.1737 | 0.22051 | 0.035951 | 0.18456 | 0.25646
[5 empty] | 273.1579 | 0.37113 |  |  |  | 120.1895 | 0.26071 | 0.033803 | 0.2269 | 0.29451
[5 empty] | 295.9211 | 0.4163 |  |  |  | 130.2053 | 0.30426 | 0.03212 | 0.27214 | 0.33638
[5 empty] | 318.6842 | 0.47177 |  |  |  | 140.2211 | 0.35117 | 0.031893 | 0.31928 | 0.38306
[5 empty] | 341.4474 | 0.53715 |  |  |  | 150.2368 | 0.40144 | 0.034173 | 0.36727 | 0.43561
[5 empty] | 364.2105 | 0.61198 |  |  |  | 160.2526 | 0.45507 | 0.039538 | 0.41553 | 0.49461
[5 empty] | 386.9737 | 0.6958 |  |  |  | 170.2684 | 0.51206 | 0.047898 | 0.46416 | 0.55995
[5 empty] | 409.7368 | 0.78823 |  |  |  | 180.2842 | 0.5724 | 0.058841 | 0.51356 | 0.63124
[5 empty] | 432.5 | 0.88893 |  |  |  | 190.3 | 0.6361 | 0.071973 | 0.56413 | 0.70808

## Function Reference (hidden)
These are the same formulas entered in the Name Manager, listed here for readability. To modify, copy them to a real code editor then paste in to the name manager. |  | Comments
General Equations
BISECT | =LAMBDA(f,a,b,[tol],[iters], LET(tol, IF(ISOMITTED(tol), 0.000000001, tol), iters, IF(ISOMITTED(iters), 100, iters), fa, f(a), fb, f(b), IF(SIGN(fa)=SIGN(fb), 0, LET(mid, (a+b)/2, fmid, f(mid), IF(OR(iters<=0, ABS(fa-fmid) <= tol), mid, IF(fa*fmid < 0, BISECT(f, a, mid, tol, iters-1), BISECT(f, mid, b, tol, iters-1)) ) ) ) ) ) | Find zeros of the function f between a and b using bisection
NONBLANK | =LAMBDA(array, FILTER(array, array<>"") ) | Remove blank items from the array
Method S2 Functions
FSCORES | =LAMBDA(ti,ri,ni,model,params, LET(n, SUM(ni), k, COUNTA(ni), gi, INT(ni*ri), m, COUNTA(params), setwo, SUM(ni*ri*(1-ri)) / (n-k), sonetwo, SUM(ni*(ri-model(ti, params))^2) / (k-m), Fval, sonetwo/setwo, Fcrit, F.INV(0.95, k-m, n-k), C, F.DIST(Fval, k-m, n-k, TRUE)*100, accept, Fval<Fcrit, VSTACK(Fval,Fcrit,C,accept) ) ) | Calculate F-score and rejection confidence for one S2 model. Returns stack of {F, Fcritical, Rejection Conf, F<Fcrit}.
MLE_MIN | =LAMBDA(ti,ri,ni,model,theta, LET(iterations, 20, thetar, theta, loop, REDUCE(thetar, SEQUENCE(iterations), LAMBDA(acc,i, LET(thetar, INDEX(acc, i, 0), rr, model(ti, thetar), rdiff, ri-rr, ww, DIAG(ni/(rr*(1-rr))), h, thetar/10000, dd, (model(ti, thetar+h) - rr)/h, ddtrans, TRANSPOSE(dd), b, TRANSPOSE(MMULT(MMULT(MMULT(MINVERSE(MMULT(MMULT(ddtrans, ww), dd)), ddtrans), ww), rdiff)), VSTACK(acc, thetar+IFERROR(b,0)) ) )), INDEX(loop, ROWS(loop), 0) ) ) | Minimize the negative log likelihood function, RP1 equation D-12
NGROUP | =LAMBDA(x, LET(xsort, SORT(x), d, DIFF(xsort), meand, AVERAGE(d)*1.5, newgroup, d>meand, groupidx, REDUCE(0,newgroup,LAMBDA(acc,k, VSTACK(acc,IF(k,INDEX(acc,ROWS(acc))+1,INDEX(acc,ROWS(acc)))) )), cnt, MAP(groupidx, LAMBDA(k, ROWS(FILTER(groupidx, groupidx=k)))), HSTACK(x, cnt) ) ) | Get number of S2 results in the same "group" of similar intervals
S2 Fit Models
S2_DRIFT | =LAMBDA(t,p, LET(a, INDEX(p,1), b, INDEX(p,2), c, INDEX(p,3), p, NORM.S.DIST(a+c*t, TRUE), q, NORM.S.DIST(b-c*t, TRUE), p+q-1 ) )
S2_EXP | =LAMBDA(t,p, EXP(-t*INDEX(p,1)))
S2_GAMMA | =LAMBDA(t,p, LET(a, INDEX(p,1), EXP(-a*t) * (1 + a*t + ((a*t)^2)/2 + ((a*t)^3)/6) ) )
S2_LOGNORM | =LAMBDA(t,p, LET(a, INDEX(p,1), b, INDEX(p,2), IFERROR(1 - NORM.S.DIST(LN(a*t)/b, TRUE), 1) ) )
S2_MIXEDEXP | =LAMBDA(t,p, LET(a, INDEX(p,1), b, INDEX(p,2), (1 + a*t)^(-b) ) )
S2_MORTALITY | =LAMBDA(t,p, LET(a, INDEX(p,1), b, INDEX(p,2), EXP(-(a*t + b*t^2)) ) )
S2_RANDWALK | =LAMBDA(t,p, LET(a, INDEX(p,1), b, INDEX(p,2), q, 1/SQRT(a + t*b), 2*NORM.S.DIST(q, TRUE)-1 ) )
S2_RESTRICTEDWALK | =LAMBDA(t,p, LET(a, INDEX(p,1), b, INDEX(p,2), c, INDEX(p,3), q, 1/SQRT(a + b*(1-EXP(-c*t))), 2*NORM.S.DIST(q, TRUE)-1 ) )
S2_WARRANTY | =LAMBDA(t,p, LET(a, INDEX(p,1), b, INDEX(p,2), 1/(1 + EXP(a*(t-b))) ) )
S2_WEIBULL | =LAMBDA(t,p, LET(a, INDEX(p,1), b, INDEX(p,2), EXP(-((a*t)^b)) ) )
GETMODEL | =LAMBDA(name, IFS(name="Exponential", S2_EXP, name="Mixed Exponential", S2_MIXEDEXP, name="Weibull", S2_WEIBULL, name="Random Walk", S2_RANDWALK, name="Restricted Walk", S2_RESTRICTEDWALK, name="Gamma", S2_GAMMA, name="Mortality Drift", S2_MORTALITY, name="Warranty", S2_WARRANTY, name="Drift", S2_DRIFT, name="Lognormal", S2_LOGNORM, TRUE(), NA() ) ) | Get S2 model function by name
V1 Functions
STUDENTT | =LAMBDA(z,v, IFS(z=0, 0.5, z<0, 0.5*(1-BETA.DIST(z*z/(v+z*z), 0.5, v/2, TRUE)), TRUE, 0.5*(1+BETA.DIST(z*z/(v+z*z), 0.5, v/2, TRUE)) ) ) | Calculate CDF of Student T. Excel's version doesn't handle fractional degrees of freedom. Based on Uses https://stats.stackexchange.com/questions/394978/how-to-approximate-the-student-t-cdf-at-a-point-without-the-hypergeometric-funct
V1FIT | =LAMBDA(t,dt,order, LET(x, MAKEARRAY(ROWS(t), order, LAMBDA(I,j, IFS(j=1, INDEX(t, I), j=2, INDEX(t, I)^2, j=3, INDEX(t, I)^3))), xt, TRANSPOSE(x), params, TRANSPOSE(MMULT(MMULT(MINVERSE(MMULT(xt, x)), xt), dt)), resid, REDUCE(dt, SEQUENCE(order), LAMBDA(acc,k, acc-INDEX(params,k)*t^k)), ss, MMULT(TRANSPOSE(resid), resid), degf, ROWS(t) - order, cov, ss/degf * MINVERSE(MMULT(xt, x)), VSTACK(params, HSTACK(SQRT(SUM(resid^2)/degf), degf), cov) ) ) | Calculate curve fit of t, dt values with given polynomial order. Returns vstack of coefficients, Syx¦Degf, covariance.
UCONF | =LAMBDA(x,fit, LET(n, ROWS(fit)-2, params, TAKE(INDEX(fit, 1, 0), 1, n), cov, TAKE(fit, -n, n), tprime, MAP(SEQUENCE(n), LAMBDA(k, x^k)), SQRT(MMULT(MMULT(TRANSPOSE(tprime), cov), tprime)) ) ) | Calculate confidence band of curve fit at x given results of V1FIT method
YVAL | =LAMBDA(x,fit, LET(n, ROWS(fit)-2, params, TAKE(INDEX(fit, 1, 0), 1, n), tprime, MAP(SEQUENCE(n), LAMBDA(k, x^k)), MMULT(TRANSPOSE(tprime), TRANSPOSE(params)) ) ) | Calculate predicted y value at x using results of V1FIT method
RELIABILITY | =LAMBDA(x,LL,UL,yinit,uinit,dfinit,fit, LET(conf, UCONF(x, fit), fitdf, INDEX(fit, 2, 2), syx, INDEX(fit, 2, 1), mu, yinit + YVAL(x, fit), sigb, SQRT(conf^2 + uinit^2), degf, sigb^4 / ((conf^4/fitdf + uinit^4/dfinit)), RU, IFERROR(STUDENTT((UL-mu)/sigb, degf), 1), RL, IFERROR(STUDENTT((LL-mu)/sigb, degf), 0), RU-RL ) ) | Calculate reliability at time x for V1 method
Utility Functions from Vertex42.com (MIT Licensed)
DIAG | =LAMBDA(array,[k], LET(doc,"https://www.vertex42.com/lambda/diag.html", version,"1/2/2026 - Added [k] option", nR, ROWS(array), nC, COLUMNS(array), k, IF(ISOMITTED(k),0,k), isVector, OR(nR=1, nC=1), IF(isVector, LET( v, TOCOL(array), nV, ROWS(v), m, nV+ABS(k), MAKEARRAY(m,m,LAMBDA(i,j, IF( j-i = k, LET(p, IF(k>=0,i,j), IF(AND(p>=1,p<=nV),INDEX(v,p),0) ), 0 ) )) ), LET( nK, IF(k>=0, MAX(0, MIN(nR, nC-k)), MAX(0, MIN(nR+k, nC))), i_0, IF(k>=0,1,1-k), j_0, IF(k>=0,1+k,1), IF(nK=0, NA(), MAKEARRAY(nK,1,LAMBDA(t,_,INDEX(array,i_0+t-1,j_0+t-1))) ) ) ) )) | Generate diagonal matrix, or get diagonal of matrix
DIFF | = LAMBDA(x,[n], LET(doc,"https://www.vertex42.com/lambda/diff.html", n,IF(ISBLANK(n),1,n), REDUCE(1,SEQUENCE(n),LAMBDA(acc,i, IF(i=1, DROP(x,1)-DROP(x,-1), DROP(acc,1)-DROP(acc,-1) ) )) )) | Elementwise difference of array
LINSPACE | = LAMBDA(start,end,n, LET(doc,"https://www.vertex42.com/lambda/linspace.html", SEQUENCE(n,1,start,(end-start)/(n-1)) )) | Array of equally spaced points