## GUM Instructions
Sandia Primary Standards Lab - Excel Uncertainty Calculator
v0.2
uncertainty@sandia.gov
This spreadsheet provides tools for determining measurement uncertainty. It includes Type A (repeatability and reproducibility) and Type B uncertainty evaluation, plus propagation of uncertainties through a measurement equation using the GUM and Monte Carlo approaches, accounting for correlated input variables. Green cells are user inputs, blue cells are calculated output values. Gray cells, often hidden, are intermediate calculated values or plot data. No Excel Macros are used, however the calculations make use of Excel's LAMBDA function which is only available in Office 365 or 2024.
Spreadsheet Tabs
GUM Uncertainty | Propagate uncertainty through a measurement model f(x1, x2, x3), where each x input has uncertainty. Uses the GUM equation and Monte Carlo method for comparison.
Units | Defines base units and custom units. Add units to the custom table if they are not defined.
Type B | Use this tab for computing and combining Type B uncertainties of measuring equipment, such as accuracy, resolution, and linearity. The results may be linked to the GUM Uncertainty Tab.
Repeatability | Calculate uncertainty due to repeatability (N repeated measurements under similar conditions). Tab may be duplicated and results linked to GUM Uncertainty tab.
Reproducibility | Calculate uncertainty due to reproducibility (N repeated measurements over M reproducibility conditions, such as M operators, methods, days, etc.)
Repeatability - Autocorrelation | Calculate Type A repeatability uncertainty accounting for autocorrelation
Student-T | Calculate coverage factor, degrees of freedom, and level of confidence from Student T distribution
GUM Examples | Examples from literature and training courses that may be copied into the GUM Uncertainty tab
Manual Monte Carlo | Generate random numbers from probability distributions and combine through a formula. Used primarily for training purposes and learning the basis of Monte Carlo techniques. Deleting this tab can greatly reduce the file size if not needed.
Formula Reference | Lists the LAMBDA functions used in Excel's Name Manager for readability. Hidden by default.
Lookups | Provides allowable options for dropdown menus. Hidden by default.
GUM Uncertainty Instructions
1. Press CTRL+F3 to open Excel's Name Manager (Or Click "Name Manager" from the "Formulas" menu bar)
2. Change the __MEASUREMENT_MODEL function to refer to the measurement equation as a LAMBDA function. The LAMBDA contains up to 8 variable names, followed by the formula to compute. For example, "=LAMBDA(delta, t, rs, p, SQRT(2*delta*t*rs/p))" computes the pitot tube example described by Kline and McClintock in their 1953 paper.
3. Enter Variable names, values, and uncertainties in the Variables table, in the same order as defined in the LAMBDA in step 2.
4. To add correlations between variables, unhide columns N-V and edit the correlation matrix.
5. GUM and Monte Carlo results will be computed, providing standard uncertainty, expanded uncertianty, and 95% coverage limits.
6. The GUM vs Monte Carlo comparison can be used to validate the GUM equation and check for nonlinearities in the model. It compares the 95% coverage limits and determines whether they fall within a tolerance δ. See Section 8 of JCGM 101:2008.
To add multiple GUM calculations to one spreadsheet, duplicate the GUM Uncertainty tab, add another function to the Name Manager, and edit the three cells with dashed outline to reference the new function instead of __MEASUREMENT_MODEL.
Limitations
1) | Excel's random number generator cannot be seeded.
2) | The measurement model may have up to 8 variables. More could be added by modifying the Named Formulas.
3) | While unit conversions are done, there is no validation of dimensional analysis. All values are converted to SI base units, processed through the GUM model, then the result converted from the assumed base unit given the provided output units.
About
Sandia National Laboratories is a multimission laboratory managed and operated by National Technology & Engineering Solutions of Sandia, LLC, a wholly owned subsidiary of Honeywell International Inc., for the U.S. Department of Energy’s National Nuclear Security Administration under contract DE-NA0003525.
SAND2026-16029O

## GUM Uncertainty
Press CTRL+F3 to open Name Manager, then edit the __MEASUREMENT_MODEL function to define the measurement equation. | [22 empty] | ← Correlation Matrix is hidden in columns here. Unhide to enter correlated variables.
Settings
Level of Confidence | 95%
Monte Carlo Samples | 1000
Sig. Figs. for Comparison | 1
Measurand units | m/s
Variable | Value | Uncertainty | Distribution | Units | Deg. Freedom | Variable | Proportion | Standard Uncertainty | Base Units | Value (Base Units) | Uncertainty (Base Units) | Sensitivity | Sensitivity Units | Correlation Matrix
Δ | 8 | 0.1 | Normal k=2 | inH2O | 1000000 | Δ | 27.27% | 12.455 | Pa | 1992.8 | 24.91 | 0.014425 | (m/s) / (Pa) |  | Δ | T | Rs | p
T | 67.4 | 0.2 | Normal k=2 | F | 1000000 | T | 0.02513% | 0.055556 | K | 292.8167 | 0.11111 | 0.098172 | (m/s) / (K) | Δ | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Rs | 287.06 | 0 | Normal k=2 | J/K/kg | 1000000 | Rs | 0% | 0 | J/K/kg | 287.06 | 0 | 0 | (m/s) / (J/K/kg) | T | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
p | 14.7 | 0.3 | Normal k=2 | psi | 1000000 | p | 72.7% | 1034.2136 | Pa | 101352.9322 | 2068.4272 | -0.00028363 | (m/s) / (Pa) | Rs | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | - | [7 empty] | p | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0
[6 empty] | - | [8 empty] | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
[6 empty] | - | [8 empty] | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0
[6 empty] | - | [8 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | Correlations | 0% | [7 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1
[6 empty] | GUM Results |  |  |  | Monte Carlo Results
[6 empty] | Measurand Value | 57.4927 | m/s |  | Mean | 57.5033 | m/s
[6 empty] | Standard Deviation | 0.34402 | m/s |  | Standard Deviation | 0.33804 | m/s
[6 empty] | Variance | 0.11835 | (m/s)^2 |  | Minimum | 56.5408 | m/s
[6 empty] | Effective Deg. Freedom | 1658574.0684 |  |  | 1st Quartile | 57.296 | m/s
[6 empty] | Coverage Factor (k) | 1.96 |  |  | Median | 57.5043 | m/s
[6 empty] | Expanded Uncertainty ± | 0.67427 | m/s |  | 3rd Quartile | 57.7237 | m/s
[10 empty] | Maximum | 58.7019 | m/s
[6 empty] | Coverage Region Low | 56.8185 | m/s |  | Coverage Region Low | 56.7745 | m/s
[6 empty] | Coverage Region High | 58.167 | m/s |  | Coverage Region High | 58.1461 | m/s | [10 empty] | ⬐ Plot histogram data spills into to hidden rows here ↴
[10 empty] | Plot Data | 56.5408 | 0.025214
[11 empty] | 56.5948 | 0.038461
[11 empty] | 56.6488 | 0.057238
[11 empty] | 56.7029 | 0.083106
[11 empty] | 56.7569 | 0.11773
[11 empty] | 56.8109 | 0.1627
[11 empty] | 56.8649 | 0.21939
[11 empty] | 56.919 | 0.28862
[11 empty] | 56.973 | 0.37044
[11 empty] | 57.027 | 0.46387
[11 empty] | 57.0811 | 0.56672
[11 empty] | 57.1351 | 0.67551
[11 empty] | 57.1891 | 0.78557
[11 empty] | 57.2431 | 0.8913
[11 empty] | 57.2972 | 0.98662
[11 empty] | 57.3512 | 1.0655
[11 empty] | 57.4052 | 1.1227
[11 empty] | 57.4592 | 1.1542
[11 empty] | 57.5133 | 1.1576
[11 empty] | 57.5673 | 1.1327
[11 empty] | 57.6213 | 1.0814
[11 empty] | 57.6754 | 1.0072
[11 empty] | 57.7294 | 0.91531
[11 empty] | 57.7834 | 0.81151
[11 empty] | 57.8374 | 0.70196
[11 empty] | 57.8915 | 0.5924
[11 empty] | 57.9455 | 0.48776
[11 empty] | 57.9995 | 0.39182
[11 empty] | 58.0535 | 0.30709
[11 empty] | 58.1076 | 0.23481
[11 empty] | 58.1616 | 0.17517
[11 empty] | 58.2156 | 0.1275
[11 empty] | 58.2697 | 0.090539
[11 empty] | 58.3237 | 0.062726
[11 empty] | 58.3777 | 0.042399
[11 empty] | 58.4317 | 0.027961
[11 empty] | 58.4858 | 0.01799
[11 empty] | 58.5398 | 0.011293
[11 empty] | 58.5938 | 0.0069161
[11 empty] | 58.6478 | 0.0041325
[11 empty] | 58.7019 | 0.002409
[10 empty] | Histogram Density | 0.074037
0.11105
0.11105
0.12956
0.14807
0.055527
0.22211
0.24062
0.4072
0.46273
0.6108
0.70335
0.64782
0.74037
1.0735
0.94397
1.4622
1.3141
1.2956
1.055
1.1291
1.018
0.8144
0.83291
0.74037
0.46273
0.35167
0.38869
0.33316
0.22211
0.11105
0.074037
0.11105
0.074037
0
0
0.018509
0
0
0
GUM versus Monte Carlo Comparison
[6 empty] | Significant Figures | 1
[6 empty] | δ | 0.05
[6 empty] | Low Limit Agree? | ✓
[6 empty] | High Limit Agree? | ✓

## Units
Base Units for Excel Built-in to CONVERT |  | Base Unit conversion factors for ones Excel doesn't know
Unit | Base | Unit | Factor | Base
m | m |  | 1
cm | m | V | 1 | V
mm | m | uV | 1e-06 | V
um | m | mV | 0.001 | V
nm | m | kV | 1000 | V
km | m | MV | 1000000 | V
in | m | deltaC | 1 | K
uin | m | deltaF | 0.55556 | K
mi | m | deltaK | 1 | K
yd | m | A | 1 | A
ang | m | uA | 1e-06 | A
ly | m | mA | 0.001 | A
ug | kg | kA | 1000 | A
mg | kg | MA | 1000000 | A
g | kg | Ohm | 1 | Ohm
kg | kg | uOhm | 1e-06 | Ohm
ozm | kg | mOhm | 0.001 | Ohm
lbm | kg | kOhm | 1000 | Ohm
s | s | MOhm | 1000000 | Ohm
us | s | GOhm | 1000000000 | Ohm
ms | s | Ω | 1 | Ω
ns | s | mΩ | 0.001 | Ω
ps | s | uΩ | 1e-06 | Ω
fs | s | kΩ | 1000 | Ω
as | s | MΩ | 1000000 | Ω
yr | s | GΩ | 1000000000 | Ω
day | s | J/K/kg | 1 | J/K/kg
hr | s | inH2O | 249.1 | Pa
min | s | F | 1 | F
Pa | Pa | mF | 0.001 | F
atm | Pa | uF | 1e-06 | F
mmHg | Pa | nF | 1e-09 | F
psi | Pa | pF | 1e-12 | F
Torr | Pa | fF | 1e-15 | F
N | N | Hz | 1 | Hz
dyn | N
lbf | N
J | J
flb | J
BTU | J
cal | J
W | W
HP | W
T | T
ga | T
C | K
F | K
K | K
Rank | K
m3 | m3
tsp | m3
tbs | m3
cup | m3
qt | m3
gal | m3
L | m3
barrel | m3
bushel | m3
ft3 | m3
in3 | m3
ly3 | m3
yd3 | m3
m2 | m2
us_acre | m2
ft2 | m2
ha | m2
in2 | m2
ly2 | m2
mi2 | m2
yd2 | m2
m/s | m/s
mph | m/s
km/h | m/s
km/s | m/s
cm/s | m/s
um/s | m/s
cm/hr | m/s
um/hr | m/s
kn | m/s

## Type B
Duplicate these boxes for each variable having multiple Type B components. Link the uncertainty and degrees of freedom to the GUM Uncertainty tab.
Variable Name: _______
| Uncertainty | Distribution | k | Degrees Freedom
Accuracy | 0.25 | Uniform k=1.73 | 1.7321 | 50
Resolution | 0.01 | Normal k=2 | 2 | 50
Linearity | 0 | Normal k=2 | 2 | 1000000
Other | 0 | Normal k=2 | 2 | 1000000
|  | Total Standard Uncertainty |  | 0.14442
|  | Effective Degrees of Freedom |  | 50.12
Variable Name: _______
| Uncertainty | Distribution | k | Degrees Freedom
Accuracy | 0.25 | Normal k=2 | 2 | 50
Resolution | 0.01 | Normal k=2 | 2 | 50
Linearity | 0 | Normal k=2 | 2 | 1000000
Other | 0 | Normal k=2 | 2 | 1000000
|  | Total Standard Uncertainty |  | 0.1251
|  | Effective Degrees of Freedom |  | 50.16
Variable Name: _______
| Uncertainty | Distribution | k | Degrees Freedom
Accuracy | 0.25 | Normal k=2 | 2 | 50
Resolution | 0.01 | Normal k=2 | 2 | 50
Linearity | 0 | Normal k=2 | 2 | 1000000
Other | 0 | Normal k=2 | 2 | 1000000
|  | Total Standard Uncertainty |  | 0.1251
|  | Effective Degrees of Freedom |  | 50.16

## Repeatability
9.9727 | ← Enter repeated measurement values in Column B
10.2405
10.2258 |  | Enter Additional Type B Uncertainties Here: |  |  |  |  | Uncertainty in THIS mean value |  | ← Use this box, for example, when measuring the temperature of an oven used during a specific test
10.0813 |  |  | Uncertainty | Distribution | k | Degrees Freedom | Standard Deviation | 0.16821
10.2493 |  | Accuracy | 0 | Normal k=2 | 2 | 50 | Standard Error of the Mean | 0.037612
10.0303 |  | Resolution | 0 | Normal k=2 | 2 | 50
9.7967 |  | Linearity | 0 | Normal k=2 | 2 | 1000000 | Mean | 10.0517
10.1907 |  | Other | 0 | Normal k=2 | 2 | 1000000 | Repeatability Standard Uncertainty | 0.037612
10.1888 | [6 empty] | Degrees Freedom | 19
10.1413
9.7663 | [6 empty] | Total Standard Uncertainty | 0.037612
10.1827 | [6 empty] | Effective Degrees of Freedom | 19
10.2104
9.8051
9.9631 | [6 empty] | Applied to uncertainty of FUTURE measurements |  | ← Use this box, for example, when characterizing the temperature of the lab to apply the temperature average and uncertainty to future measurements
10.0689 | [6 empty] | Number of future repeat measurements | 1
10.0694
9.8415 | [6 empty] | Mean | 10.0517
9.8147 | [6 empty] | Repeatability Standard Uncertainty | 0.16821
10.1946 | [6 empty] | Degrees Freedom | 19
[7 empty] | Total Standard Uncertainty | 0.16821
[7 empty] | Effective Degrees of Freedom | 19

## Reproducibility
Enter reproducibility measurements in table below (expand as necessary)
Run ↴ \ Group → | A | B | C | D | E | F | G | H | I | J | Enter Additional Type B Uncertainties Here:
1 | 0.95101 | 1.0089 | 1.0132 | 0.96616 | 1.0071 | 0.96762 | 0.95905 | 1.0167 | 1.0331 | 0.98005 |  | Uncertainty | Distribution | k | Degrees Freedom
2 | 0.95833 | 0.95468 | 0.99785 | 0.98744 | 1.0267 | 1.0039 | 1.0059 | 1.018 | 1.04 | 0.9586 | Accuracy | 0 | Normal k=2 | 2 | 50
3 | 0.95919 | 0.96017 | 0.99752 | 0.95252 | 1.0392 | 0.97402 | 1.0059 | 1.0222 | 1.0452 | 0.96819 | Resolution | 0 | Normal k=2 | 2 | 50
4 | 1.0496 | 1.0379 | 1.0229 | 0.98583 | 1.018 | 0.95183 | 0.97295 | 0.95036 | 1.0358 | 0.98732 | Linearity | 0 | Normal k=2 | 2 | 1000000
5 | 1.0319 | 0.9628 | 0.99968 | 0.98254 | 1.0365 | 1.0035 | 1.0223 | 0.9701 | 1.0364 | 0.96474 | Other | 0 | Normal k=2 | 2 | 1000000
6 | 1.0327 | 0.96672 | 1.0192 | 0.99218 | 0.98317 | 0.97926 | 1.008 | 1.0382 | 0.98304 | 1.0278
7 | 1.0048 | 1.0477 | 0.99156 | 0.99653 | 1.0337 | 1.0236 | 1.0197 | 0.95418 | 1.0414 | 0.97323
8 | 1.0475 | 0.95285 | 0.97403 | 0.96524 | 0.95035 | 1.0291 | 1.0167 | 0.97867 | 0.9922 | 1.0353
9 | 0.95565 | 1.0293 | 0.98272 | 1.0131 | 0.96504 | 0.96971 | 0.96446 | 0.96961 | 0.96958 | 1.0094
10 | 0.99312 | 0.95638 | 1.0153 | 0.95917 | 0.95145 | 1.0214 | 1.0486 | 1.0413 | 1.0279 | 1.037
Reproducibility | B: Deg. F.; D: Uncertainty in THIS mean value; J: ← Use this box, for example, when measuring the temperature of an oven used during a specific test
Grand Mean | A: 0.9974; D: Standard Uncertainty; H: 0.0031037
Repeatability Std. Dev. | A: 0.030765; B: 90; D: Degrees Freedom; H: 99
Reproducibility Std. Dev. | A: 0.010637; B: 9
|  |  |  | Total Standard Uncertainty |  |  |  | 0.0031037
|  |  |  | Effective Degrees of Freedom |  |  |  | 99
|  |  |  | Applied to uncertainty of FUTURE measurements | [5 empty] | ← Use this box, for example, when characterizing the temperature of the lab to apply that uncertainty to future measurements
|  |  |  | Number of new repeat measurements |  |  |  | 1
|  |  |  | Standard Uncertainty |  |  |  | 0.030765
|  |  |  | Degrees Freedom |  |  |  | 99
[7 empty] | Total Standard Uncertainty | 0.030765
[7 empty] | Effective Degrees of Freedom | 99
Group Averages Vj: | 0.99838 | 0.98774 | 1.0014 | 0.98007 | 1.0011 | 0.9924 | 1.0024 | 0.99593 | 1.0205 | 0.99416
Group Variances s^2(Vjk): | 0.001624 | 0.001489 | 0.00025959 | 0.00035653 | 0.0012679 | 0.00074507 | 0.00080825 | 0.0012143 | 0.00076907 | 0.00093093
Group Std. Devs s(Vjk): | 0.040299 | 0.038588 | 0.016112 | 0.018882 | 0.035607 | 0.027296 | 0.02843 | 0.034847 | 0.027732 | 0.030511
# Measurements nj: | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
Deg. Freedom | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9
# Groups J: | A: 10
Total Measurements: | A: 100
|  | ANOVA | Value | Deg. Freedom
|  | F-test Confidence | 95%
|  | Variance of Means s^2(Vj) | 0.00011315 | 9
| Between Group Variance | sa^2 | 0.0011315 | 9
| Repeatability Variance | sb^2 | 0.00094646 | 90
|  | F | 1.1955
|  | F critical | 1.9856
|  | Between-Group Effect? | FALSE
|  | Uncertainty if no between-group effect | 0.0031037 | 99
|  | Uncertainty if between-group effect | 0.0033637 | 9

## Student-T
Student-T Distribution Calculators
Degrees of Freedom | 30 | Degrees of Freedom | 30 | Confidence | 95%
Confidence | 95% | Coverage Factor | 2 | Coverage Factor | 2.7
Coverage Factor | 2.0423 | Confidence | 94.54% | Degrees Freedom | 4.3075
Note: Excel's T.INV.2T is limited to integer degrees of freedom. STUDENTT lambda function used here implements G.W.Hill, "Algorithm 396".

## GUM Examples
Example GUM Calculations. Copy the LAMBDA equation into Excel's Name Manager in the __MEASUREMENT_MODEL function. Then copy the variables table into the GUM Uncertainty worksheet.
Film Thickness corrected for bias | [5 empty] | Example from SNL ENGR224 Uncertainty class
=LAMBDA(a,b,c, a/b*c) | [5 empty] | Note: Don't use "X1", "X2" in formula because Excel thinks these refer to cells X1 and X2 in the sheet!
Variable | Value | Uncertainty | Distribution | Units | Deg. Freedom
X1 | 0.182 | 0.00093 | Normal k=1 | um | 9 | Calibration lab's measurement of thickness standard
X2 | 0.1823 | 0.00058 | Normal k=1 | um | 19 | Our measurement of thickness standard
Yu | 0.6978 | 0.0026 | Normal k=1 | um | 19 | Our uncorrected measurement of item under test
Pitot Tube Airspeed Calculation | [5 empty] | Example from Kline and McClintock's 1953 paper originating the GUM method
=LAMBDA(delta,t,rs,p,SQRT(2*delta*t*rs/p))
Variable | Value | Uncertainty | Distribution | Units | Deg. Freedom
Δ | 8 | 0.1 | Normal k=2 | inH2O | 1000000 | Measured difference in static vs total pressure
T | 67.4 | 0.2 | Normal k=2 | F | 1000000 | Ambient Temperature
Rs | 287.06 | 0 | Normal k=2 | J/K/kg | 1000000 | Specific gas constant for dry air. Assuming no uncertainty.
p | 14.7 | 0.3 | Normal k=2 | psi | 1000000 | Ambient pressure
Time constant of RC circuit | [5 empty] | Example from SNL ENGR224 Uncertainty class
=LAMBDA(x,y,z, x*(y+z)) | [5 empty] | Note: Don't use "R1", "C2", "C3" in formula because Excel thinks these refer to cells R1, C2, and C3 in the sheet!
Variable | Value | Uncertainty | Distribution | Units | Deg. Freedom
R1 | 32.201 | 0.0044 | Normal k=2 | kΩ | 15.7 | Resistance
C2 | 0.1 | 0.005 | Uniform k=1.73 | uF | 1000000 | Capacitance
C3 | 0.22 | 0.01 | Uniform k=1.73 | uF | 1000000 | Capacitance
Viscosity of test liquid | [5 empty] | Example E3 from NIST TN-1900
=LAMBDA(muc,rhob,rhoc,rhom,tc,tm, muc*tm*(rhob-rhom)/(tc*(rhob-rhoc)))
Variable | Value | Uncertainty | Distribution | Units | Deg. Freedom
mu | 4.63 | 0.0463 | Normal k=1 | mPa.s | 1000000 | Viscosity of calibration liquid
rho_b | 2217 | 0.5 | Normal k=1 | kg/m3 | 1000000 | Density of ball
rho_c | 810 | 0.5 | Normal k=1 | kg/m3 | 1000000 | Density of calibration liquid
rho_m | 1180 | 0.5 | Normal k=1 | kg/m3 | 1000000 | Densityof sodium hydroxide
t_c | 36.6 | 5.49 | Normal k=1 | s | 1000000 | Travel time in calibration liquid
t_m | 61 | 6.1 | Normal k=1 | s | 1000000 | Travel time in sodium hydroxide

## Manual Monte Carlo
This tab is primarily for training purposes, allowing for manual exploration of probability distributions and combining them through simple formulas.
Settings | [5 empty] | Statistics | [13 empty] | Distributions
Random Samples | 5000 | [5 empty] | X | Y | Z | F(X, Y, Z) | [9 empty] | Normal
Histogram Bins | 100 |  |  |  |  | Mean | 0.015656 | 0.026106 | 0.014864 | 2.9888 | [9 empty] | Uniform | ←Plot data hidden in these columns
[6 empty] | Minimum | -3.358 | -3.6652 | -3.9769 | 0.00040454 | [9 empty] | Triangular
Variables | [5 empty] | 1st Quartile | -0.6211 | -0.6518 | -0.66678 | 1.1696 |  | x1 |  | x2 |  | x3 |  | Combined
| X | Y | Z |  |  | Median | 0.01294 | 0.01473 | 0.020174 | 2.3748 | Max | 3.6958 |  | 3.3986 |  | 4.2633 |  | 20.1422
Distribution | Normal | Normal | Normal |  |  | 3rd Quartile | 0.6784 | 0.70056 | 0.6856 | 4.0987 | Min | -3.358 |  | -3.6652 |  | -3.9769 |  | 0.00040454
Mean | 0 | 0 | 0 |  |  | Maximum | 3.6958 | 3.3986 | 4.2633 | 20.1422 | Bin Width | 0.070538 |  | 0.070638 |  | 0.082402 |  | 0.20142
a | 1 | 1 | 1 |  |  | 95% Interval Low | -1.9428 | -1.9428 | -1.9474 | 0.20881
[6 empty] | 95% Interval High | 1.9474 | 2.0153 | 1.9774 | 9.4262 |  | Bins | Counts | Bins | Counts | Bins | Counts | Bins | Counts
|  |  |  | F(X, Y, Z) | =X^2+Y^2+Z^2 | [6 empty] | -3.2874 | 1 | -3.5946 | 2 | -3.8945 | 1 | 0.20182 | 122
| 1.2013 | -0.67823 | 0.18827 | 1.9387 | ← Enter formula for F in terms of named ranges "X", Y", "Z" | [6 empty] | -3.2169 | 0 | -3.524 | 1 | -3.8121 | 0 | 0.40324 | 203
| -1.0704 | 0.0049948 | 0.077353 | 1.1519 | [7 empty] | -3.1464 | 0 | -3.4533 | 1 | -3.7297 | 1 | 0.60466 | 216
| 0.26705 | -0.44095 | 0.99763 | 1.261 | [7 empty] | -3.0758 | 1 | -3.3827 | 1 | -3.6473 | 0 | 0.80608 | 243
| 1.2029 | -0.19614 | 2.0091 | 5.5217 | [7 empty] | -3.0053 | 1 | -3.312 | 0 | -3.5649 | 0 | 1.0075 | 271
| -0.53585 | 2.9186 | -0.38049 | 8.9503 | [7 empty] | -2.9347 | 1 | -3.2414 | 0 | -3.4825 | 0 | 1.2089 | 248
| 0.87614 | 0.56556 | 0.44625 | 1.2866 | [7 empty] | -2.8642 | 1 | -3.1708 | 0 | -3.4001 | 0 | 1.4103 | 239
| 0.30452 | -0.22545 | -1.0423 | 1.23 | [7 empty] | -2.7937 | 1 | -3.1001 | 1 | -3.3177 | 0 | 1.6117 | 227
| 0.96474 | 1.1084 | -0.86505 | 2.9076 | [7 empty] | -2.7231 | 7 | -3.0295 | 1 | -3.2353 | 1 | 1.8132 | 235
| -1.4141 | -0.40997 | -0.2412 | 2.2259 | [7 empty] | -2.6526 | 3 | -2.9588 | 1 | -3.1529 | 1 | 2.0146 | 203
| -1.6397 | 0.94842 | -0.29986 | 3.6781 | [7 empty] | -2.5821 | 5 | -2.8882 | 3 | -3.0705 | 0 | 2.216 | 166
| -2.2348 | 1.4805 | 1.0738 | 8.3392 | [7 empty] | -2.5115 | 4 | -2.8176 | 2 | -2.9881 | 3 | 2.4174 | 170
| -1.5855 | -0.43229 | 0.22044 | 2.7492 | [7 empty] | -2.441 | 11 | -2.7469 | 0 | -2.9057 | 2 | 2.6188 | 168
| 0.16651 | -1.4008 | -0.11687 | 2.0037 | [7 empty] | -2.3704 | 11 | -2.6763 | 3 | -2.8233 | 3 | 2.8203 | 165
| -0.39007 | -0.69488 | -0.75531 | 1.2055 | [7 empty] | -2.2999 | 8 | -2.6057 | 2 | -2.7409 | 4 | 3.0217 | 152
| 0.95093 | -0.10265 | 0.30132 | 1.0056 | [7 empty] | -2.2294 | 10 | -2.535 | 4 | -2.6585 | 5 | 3.2231 | 161
| 1.7436 | -1.1876 | 1.5491 | 6.8504 | [7 empty] | -2.1588 | 16 | -2.4644 | 9 | -2.5761 | 4 | 3.4245 | 134
| 0.55552 | -0.75541 | 0.87272 | 1.6409 | [7 empty] | -2.0883 | 13 | -2.3937 | 10 | -2.4937 | 11 | 3.6259 | 138
| -1.542 | -0.9773 | 0.25492 | 3.3977 | [7 empty] | -2.0178 | 11 | -2.3231 | 6 | -2.4113 | 5 | 3.8273 | 139
| 1.0392 | 0.18875 | 0.077678 | 1.1216 | [7 empty] | -1.9472 | 18 | -2.2525 | 9 | -2.3289 | 13 | 4.0288 | 114
| 0.85989 | -0.20309 | -0.91113 | 1.6108 | [7 empty] | -1.8767 | 27 | -2.1818 | 16 | -2.2465 | 12 | 4.2302 | 104
| -0.5243 | -0.89691 | 1.3926 | 3.0186 | [7 empty] | -1.8061 | 24 | -2.1112 | 8 | -2.1641 | 20 | 4.4316 | 102
| 0.091518 | 0.64497 | 0.22032 | 0.4729 | [7 empty] | -1.7356 | 30 | -2.0406 | 19 | -2.0817 | 13 | 4.633 | 87
| -1.7579 | -0.52764 | 0.82105 | 4.0426 | [7 empty] | -1.6651 | 31 | -1.9699 | 16 | -1.9993 | 14 | 4.8344 | 93
| 0.63434 | 1.7224 | 0.64477 | 3.7849 | [7 empty] | -1.5945 | 41 | -1.8993 | 26 | -1.9168 | 26 | 5.0358 | 70
| 0.65392 | -1.0449 | -1.1041 | 2.7385 | [7 empty] | -1.524 | 27 | -1.8286 | 26 | -1.8344 | 33 | 5.2373 | 72
| 0.042032 | 0.13664 | -1.2107 | 1.4862 | [7 empty] | -1.4534 | 47 | -1.758 | 21 | -1.752 | 26 | 5.4387 | 56
| -0.23985 | -0.61781 | -1.569 | 2.9009 | [7 empty] | -1.3829 | 46 | -1.6874 | 29 | -1.6696 | 27 | 5.6401 | 53
| -1.4539 | 1.1515 | -0.33644 | 3.5531 | [7 empty] | -1.3124 | 55 | -1.6167 | 32 | -1.5872 | 38 | 5.8415 | 55
| 0.13219 | 0.19469 | 1.046 | 1.1494 | [7 empty] | -1.2418 | 57 | -1.5461 | 40 | -1.5048 | 51 | 6.0429 | 49
| 0.53954 | -0.30612 | 1.0557 | 1.4994 | [7 empty] | -1.1713 | 70 | -1.4754 | 43 | -1.4224 | 52 | 6.2444 | 45
| -0.23035 | 0.50781 | 1.12 | 1.5654 | [7 empty] | -1.1008 | 63 | -1.4048 | 46 | -1.34 | 68 | 6.4458 | 46
| -0.43864 | -0.69755 | -0.44486 | 0.87688 | [7 empty] | -1.0302 | 82 | -1.3342 | 66 | -1.2576 | 68 | 6.6472 | 31
| 0.30274 | 0.025141 | -2.5347 | 6.517 | [7 empty] | -0.95969 | 73 | -1.2635 | 60 | -1.1752 | 86 | 6.8486 | 31
| -0.26097 | 0.25028 | 0.68194 | 0.59579 | [7 empty] | -0.88915 | 84 | -1.1929 | 46 | -1.0928 | 71 | 7.05 | 33
| 0.61834 | 0.27442 | -0.36137 | 0.58825 | [7 empty] | -0.81861 | 81 | -1.1223 | 63 | -1.0104 | 81 | 7.2514 | 25
| 1.0322 | -1.6121 | 0.22526 | 3.7149 | [7 empty] | -0.74807 | 92 | -1.0516 | 72 | -0.92802 | 116 | 7.4529 | 27
| -1.9463 | -1.0598 | -0.85354 | 5.6399 | [7 empty] | -0.67754 | 118 | -0.98098 | 67 | -0.84562 | 107 | 7.6543 | 23
| 1.0532 | -0.52328 | -0.46404 | 1.5983 | [7 empty] | -0.607 | 102 | -0.91034 | 86 | -0.76321 | 133 | 7.8557 | 20
| 0.55346 | 1.6964 | 0.83999 | 3.8895 | [7 empty] | -0.53646 | 126 | -0.8397 | 90 | -0.68081 | 133 | 8.0571 | 29
| 0.74551 | -1.152 | 0.89475 | 2.6835 | [7 empty] | -0.46592 | 145 | -0.76906 | 117 | -0.59841 | 135 | 8.2585 | 20
| 1.0965 | 2.1413 | -0.3099 | 5.8835 | [7 empty] | -0.39539 | 158 | -0.69843 | 118 | -0.51601 | 156 | 8.46 | 23
| 1.3451 | 0.40165 | -0.51645 | 2.2372 | [7 empty] | -0.32485 | 137 | -0.62779 | 124 | -0.4336 | 118 | 8.6614 | 11
| 0.29361 | 1.9169 | 1.7102 | 6.6854 | [7 empty] | -0.25431 | 131 | -0.55715 | 112 | -0.3512 | 137 | 8.8628 | 12
| -1.1519 | 0.21511 | -1.8571 | 4.8218 | [7 empty] | -0.18377 | 141 | -0.48651 | 119 | -0.2688 | 141 | 9.0642 | 18
| 0.33208 | -0.55967 | -0.88571 | 1.208 | [7 empty] | -0.11324 | 123 | -0.41587 | 136 | -0.1864 | 164 | 9.2656 | 17
| -1.2762 | -1.2894 | -0.91377 | 4.1264 | [7 empty] | -0.042699 | 161 | -0.34523 | 148 | -0.10399 | 149 | 9.467 | 16
| -0.010278 | 2.261 | -0.20728 | 5.1551 | [7 empty] | 0.027838 | 132 | -0.2746 | 153 | -0.021592 | 179 | 9.6685 | 8
| 0.30537 | -1.1595 | 0.94445 | 2.3296 | [7 empty] | 0.098376 | 142 | -0.20396 | 133 | 0.06081 | 170 | 9.8699 | 7
| 0.95887 | 0.47596 | 0.17608 | 1.177 | [7 empty] | 0.16891 | 144 | -0.13332 | 138 | 0.14321 | 181 | 10.0713 | 15
| -0.93075 | 0.81548 | -0.65176 | 1.9561 | [7 empty] | 0.23945 | 125 | -0.062682 | 126 | 0.22562 | 163 | 10.2727 | 7
| 1.0444 | -0.75701 | -1.0309 | 2.7265 | [7 empty] | 0.30999 | 134 | 0.0079565 | 139 | 0.30802 | 169 | 10.4741 | 14
| 1.5123 | -0.60527 | -0.9429 | 3.5424 | [7 empty] | 0.38053 | 146 | 0.078595 | 135 | 0.39042 | 164 | 10.6755 | 2
| 0.52985 | 0.94238 | 0.041542 | 1.1705 | [7 empty] | 0.45106 | 150 | 0.14923 | 134 | 0.47282 | 135 | 10.877 | 5
| -0.18164 | 1.0666 | -0.778 | 1.7759 | [7 empty] | 0.5216 | 104 | 0.21987 | 134 | 0.55522 | 138 | 11.0784 | 4
| 1.0635 | 1.5072 | -0.39142 | 3.5558 | [7 empty] | 0.59214 | 124 | 0.29051 | 154 | 0.63763 | 133 | 11.2798 | 5
| -0.15993 | -1.6404 | -1.5278 | 5.0506 | [7 empty] | 0.66268 | 121 | 0.36115 | 134 | 0.72003 | 138 | 11.4812 | 4
| 3.1033 | 1.7288 | 0.4708 | 12.8407 | [7 empty] | 0.73321 | 118 | 0.43179 | 103 | 0.80243 | 119 | 11.6826 | 6
| -0.64641 | -1.0074 | -1.2114 | 2.9001 | [7 empty] | 0.80375 | 111 | 0.50242 | 138 | 0.88483 | 118 | 11.8841 | 3
| 1.0734 | -0.84866 | 0.043423 | 1.8744 | [7 empty] | 0.87429 | 107 | 0.57306 | 127 | 0.96724 | 105 | 12.0855 | 3
| -1.0434 | 0.18357 | -2.1608 | 5.7915 | [7 empty] | 0.94483 | 89 | 0.6437 | 119 | 1.0496 | 114 | 12.2869 | 3
| 0.20858 | -0.33601 | -1.3134 | 1.8815 | [7 empty] | 1.0154 | 90 | 0.71434 | 99 | 1.132 | 113 | 12.4883 | 3
| 1.2727 | -1.6458 | 1.1723 | 5.7028 | [7 empty] | 1.0859 | 82 | 0.78498 | 113 | 1.2144 | 75 | 12.6897 | 3
| -0.32111 | -1.7586 | -1.5361 | 5.5555 | [7 empty] | 1.1564 | 83 | 0.85561 | 103 | 1.2968 | 73 | 12.8911 | 1
| -1.0115 | 0.233 | 0.66021 | 1.5134 | [7 empty] | 1.227 | 65 | 0.92625 | 99 | 1.3792 | 65 | 13.0926 | 2
| -2.5577 | -0.32949 | -0.59427 | 7.0033 | [7 empty] | 1.2975 | 56 | 0.99689 | 84 | 1.4617 | 55 | 13.294 | 1
| 0.016048 | -1.8514 | -0.26254 | 3.4969 | [7 empty] | 1.3681 | 58 | 1.0675 | 77 | 1.5441 | 38 | 13.4954 | 4
| 0.99939 | -1.0952 | -0.27133 | 2.2719 | [7 empty] | 1.4386 | 53 | 1.1382 | 93 | 1.6265 | 44 | 13.6968 | 3
| 0.90164 | 0.69631 | -0.021182 | 1.2983 | [7 empty] | 1.5091 | 47 | 1.2088 | 72 | 1.7089 | 45 | 13.8982 | 1
| -1.1598 | -0.66371 | 0.8788 | 2.558 | [7 empty] | 1.5797 | 46 | 1.2794 | 52 | 1.7913 | 41 | 14.0996 | 0
| -1.9261 | -0.88167 | 0.72896 | 5.0184 | [7 empty] | 1.6502 | 36 | 1.3501 | 57 | 1.8737 | 35 | 14.3011 | 2
| -0.45867 | 0.28599 | -0.25596 | 0.35768 | [7 empty] | 1.7207 | 39 | 1.4207 | 52 | 1.9561 | 29 | 14.5025 | 0
| -0.14581 | -1.1728 | -1.3574 | 3.2393 | [7 empty] | 1.7913 | 29 | 1.4914 | 52 | 2.0385 | 25 | 14.7039 | 1
| -1.633 | 0.55908 | 0.33478 | 3.0912 | [7 empty] | 1.8618 | 21 | 1.562 | 39 | 2.1209 | 19 | 14.9053 | 1
| 0.37695 | -1.9347 | -0.83733 | 4.5862 | [7 empty] | 1.9324 | 16 | 1.6326 | 48 | 2.2033 | 18 | 15.1067 | 2
| 2.5399 | -0.46218 | 1.023 | 7.7112 | [7 empty] | 2.0029 | 29 | 1.7033 | 34 | 2.2857 | 10 | 15.3082 | 1
| 0.49621 | 0.63464 | 0.55441 | 0.95637 | [7 empty] | 2.0734 | 16 | 1.7739 | 43 | 2.3681 | 8 | 15.5096 | 0
| -0.50427 | -0.15347 | 0.35294 | 0.40241 | [7 empty] | 2.144 | 16 | 1.8445 | 32 | 2.4505 | 16 | 15.711 | 2
| 0.11159 | -0.8584 | 0.40558 | 0.9138 | [7 empty] | 2.2145 | 12 | 1.9152 | 25 | 2.5329 | 7 | 15.9124 | 0
| 0.18621 | -1.9841 | -1.5771 | 6.4583 | [7 empty] | 2.285 | 12 | 1.9858 | 26 | 2.6153 | 9 | 16.1138 | 0
| 0.17318 | -0.4766 | 1.6872 | 3.1039 | [7 empty] | 2.3556 | 5 | 2.0565 | 14 | 2.6977 | 4 | 16.3152 | 0
| 0.53643 | -0.28636 | -2.0631 | 4.6263 | [7 empty] | 2.4261 | 7 | 2.1271 | 24 | 2.7801 | 2 | 16.5167 | 0
| 1.6794 | -0.60499 | -0.24781 | 3.2477 | [7 empty] | 2.4967 | 10 | 2.1977 | 17 | 2.8625 | 5 | 16.7181 | 1
| 0.044446 | -0.70639 | -0.17405 | 0.53125 | [7 empty] | 2.5672 | 5 | 2.2684 | 16 | 2.9449 | 3 | 16.9195 | 0
| 2.1011 | -1.0376 | -0.32677 | 5.598 | [7 empty] | 2.6377 | 4 | 2.339 | 5 | 3.0273 | 1 | 17.1209 | 0
| 1.1922 | -0.15476 | -0.71293 | 1.9536 | [7 empty] | 2.7083 | 6 | 2.4097 | 12 | 3.1097 | 0 | 17.3223 | 1
| 1.1295 | 1.1571 | 0.51263 | 2.8775 | [7 empty] | 2.7788 | 1 | 2.4803 | 8 | 3.1921 | 1 | 17.5237 | 0
| -1.9322 | -0.55901 | -0.41237 | 4.2158 | [7 empty] | 2.8493 | 0 | 2.5509 | 2 | 3.2745 | 2 | 17.7252 | 0
| -0.092672 | 0.78128 | 0.58131 | 0.95692 | [7 empty] | 2.9199 | 3 | 2.6216 | 5 | 3.3569 | 0 | 17.9266 | 0
| 0.29593 | 0.41682 | 0.77085 | 0.85552 | [7 empty] | 2.9904 | 2 | 2.6922 | 2 | 3.4393 | 0 | 18.128 | 0
| 0.66236 | -1.1022 | -0.72668 | 2.1815 | [7 empty] | 3.061 | 2 | 2.7628 | 5 | 3.5217 | 1 | 18.3294 | 0
| -1.0092 | -0.85725 | -0.19694 | 1.7921 | [7 empty] | 3.1315 | 2 | 2.8335 | 3 | 3.6041 | 0 | 18.5308 | 1
| 0.89031 | 0.53343 | -0.73916 | 1.6235 | [7 empty] | 3.202 | 1 | 2.9041 | 1 | 3.6865 | 0 | 18.7323 | 0
| 0.23446 | 0.20987 | -2.2496 | 5.1597 | [7 empty] | 3.2726 | 1 | 2.9748 | 4 | 3.7689 | 0 | 18.9337 | 0
| -0.44014 | -0.42894 | -1.0184 | 1.4148 | [7 empty] | 3.3431 | 0 | 3.0454 | 5 | 3.8513 | 0 | 19.1351 | 0
| 0.79884 | -1.2155 | -0.87969 | 2.8896 | [7 empty] | 3.4136 | 2 | 3.116 | 3 | 3.9337 | 0 | 19.3365 | 0
| -0.93459 | 0.11552 | 0.20244 | 0.92779 | [7 empty] | 3.4842 | 0 | 3.1867 | 0 | 4.0161 | 0 | 19.5379 | 1
| -0.44646 | -0.084655 | 0.093225 | 0.21518 | [7 empty] | 3.5547 | 0 | 3.2573 | 1 | 4.0985 | 0 | 19.7393 | 0
| -0.14824 | 0.50005 | 1.6804 | 3.0958 | [7 empty] | 3.6253 | 0 | 3.328 | 1 | 4.1809 | 0 | 19.9408 | 0
| 1.4207 | 0.28533 | 1.2667 | 3.7042 | [7 empty] | 3.6958 | 0 | 3.3986 | 2 | 4.2633 | 1 | 20.1422 | 1
| 1.7642 | -1.5586 | 0.39623 | 5.6985 | [8 empty] | 1 |  | 1 |  | 0 |  | 0
| 2.0747 | 0.61166 | 0.06794 | 4.6831
| 1.6426 | -0.76422 | -1.3859 | 5.2027
| 1.1927 | -1.2545 | 1.9585 | 6.832
| -0.71281 | -0.05974 | 1.0841 | 1.687
| -0.10043 | -0.96927 | 0.71098 | 1.4551
| -1.6664 | 0.89564 | 2.0701 | 7.8642
| -1.3696 | 0.52629 | 0.40173 | 2.314
| -0.96195 | -2.3015 | -0.54807 | 6.5227
| -0.35158 | -0.075432 | 0.24857 | 0.19109
| -0.52302 | -0.40098 | -0.14138 | 0.45432
| 1.3391 | -0.41178 | -1.7446 | 5.0063
| -2.1531 | -0.027975 | -1.0683 | 5.7778
| -0.87959 | -1.097 | 0.38448 | 2.1249
| -0.48993 | 0.30455 | -0.88525 | 1.1164
| -1.6328 | 0.029925 | 0.87546 | 3.4335
| 0.29879 | 0.15159 | 0.68056 | 0.57541
| -0.85028 | 0.80487 | -0.21212 | 1.4158
| -0.91388 | 0.039483 | 0.58207 | 1.1755
| 1.1889 | 0.97174 | -0.82723 | 3.042
| 0.1741 | 0.38905 | -0.029324 | 0.18253
| -0.1008 | -0.61673 | 0.52504 | 0.66619
| -0.39976 | -0.90966 | 0.44241 | 1.183
| -0.26704 | 0.91754 | -0.085639 | 0.92052
| 0.08902 | -0.31741 | 1.6371 | 2.7888
| 0.57708 | 0.015647 | -1.0297 | 1.3935
| 0.036619 | -1.0244 | -1.3286 | 2.8157
| -0.6961 | -1.3131 | 0.91578 | 3.0473
| 0.49779 | -0.23348 | 1.096 | 1.5035
| -1.9355 | 0.7788 | 2.4971 | 10.5884
| 0.37557 | -0.92619 | 1.5826 | 3.5035
| 0.7713 | -0.4769 | 0.47665 | 1.0495
| 0.75777 | 1.0847 | 0.43907 | 1.9435
| -1.1474 | -1.6768 | -0.40933 | 4.2958
| 0.77637 | 1.509 | -0.79726 | 3.5156
| -0.35529 | 1.5258 | -1.3543 | 4.2882
| 0.95064 | -1.2992 | -1.1449 | 3.9024
| -0.22757 | -1.5977 | -0.66801 | 3.0506
| 0.83923 | -0.84222 | 0.015094 | 1.4139
| -1.2797 | -0.7764 | -1.8661 | 5.7227
| -0.34433 | 0.51776 | -0.080628 | 0.39314
| 0.33691 | 1.5252 | -0.90055 | 3.2508
| -1.1079 | -0.92 | -0.59648 | 2.4296
| -0.25893 | 0.29665 | -0.25897 | 0.22212
| -1.7377 | 1.1174 | -0.82816 | 4.9541
| -1.1439 | 1.6737 | -0.59238 | 4.4608
| -0.85516 | -0.65327 | -1.9226 | 4.8545
| -0.18429 | 0.7273 | -0.39393 | 0.71811
| 0.34913 | -0.36653 | -0.90996 | 1.0843
| 0.40084 | -0.57461 | 2.4743 | 6.6128
| -1.4871 | 1.03 | -0.01046 | 3.2726
| 0.98513 | -1.4791 | -0.30204 | 3.2494
| -0.60095 | -0.59797 | 1.1927 | 2.1414
| -0.21326 | -1.2007 | -0.10089 | 1.4973
| -1.0703 | 0.83519 | 0.65147 | 2.2676
| -0.66945 | -2.0086 | -0.84281 | 5.1931
| -0.8944 | -0.1429 | 0.5706 | 1.1459
| -2.1436 | -1.1572 | 0.89866 | 6.7417
| 1.3979 | -1.7149 | 0.12173 | 4.9099
| 1.7921 | 0.84059 | -0.48905 | 4.1573
| 0.78065 | 0.9599 | -2.2304 | 6.5056
| -1.1968 | -0.33348 | 0.16969 | 1.5722
| -0.42043 | 2.3603 | -0.055685 | 5.7506
| -0.28992 | -0.7426 | 0.63712 | 1.0414
| 0.15395 | 1.1125 | -0.53479 | 1.5473
| -0.40913 | 0.056799 | 0.097647 | 0.18015
| 0.13097 | 1.0596 | 0.67636 | 1.5973
| -0.31882 | 1.8663 | 0.55745 | 3.8956
| 0.58795 | 0.42424 | -0.40214 | 0.68738
| 0.13655 | 0.30697 | -1.7483 | 3.1694
| 0.75306 | 0.060184 | 0.50781 | 0.82859
| 0.23706 | 0.19977 | -1.3311 | 1.868
| 0.7485 | -2.0611 | -0.17961 | 4.8405
| 0.80708 | 0.73912 | -0.11213 | 1.2102
| 0.031949 | 0.22713 | -1.7258 | 3.0312
| -0.23578 | 2.0581 | 1.1328 | 5.5747
| -2.1116 | -0.2329 | -0.09733 | 4.5226
| -1.2837 | -0.025391 | 0.27244 | 1.7227
| 0.39782 | -1.428 | 0.96032 | 3.1197
| 0.3571 | 2.8426 | -0.18152 | 8.241
| -0.4444 | 0.74093 | 0.51427 | 1.0109
| 0.12896 | 0.76655 | -0.68779 | 1.0773
| -0.063471 | -0.90675 | -0.69546 | 1.3099
| -1.0788 | 0.87816 | -0.94184 | 2.822
| -1.0191 | -0.38077 | -0.00084787 | 1.1835
| 0.46154 | 0.64595 | 1.2619 | 2.2225
| 0.49104 | 0.45661 | -0.87231 | 1.2105
| 1.9852 | 0.70001 | 1.5675 | 6.8883
| 0.081209 | -1.3382 | -0.2714 | 1.8711
| -0.30321 | -0.84193 | 0.72869 | 1.3318
| 0.0055282 | 1.2292 | 2.5351 | 7.9375
| 1.6823 | -2.0912 | 0.16207 | 7.2296
| -0.40773 | 0.70395 | 0.75944 | 1.2385
| 0.30912 | 0.75292 | 0.61985 | 1.0467
| -0.59342 | -0.054054 | -1.0212 | 1.3979
| -0.61456 | 0.67521 | -0.20534 | 0.87575
| -1.3826 | 0.14592 | -0.63161 | 2.3319
| 0.81389 | 1.8271 | -0.81205 | 4.6601
| 0.12365 | -0.13881 | -0.57198 | 0.36172
| 0.11517 | -0.5741 | -1.5323 | 2.6907
| 1.3133 | 0.59846 | 0.46977 | 2.3035
| 0.76602 | 0.10328 | 0.91736 | 1.439
| -1.0716 | -1.4257 | -1.4797 | 5.3705
| -0.68616 | 1.2085 | -2.4411 | 7.8901
| 0.83349 | 0.7018 | -0.18721 | 1.2223
| 0.69711 | 0.53166 | -0.16434 | 0.79564
| 0.28177 | -0.46407 | 0.34799 | 0.41585
| -1.511 | -1.5822 | 0.78048 | 5.3959
| -0.57287 | -1.0428 | -0.86409 | 2.1623
| -1.008 | 0.30479 | 0.022801 | 1.1095
| -0.18723 | -0.081023 | -0.59354 | 0.39391
| -0.77902 | -0.45037 | -0.85889 | 1.5474
| -0.58837 | -0.05098 | 2.0373 | 4.4993
| -1.388 | -0.26528 | 0.86348 | 2.7425
| 0.58798 | -0.88877 | 1.5921 | 3.6706
| 0.35438 | -0.42599 | -0.34906 | 0.42889
| 0.65207 | -0.82688 | 0.45959 | 1.3201
| 0.8977 | 1.5109 | 1.8908 | 6.6638
| 0.88615 | -0.28199 | -0.96115 | 1.7886
| 0.7317 | 0.039995 | 0.077006 | 0.54292
| -1.63 | 1.0904 | -0.66693 | 4.2905
| -0.50933 | 0.59665 | 0.19054 | 0.65171
| -1.3282 | -0.59662 | -2.1755 | 6.8528
| 0.5798 | 1.5512 | -1.5597 | 5.1751
| -1.7919 | 0.014258 | 1.6427 | 5.9097
| 0.17529 | 0.47514 | -0.035301 | 0.25773
| -0.48147 | 0.26906 | -0.35256 | 0.4285
| 0.39702 | -0.71131 | 1.3248 | 2.4186
| 0.74227 | -3.0402 | -0.65244 | 10.2194
| -0.47899 | -1.1662 | -1.8551 | 5.031
| -0.67894 | -1.999 | -1.3729 | 6.3418
| -0.80667 | 1.8714 | 1.2356 | 5.6795
| 0.87882 | -0.055104 | -3.0408 | 10.0216
| -0.010548 | 0.10358 | 1.0142 | 1.0395
| -0.38791 | -0.73394 | -1.4608 | 2.8231
| 2.6051 | 0.15964 | 0.57271 | 7.1398
| 0.21144 | 0.40756 | -0.65867 | 0.64466
| 0.98571 | -0.90342 | 0.41111 | 1.9568
| -0.98174 | -1.0867 | 0.6177 | 2.5262
| -1.0785 | 1.808 | -1.0022 | 5.4364
| 0.010175 | -0.35309 | 2.1339 | 4.6783
| -0.60322 | 0.25812 | 0.76599 | 1.0172
| 0.60974 | 0.33332 | 0.83699 | 1.1834
| 0.0037773 | 0.11662 | -1.2459 | 1.5658
| 1.096 | -0.45148 | 1.0356 | 2.4776
| -0.37793 | -1.0343 | 1.454 | 3.3268
| -0.4028 | 3.0266 | -0.69909 | 9.811
| -0.57394 | -2.9335 | 0.19375 | 8.9722
| -1.4912 | -1.9255 | 1.5225 | 8.2494
| -0.5644 | -0.74401 | -0.6127 | 1.2475
| 1.942 | 0.48009 | 0.35775 | 4.1299
| 0.5362 | -0.77603 | 1.2075 | 2.3478
| -0.52011 | 0.51593 | -0.3885 | 0.68763
| -0.48708 | 1.3029 | 1.6155 | 4.5446
| 0.11451 | -1.0707 | 0.52573 | 1.4359
| 1.3803 | 0.60402 | 0.092325 | 2.2785
| -0.068696 | 2.3719 | 0.44361 | 5.8273
| -0.48414 | 0.87396 | 1.1506 | 2.322
| -0.32353 | -0.45505 | -1.5631 | 2.755
| -0.14675 | -0.2378 | 0.27702 | 0.15483
| -0.92924 | -0.39346 | 1.9049 | 4.6468
| 0.46408 | -0.53136 | -1.2862 | 2.152
| -0.40429 | -1.048 | -1.754 | 4.3381
| -0.099622 | 0.72516 | -1.1385 | 1.8319
| -0.5992 | 0.20786 | 0.46124 | 0.61499
| -0.50553 | -0.16162 | -0.32801 | 0.38928
| 0.023757 | 0.032473 | 0.84463 | 0.71501
| 0.41547 | 0.065712 | -0.71933 | 0.69437
| 0.50956 | 0.70929 | 0.27056 | 0.83594
| -0.55409 | -0.53645 | 1.3477 | 2.4111
| 0.48176 | -0.61717 | -1.6658 | 3.3879
| -0.093336 | 1.6926 | 0.35098 | 2.9969
| 0.86654 | -0.26798 | 1.0653 | 1.9575
| -0.20697 | 1.2053 | 0.78729 | 2.1153
| -0.11452 | -0.7892 | 0.51951 | 0.90584
| 1.152 | 0.8626 | 1.994 | 6.0471
| -2.1623 | 2.6167 | -0.72279 | 12.0451
| -0.34741 | 1.8148 | -0.3498 | 3.5366
| 0.8221 | -0.81109 | 1.34 | 3.1293
| 2.2265 | -0.57559 | -1.6664 | 8.0655
| -1.2257 | 1.1565 | -0.097554 | 2.8494
| 2.1372 | -1.2692 | 0.96276 | 7.1055
| 1.1463 | -0.57335 | -1.1191 | 2.8952
| -1.2305 | -0.58373 | -0.24763 | 1.9163
| -0.094086 | -0.093859 | 0.25105 | 0.080688
| 2.3855 | 2.2133 | -1.3747 | 12.4793
| 1.7235 | 0.81499 | -0.88018 | 4.4095
| 0.25975 | 1.9262 | 0.63816 | 4.1848
| -0.42556 | -0.87367 | 1.0435 | 2.0332
| -0.13604 | 0.26129 | 1.0045 | 1.0958
| -1.9068 | -0.0031394 | 0.66816 | 4.0822
| 0.12416 | -1.1316 | -0.80503 | 1.9439
| 0.83887 | 0.56272 | 0.52748 | 1.2986
| -0.65433 | 0.51687 | 0.52572 | 0.97168
| 1.465 | -0.076584 | 2.341 | 7.6323
| 1.1394 | 0.15617 | -0.42318 | 1.5017
| 1.0435 | -0.19855 | 1.7662 | 4.2477
| -0.80287 | 0.5295 | 0.65194 | 1.35
| -1.3747 | 1.265 | -0.20786 | 3.5332
| 0.87096 | 1.6549 | -0.30335 | 3.5891
| 0.49335 | 1.451 | 0.34895 | 2.4705
| -0.70142 | -0.42854 | 1.5525 | 3.086
| -2.1904 | -0.23983 | -1.1697 | 6.2235
| -0.88479 | 1.0942 | -0.66673 | 2.4247
| -0.71439 | -0.44457 | 0.29102 | 0.79269
| 0.33814 | -0.15022 | -0.15901 | 0.16219
| 1.7504 | -1.2059 | -0.64422 | 4.933
| -0.73301 | -0.50174 | -1.3793 | 2.6916
| -0.075099 | 0.1718 | -0.1327 | 0.052764
| 1.2066 | 0.30976 | 0.55368 | 1.8584
| 0.12216 | -0.30918 | -0.61458 | 0.48823
| 1.3692 | 0.13117 | 0.12355 | 1.9071
| 1.268 | -0.47582 | -0.2841 | 1.9149
| -0.73077 | -1.4686 | 1.0114 | 3.7138
| 0.1249 | 0.25169 | 0.94145 | 0.96527
| -1.5233 | -0.73811 | -1.5612 | 5.3026
| 0.86424 | -0.93467 | 0.34793 | 1.7416
| 0.021175 | -0.01187 | 0.59598 | 0.35578
| 1.0719 | 0.4909 | 0.13437 | 1.4079
| -0.20226 | 1.1538 | 0.01801 | 1.3724
| 0.016246 | -0.19716 | -1.1291 | 1.314
| 0.18185 | 1.1905 | 0.066855 | 1.4547
| 0.41751 | -0.33382 | 0.76023 | 0.8637
| 0.94833 | 0.4421 | -2.2356 | 6.0926
| -1.1197 | 0.22621 | 0.57054 | 1.6303
| -0.48434 | -1.3375 | 0.043318 | 2.0255
| -0.055778 | 2.2847 | 0.97885 | 6.1812
| -0.43451 | 0.71209 | -0.97103 | 1.6388
| 0.1576 | 0.25798 | 1.0567 | 1.2081
| -0.6969 | -0.61679 | 0.78931 | 1.4891
| -0.749 | 1.0998 | -0.79302 | 2.3995
| 0.77371 | -1.3275 | -1.1999 | 3.8007
| 1.601 | 1.0552 | -0.6658 | 4.1199
| 1.4638 | 0.36116 | 0.37793 | 2.416
| 0.68666 | 0.32514 | 0.065603 | 0.58153
| 1.8183 | 0.15991 | -0.47249 | 3.5551
| -0.89496 | 0.056093 | 0.1252 | 0.81977
| -0.84982 | -0.71425 | 0.38744 | 1.3825
| -0.96221 | -0.12548 | 0.091591 | 0.94997
| 1.2356 | 0.6643 | 1.0468 | 3.0637
| 2.0971 | 0.81606 | -0.52926 | 5.344
| -0.77646 | -1.756 | -0.92826 | 4.5481
| -0.22184 | 0.72236 | -0.50288 | 0.82391
| 0.83333 | 0.33457 | -0.79314 | 1.4355
| 0.67828 | -0.47077 | 0.28659 | 0.76381
| 0.27821 | -0.30291 | 0.34851 | 0.29062
| 1.2282 | -0.77234 | -1.2009 | 3.5472
| 0.68882 | -1.3448 | 0.81141 | 2.9415
| 0.96157 | 0.4835 | 1.2808 | 2.7988
| -1.7275 | 0.83271 | -0.25554 | 3.7428
| 1.732 | 0.87712 | -0.58898 | 4.1161
| -0.7351 | 1.4678 | -2.0573 | 6.9274
| -1.5697 | -1.0744 | 1.054 | 4.729
| -0.021157 | 0.3854 | -1.1751 | 1.5298
| -0.64369 | -0.55019 | -1.9725 | 4.6078
| 1.2151 | -1.9305 | 0.046259 | 5.2053
| -1.1575 | -0.73324 | -0.22468 | 1.928
| -1.0275 | 0.13986 | -0.84153 | 1.7835
| -0.45854 | 3.3986 | 0.40099 | 11.9215
| -1.0667 | -0.65846 | 0.25282 | 1.6353
| -1.0517 | -1.3474 | -0.72012 | 3.4401
| -0.60684 | 0.73769 | -0.51745 | 1.1802
| -0.058464 | 0.57425 | -0.42101 | 0.51043
| -1.0508 | -0.73847 | -1.0527 | 2.7577
| 2.0826 | 0.28883 | 1.2309 | 5.9357
| 1.3536 | 0.40626 | -0.6978 | 2.4841
| -0.49562 | -0.22723 | -0.79396 | 0.92766
| -0.55289 | 0.76064 | -1.2079 | 2.3434
| 0.92503 | 1.1166 | -0.63613 | 2.5071
| 0.96316 | 0.14714 | 0.73913 | 1.4956
| -0.015385 | 0.45889 | -0.99236 | 1.1956
| 0.70295 | 0.28357 | -0.24639 | 0.63526
| -1.6189 | -0.92293 | 0.43566 | 3.6623
| 0.21456 | -0.0035105 | 1.6188 | 2.6666
| 0.037346 | 0.86285 | -1.0582 | 1.8657
| 2.0469 | 0.64097 | 0.16019 | 4.6263
| 0.1883 | -0.45357 | 0.76199 | 0.82182
| 0.98039 | 1.458 | 0.13521 | 3.1053
| -0.45952 | -1.2956 | -1.1704 | 3.2596
| 0.33564 | 0.26359 | 0.23889 | 0.2392
| 1.7285 | -0.68197 | -0.46978 | 3.6735
| -0.29097 | 0.65461 | 0.61932 | 0.89674
| 0.093051 | 0.86656 | -0.19468 | 0.79749
| -0.33685 | 1.3595 | -0.53202 | 2.2449
| -0.6135 | 1.0358 | -2.16 | 6.1149
| 0.043075 | 0.19156 | 0.63368 | 0.4401
| -0.12313 | 0.20221 | 1.6605 | 2.8131
| 1.1982 | 0.029266 | -0.1329 | 1.4542
| 1.0764 | 0.46395 | -0.027867 | 1.3746
| -0.47059 | -0.79785 | -0.68197 | 1.3231
| -1.0806 | -0.31834 | -0.71912 | 1.7862
| -0.45782 | 1.7724 | 0.50056 | 3.6017
| 0.44744 | -0.23762 | 1.6844 | 3.0939
| -1.4795 | -1.2039 | 0.7809 | 4.248
| 0.40107 | 0.19202 | -0.98711 | 1.1721
| 0.41776 | -0.90886 | -0.85894 | 1.7383
| 0.41256 | -0.24182 | 1.4975 | 2.4713
| -0.20361 | -0.087256 | -0.38014 | 0.19358
| -0.58141 | -1.4209 | 0.75238 | 2.9231
| 0.49607 | 3.1917 | -0.61969 | 10.8168
| -1.8526 | 0.3299 | -1.1858 | 4.9471
| -0.14741 | 1.2529 | -0.29171 | 1.6765
| 2.2648 | -1.1624 | 2.0255 | 10.5836
| -1.3518 | -0.39784 | -1.2015 | 3.4292
| 0.083645 | -0.19231 | 0.62936 | 0.44008
| -1.3791 | -0.88173 | -1.9295 | 6.4024
| 0.55246 | -0.3116 | 1.3582 | 2.2469
| -1.1622 | -0.50729 | 0.5214 | 1.88
| -0.54615 | 0.51355 | 1.1233 | 1.8238
| 0.32322 | 1.0651 | 0.66165 | 1.6768
| 0.76286 | -0.50077 | -0.37234 | 0.97136
| 1.1041 | -0.34187 | 0.22603 | 1.3869
| 1.1398 | -0.27638 | -0.87706 | 2.1447
| 1.3017 | 0.11236 | -0.53555 | 1.9939
| 0.38001 | 1.5254 | 1.4884 | 4.6866
| -1.4803 | 0.22074 | 1.2951 | 3.9173
| -0.56446 | 0.80095 | 0.39003 | 1.1123
| 0.45026 | 0.10064 | 0.83941 | 0.91747
| 0.41256 | -1.0106 | -1.8727 | 4.6987
| 0.07206 | -1.4539 | 0.34175 | 2.2357
| -0.72607 | -0.17469 | 0.91701 | 1.3986
| 0.54649 | 0.63494 | 0.33023 | 0.81086
| -0.98748 | -0.83665 | 0.674 | 2.1294
| 1.8364 | 0.098867 | -1.233 | 4.9025
| -0.48239 | -0.3614 | 0.26216 | 0.43203
| 0.59887 | -0.488 | -0.23008 | 0.64973
| -0.5309 | -1.4006 | 1.4346 | 4.3017
| 1.39 | 1.3154 | -1.4174 | 5.6714
| 0.34137 | -1.0455 | -0.12385 | 1.225
| 1.1877 | 1.1269 | 1.5819 | 5.1829
| 0.93746 | -1.2113 | -0.42203 | 2.5242
| -0.81205 | -1.6822 | 0.25664 | 3.5549
| -0.35693 | 1.0163 | 0.59438 | 1.5136
| -1.0619 | 0.72046 | 0.13748 | 1.6656
| -0.090974 | -0.050717 | 1.3933 | 1.952
| 0.6497 | -1.7084 | 0.34966 | 3.4631
| -0.54072 | -0.37353 | 0.29568 | 0.51932
| -2.3092 | -0.23103 | -0.40559 | 5.5502
| 1.0368 | -0.076745 | -1.3499 | 2.9031
| -1.7734 | 0.17028 | -0.85341 | 3.9022
| 0.98234 | -1.5236 | -0.45025 | 3.489
| -0.41867 | -1.3347 | 1.8819 | 5.4982
| 0.61155 | 0.0093598 | 0.4071 | 0.53981
| 0.67005 | 0.19125 | -1.2646 | 2.0848
| 0.62055 | -1.0052 | 0.27146 | 1.4692
| -1.0982 | -0.058528 | -1.2346 | 2.7337
| 0.62165 | -0.14422 | -0.0019997 | 0.40725
| -0.022683 | 0.47566 | -1.9947 | 4.2057
| -0.65854 | -2.4604 | 1.2662 | 8.0904
| -0.2335 | -0.32863 | 0.77931 | 0.76985
| 0.33447 | 0.024088 | 1.3876 | 2.038
| -1.4757 | -1.0329 | -0.052688 | 3.2473
| 1.4532 | -0.8623 | 0.0018177 | 2.8553
| 0.39124 | 1.0251 | 0.4188 | 1.3793
| 1.2336 | 1.094 | 0.047321 | 2.7211
| -1.292 | -1.0068 | 1.3983 | 4.6379
| -0.64371 | 0.50829 | 0.81752 | 1.3411
| -2.2338 | 0.38607 | 1.0672 | 6.2778
| 1.2098 | -0.07141 | 0.9455 | 2.3628
| 0.56455 | 0.48194 | 1.8748 | 4.066
| 1.0797 | -0.1773 | -0.55388 | 1.5039
| 0.84545 | 0.019264 | 0.085856 | 0.72253
| 0.8431 | -0.52566 | 1.4173 | 2.9958
| 1.2224 | 0.50048 | -0.69738 | 2.231
| -0.096642 | 0.96497 | 0.0032628 | 0.94052
| -0.4526 | -0.81127 | -0.99454 | 1.8521
| 0.60382 | 1.098 | 2.1633 | 6.2499
| 0.45068 | -1.1341 | 1.5213 | 3.8037
| -0.82182 | -0.45378 | 0.68309 | 1.3479
| -2.3768 | 1.8122 | 0.17083 | 8.9626
| -0.4277 | 0.18915 | 0.062535 | 0.22262
| 0.57755 | 0.61633 | 2.8862 | 9.0435
| -0.18722 | -0.27814 | -0.718 | 0.62793
| -0.65282 | -1.209 | -0.18917 | 1.9235
| 0.72697 | -0.599 | 1.0897 | 2.0747
| -0.53689 | -0.40721 | -0.41614 | 0.62725
| 1.0805 | 0.31014 | 0.19399 | 1.3014
| 0.80139 | -0.27215 | -2.216 | 5.6268
| 0.31648 | -0.25172 | -0.4355 | 0.35319
| 1.4009 | -0.21699 | 0.51577 | 2.2755
| -0.35756 | 1.764 | 0.51569 | 3.5056
| 0.30206 | -0.22698 | 0.26344 | 0.21216
| 0.93832 | 0.45535 | 0.63034 | 1.4851
| 0.2789 | -1.1678 | 1.061 | 2.5674
| -0.80631 | -0.19431 | -0.52042 | 0.95872
| 2.6769 | -0.31979 | -0.14616 | 7.2894
| 0.083772 | -0.55039 | -0.34303 | 0.42762
| -0.063768 | -0.15367 | -0.63648 | 0.43279
| -0.30422 | -0.46147 | 1.1696 | 1.6735
| 0.15134 | -0.46699 | 0.35143 | 0.36448
| -1.9347 | -0.80475 | -1.0436 | 5.4795
| -0.47229 | 0.018542 | 0.44456 | 0.42104
| -0.29115 | 0.77615 | 0.59096 | 1.0364
| -1.1578 | 0.54243 | 0.70169 | 2.127
| 0.16935 | -0.22559 | -0.86433 | 0.82665
| -2.028 | -0.4268 | -1.6088 | 6.8833
| 0.47171 | -0.00066449 | 1.9747 | 4.1219
| 0.21235 | -0.5537 | -1.365 | 2.2148
| -1.2614 | -1.3761 | 0.1777 | 3.5165
| 1.0471 | 0.11491 | -0.25646 | 1.1753
| 1.8304 | -0.51202 | -0.37072 | 3.7498
| -0.14441 | -0.70001 | 2.1559 | 5.1586
| -1.7694 | -0.018462 | -0.52071 | 3.4021
| -0.19171 | 2.2067 | -1.0499 | 6.0087
| 0.79169 | 1.4087 | -1.2166 | 4.0911
| -1.4271 | -0.037415 | 1.6502 | 4.7611
| 0.21525 | -0.23695 | -0.0084652 | 0.10255
| 0.91341 | -1.2563 | 0.27766 | 2.4897
| 0.94742 | -0.29556 | 0.41775 | 1.1595
| 0.50917 | 0.80841 | -0.79841 | 1.5502
| -0.6195 | 0.27507 | -2.1971 | 5.2866
| -0.81441 | -0.20907 | -0.35669 | 0.83421
| -1.4415 | -0.86133 | 0.40341 | 2.9824
| -0.76053 | -0.72863 | -0.76988 | 1.702
| -0.82194 | -0.88522 | 0.5657 | 1.7792
| -0.053699 | -2.4336 | -0.27723 | 6.0023
| 0.47666 | 0.12126 | -1.3308 | 2.0129
| -0.70037 | -0.9866 | 1.7345 | 4.4723
| -0.31527 | 0.14427 | 1.6377 | 2.8023
| 1.9829 | 0.56984 | 0.092541 | 4.2653
| -0.30423 | -0.75995 | -0.57951 | 1.0059
| 1.5515 | -0.34359 | -0.18578 | 2.5596
| -0.073209 | 0.58579 | 0.95835 | 1.2669
| 0.54332 | -0.79436 | -0.73755 | 1.4702
| 0.58788 | 0.2386 | 0.54308 | 0.69747
| -2.4206 | 0.37819 | 0.078704 | 6.0086
| 0.65428 | -1.8184 | -0.75642 | 4.3067
| 0.10081 | -1.0732 | 0.11469 | 1.1751
| -1.1412 | -2.7115 | 1.2317 | 10.1719
| 1.417 | 1.3471 | 0.41346 | 3.9934
| 0.3757 | -0.5875 | -0.21401 | 0.53212
| 1.2584 | 0.076916 | -1.5442 | 3.974
| -1.7273 | 0.95704 | -0.016303 | 3.8999
| -1.5782 | 0.45381 | -1.2951 | 4.374
| 0.35788 | 1.1169 | -1.6683 | 4.1587
| -1.0268 | -1.8791 | -1.7084 | 7.504
| 0.57324 | 0.33832 | 1.048 | 1.5414
| -1.14 | 2.068 | -1.0209 | 6.6184
| -0.51418 | -0.7243 | 0.55561 | 1.0977
| 1.6985 | 0.21342 | -0.86689 | 3.682
| -1.3733 | -0.58568 | 0.98436 | 3.1981
| -0.71898 | 0.74129 | 0.95132 | 1.9714
| 0.086734 | -0.91631 | 0.80183 | 1.4901
| -0.9997 | -1.4703 | -1.1384 | 4.4573
| 1.4161 | 1.0391 | 0.31638 | 3.1851
| 0.91505 | 0.69449 | -1.1007 | 2.5312
| 1.3957 | 2.009 | 0.36925 | 6.1203
| 1.5713 | -0.67952 | -1.0421 | 4.0167
| -1.6868 | 0.83827 | 0.68223 | 4.0135
| -0.80215 | -0.25594 | -0.65654 | 1.14
| 0.039983 | 0.52422 | 0.61625 | 0.65618
| -0.50817 | 0.039917 | 0.86812 | 1.0135
| 0.72735 | -0.61666 | 1.4138 | 2.9082
| -0.81744 | 0.54919 | 0.64359 | 1.384
| 0.27591 | 0.027956 | 0.56241 | 0.39321
| 0.43677 | 0.20793 | 0.9409 | 1.1193
| 0.75681 | -0.057152 | 0.52346 | 0.85003
| -0.37902 | 2.2564 | -1.0392 | 6.315
| 0.78316 | 1.1548 | -0.10005 | 1.9569
| -0.88011 | 1.5276 | -1.4771 | 5.29
| 0.3872 | 0.61368 | 0.072756 | 0.53182
| 0.88208 | -1.5121 | 0.5531 | 3.3703
| -0.27633 | -0.024231 | -0.024585 | 0.077551
| 0.19727 | 0.026422 | 0.8546 | 0.76996
| -0.067639 | -0.41395 | -0.72212 | 0.69738
| 0.59584 | -0.32555 | 0.26903 | 0.53339
| -0.78194 | 1.0525 | 0.83178 | 2.4111
| -0.69551 | -2.9383 | -1.5497 | 11.5192
| 2.184 | -0.49178 | -0.29901 | 5.1011
| 0.0047867 | -0.94378 | -0.051083 | 0.89335
| -1.2144 | -0.66858 | -2.4312 | 7.8326
| -0.10563 | 0.30133 | -0.73284 | 0.63901
| 1.5441 | -1.4949 | 0.31459 | 4.7178
| -0.40399 | 1.2984 | -0.36802 | 1.9846
| 0.42602 | -0.061187 | -1.6914 | 3.046
| 0.14439 | 0.41443 | 0.59279 | 0.544
| 0.19831 | 0.81523 | -0.51172 | 0.96577
| -0.2318 | 0.45335 | -1.3814 | 2.1675
| 2.0118 | -1.5504 | -1.2227 | 7.9462
| 1.1903 | -1.1116 | -0.95838 | 3.571
| 0.71321 | -1.6535 | -0.33766 | 3.3568
| 0.9896 | 0.42069 | -0.85694 | 1.8906
| -0.74805 | -0.63478 | 0.09442 | 0.97145
| 0.95738 | -0.34604 | 2.1448 | 5.6365
| 0.4115 | 1.0917 | 0.40442 | 1.5247
| -1.2456 | 0.72018 | 0.56886 | 2.3937
| -0.10003 | -1.4488 | 0.34565 | 2.2285
| 0.97072 | 1.0683 | -2.2834 | 7.2975
| -0.46669 | 1.1129 | -0.3363 | 1.5694
| 0.23755 | -1.4852 | -1.3327 | 4.0383
| 1.0219 | -0.39648 | -0.67694 | 1.6598
| -0.76562 | 2.9789 | -0.074664 | 9.4655
| 1.9474 | 1.0465 | -0.87743 | 5.6574
| -1.775 | 1.1791 | -1.5765 | 7.0261
| 0.069547 | 0.32967 | 0.1193 | 0.12775
| 1.5737 | -0.59536 | 0.57841 | 3.1656
| 0.92124 | 0.55535 | 1.2533 | 2.7278
| -0.15113 | 2.7249 | 1.413 | 9.4441
| -0.522 | 1.6106 | -0.084341 | 2.8735
| 2.3368 | -1.2198 | -0.078322 | 6.9545
| -0.6099 | 0.8053 | -1.2383 | 2.5538
| -2.3856 | -1.1938 | 0.30409 | 7.209
| -0.77596 | 0.75143 | 1.1284 | 2.4401
| 1.0728 | -3.6563 | 1.9852 | 18.4604
| 0.697 | 1.3744 | 0.54546 | 2.6722
| -0.62972 | 1.1328 | 0.2484 | 1.7415
| -0.20973 | 2.0285 | 1.7266 | 7.1398
| -1.4823 | -0.55213 | -0.2067 | 2.5447
| 0.22361 | -0.33103 | 0.0072381 | 0.15963
| -1.2302 | 0.41111 | 0.3606 | 1.8124
| -1.1197 | -1.3904 | 0.79702 | 3.8221
| -2.4557 | 0.24979 | 1.0317 | 7.1574
| 0.46161 | 0.059651 | 0.48211 | 0.44907
| 0.19687 | -2.688 | 1.1757 | 8.6466
| -0.35942 | 0.61248 | 0.072864 | 0.50962
| 0.093055 | 0.31781 | 1.4703 | 2.2716
| 1.4622 | -0.75457 | -0.8302 | 3.3965
| -0.55757 | 0.88899 | -0.30038 | 1.1914
| 0.81247 | 0.25694 | 0.76878 | 1.3172
| 1.1493 | 0.019207 | 0.56424 | 1.6397
| 2.0165 | -0.65834 | -1.4622 | 6.6378
| -1.4582 | -0.082388 | -1.1743 | 3.512
| 1.1839 | -0.83273 | -1.4804 | 4.2868
| 0.047431 | -0.46045 | 0.73678 | 0.75711
| 0.28927 | -1.3163 | 1.1714 | 3.1886
| -0.20456 | -0.61831 | -1.2999 | 2.114
| 0.45798 | 0.13538 | -0.02228 | 0.22857
| 0.4388 | 0.52876 | -0.69159 | 0.95043
| -0.6025 | -1.512 | 0.85292 | 3.3765
| 0.58019 | -1.2411 | 0.08775 | 1.8846
| 0.277 | 0.18303 | -0.25049 | 0.17297
| 0.051805 | 0.57861 | -1.1329 | 1.6209
| 1.2962 | -1.4726 | -0.70436 | 4.3447
| -0.29036 | 0.098371 | 0.15134 | 0.11689
| 0.43764 | -1.3304 | -0.52943 | 2.2417
| -0.54763 | -0.10982 | -0.1999 | 0.35191
| -1.4032 | 0.0072016 | 0.72577 | 2.4958
| -0.64706 | 0.45381 | -0.35278 | 0.74908
| -0.72096 | 0.31771 | 1.2243 | 2.1195
| 0.24433 | -1.9502 | -0.98442 | 4.8321
| 0.62023 | -0.73189 | -1.0827 | 2.0926
| -0.58039 | 1.9641 | -1.1714 | 5.5667
| 1.2778 | 0.52337 | -0.42586 | 2.088
| 0.1118 | -0.60061 | 0.87352 | 1.1363
| 0.48826 | 0.44714 | 2.5969 | 7.1821
| 0.64508 | 0.50089 | 1.1661 | 2.0269
| 2.6599 | 0.75362 | 0.40995 | 7.8113
| -1.7093 | 0.9865 | 0.88145 | 4.6719
| 0.15451 | -0.31786 | -0.3908 | 0.27763
| 1.2439 | -0.65774 | -0.31328 | 2.0782
| -0.087185 | 0.24107 | -1.0051 | 1.0759
| -0.26709 | 1.8249 | -0.70701 | 3.9014
| -0.32128 | 0.59 | -1.5665 | 2.9054
| 0.086371 | -0.3759 | -2.6701 | 7.278
| -0.40174 | 1.2152 | -0.064718 | 1.6422
| 1.5668 | -0.26594 | -0.58931 | 2.8727
| 1.7645 | -1.6829 | 0.71135 | 6.4516
| 1.4569 | -1.2008 | 1.1411 | 4.8665
| 1.5735 | 0.29287 | -0.26886 | 2.6339
| -0.33745 | -0.05733 | 0.58549 | 0.45996
| 1.219 | -1.1151 | -1.3794 | 4.632
| 0.99585 | 1.1056 | 1.111 | 3.4485
| -0.2449 | -0.9963 | 0.82622 | 1.7352
| 1.1019 | -0.24228 | -0.012796 | 1.2731
| -0.28099 | -1.0477 | -0.56255 | 1.493
| 0.70692 | -0.19857 | 0.29259 | 0.62477
| 1.1123 | 0.36865 | -0.22921 | 1.4257
| -0.72751 | 0.58036 | -0.47329 | 1.0901
| -0.21825 | -1.3105 | -0.96946 | 2.7048
| -2.553 | -1.211 | 1.8653 | 11.464
| -0.53557 | 1.6074 | 0.46486 | 3.0868
| 0.37226 | 0.089741 | -0.70464 | 0.64314
| -1.1168 | 0.64834 | 0.27429 | 1.7428
| 0.69896 | 0.245 | 0.31568 | 0.64822
| -0.69228 | 0.8794 | 0.20891 | 1.2962
| 0.36805 | -0.24629 | -1.7895 | 3.3983
| -1.1834 | 0.45207 | -0.8624 | 2.3486
| -0.3157 | -0.47592 | -2.3438 | 5.8195
| -1.0659 | 0.33543 | -0.26429 | 1.3186
| 0.65153 | 1.3723 | -0.65832 | 2.741
| 1.3931 | 2.2111 | 0.098701 | 6.8394
| 0.85819 | -1.4171 | 0.68186 | 3.2095
| -1.1452 | -0.56073 | 0.98632 | 2.5988
| -1.7258 | -0.20318 | 0.29407 | 3.1062
| -0.92858 | -0.37726 | 0.31475 | 1.1037
| -0.44766 | -1.4468 | -0.43655 | 2.4843
| -0.32893 | -0.24197 | 1.1162 | 1.4127
| -0.94491 | 1.6194 | -0.71348 | 4.0243
| 0.44148 | -0.26216 | -0.82899 | 0.95086
| -0.29316 | -0.9476 | 0.14665 | 1.0054
| -2.4741 | -1.4216 | 0.52825 | 8.4214
| 0.20195 | 2.4404 | -1.233 | 7.5164
| 0.084956 | -1.2778 | -1.0841 | 2.8153
| -2.0934 | 1.7139 | 0.82904 | 8.0071
| -0.14982 | -0.4784 | 1.1817 | 1.6478
| 1.7155 | 0.56178 | 0.1512 | 3.2814
| -0.065533 | -0.53621 | 0.076304 | 0.29763
| 0.031712 | -0.76234 | -0.86569 | 1.3316
| 1.0693 | -0.012865 | -0.032981 | 1.1446
| -1.4031 | -1.2731 | 1.0178 | 4.6253
| 1.1623 | 0.47467 | -2.1156 | 6.0518
| -0.81503 | 0.20213 | -0.10044 | 0.71523
| 1.3057 | 1.6772 | -0.60262 | 4.8811
| 1.4308 | -1.839 | 0.98396 | 6.3974
| -1.2132 | 0.93708 | -1.9004 | 5.9617
| -1.2972 | 0.74067 | 0.37572 | 2.3725
| 0.36265 | 0.18162 | 0.069196 | 0.16929
| 1.4115 | 0.88431 | -0.45954 | 2.9856
| -0.18116 | 0.35461 | 0.29065 | 0.24304
| 1.6595 | -1.4772 | 0.30258 | 5.0277
| -1.5823 | 0.13042 | 0.38176 | 2.6665
| 1.1422 | 1.5698 | 0.89068 | 4.5622
| -1.2115 | -0.54462 | 0.46646 | 1.9819
| 0.80878 | 1.1379 | -1.5336 | 4.3008
| -2.0188 | -1.0911 | 1.8702 | 8.7637
| 0.72026 | -1.5663 | 0.38416 | 3.1196
| -0.26148 | 0.21264 | -0.87653 | 0.88189
| 0.65422 | 1.6999 | 0.30629 | 3.4115
| 0.4042 | -2.0697 | -1.4659 | 6.5959
| 2.5439 | -0.29452 | 1.3379 | 8.3481
| 0.14304 | -0.52277 | 1.2767 | 1.9236
| -2.6659 | 1.082 | -0.30625 | 8.3715
| -0.30997 | 2.4004 | -0.12892 | 5.8744
| 0.98427 | 0.28885 | -1.2701 | 2.6653
| -0.6828 | 1.3921 | -0.93525 | 3.2789
| -0.48244 | 0.33435 | 0.61879 | 0.72744
| 0.57053 | -0.62347 | 0.12002 | 0.72862
| 0.060762 | 0.29721 | 0.46072 | 0.30429
| 0.49307 | 2.4514 | 0.91698 | 7.0935
| 1.1227 | 0.48848 | -0.40671 | 1.6645
| 0.33608 | 0.40764 | 0.08964 | 0.28716
| 0.27557 | -0.62154 | 1.804 | 3.7168
| -0.73653 | -0.10033 | -1.0759 | 1.7102
| -2.7257 | 0.26438 | 1.3112 | 9.2187
| -0.88825 | -0.057326 | 0.52066 | 1.0634
| -0.23479 | -0.49907 | -0.20196 | 0.34498
| -0.16203 | -0.1205 | 0.15598 | 0.065105
| 1.3797 | -0.52295 | 0.86901 | 2.9323
| 0.86771 | 0.58522 | 0.68397 | 1.5632
| 0.46811 | 0.65762 | 0.74223 | 1.2025
| 0.94494 | 0.64459 | -0.24035 | 1.3662
| 0.48982 | 0.040935 | -0.60926 | 0.61279
| -0.4653 | -0.35656 | 0.055816 | 0.34676
| 0.3423 | -0.50908 | -0.21578 | 0.4229
| -0.87591 | -1.717 | 1.1163 | 4.9614
| -0.14962 | -0.9263 | -0.61237 | 1.2554
| 1.774 | 0.21944 | -2.6794 | 10.3749
| -1.4346 | 1.924 | -0.58786 | 6.1052
| -0.38177 | -0.17365 | 0.61342 | 0.55219
| -2.1321 | 0.73139 | 0.88793 | 5.8692
| 1.5958 | -0.55504 | 0.34831 | 2.9758
| 0.11177 | 0.55826 | 1.9061 | 3.9574
| -0.015011 | 0.87087 | 0.53342 | 1.0432
| -0.38561 | 0.74062 | 0.37698 | 0.83933
| 1.0242 | 1.9304 | 0.18988 | 4.8117
| 0.1488 | 0.066882 | 1.7478 | 3.0816
| -0.36022 | -0.062511 | 1.0948 | 1.3322
| -0.2529 | 1.2388 | 0.33067 | 1.708
| 0.26388 | 0.32486 | -0.39998 | 0.33515
| -0.27886 | -0.74335 | 0.32606 | 0.73664
| -1.0218 | 0.88023 | -0.71255 | 2.3266
| -0.36213 | -0.74206 | -0.10448 | 0.69271
| 1.3192 | 0.51705 | 0.1286 | 2.0242
| 0.94086 | 0.13668 | -0.065053 | 0.90814
| -1.7871 | 0.44797 | 1.4557 | 5.5135
| 0.23751 | 1.0609 | 1.2332 | 2.7027
| -0.85554 | -0.016951 | 1.4577 | 2.857
| -0.76854 | 0.94271 | 0.81401 | 2.142
| -0.59301 | -1.2182 | 1.3112 | 3.555
| 1.6951 | -0.82575 | 0.057246 | 3.5587
| 1.3835 | 1.4172 | -0.04724 | 3.9247
| -1.0406 | 1.174 | -1.0572 | 3.5788
| -0.52948 | -0.24849 | 0.70389 | 0.83755
| 1.3097 | -0.0017011 | -0.70106 | 2.2067
| -0.66783 | -0.31983 | -0.7052 | 1.0456
| -0.1348 | -2.1632 | 1.2685 | 6.3067
| -0.0031941 | 1.1848 | 0.97343 | 2.3513
| 0.67484 | 0.37709 | 0.98909 | 1.5759
| 0.57795 | 0.20343 | 1.3076 | 2.0852
| -0.60407 | 0.15967 | 0.35097 | 0.51357
| -0.62459 | -1.6727 | -1.0446 | 4.2792
| 0.61414 | 2.0715 | -0.048642 | 4.6706
| -0.70621 | -0.90355 | -0.79539 | 1.9478
| -0.25504 | 1.0707 | -0.1964 | 1.25
| 0.36766 | -0.90851 | 0.4729 | 1.1842
| 0.84192 | -0.6505 | 0.13095 | 1.1491
| -0.21263 | -0.38175 | 1.0742 | 1.3449
| 0.12698 | -0.30832 | 0.62916 | 0.50702
| 0.816 | 0.01716 | 0.073683 | 0.67159
| 0.18176 | -1.0645 | 0.52812 | 1.4451
| 0.25798 | -0.9903 | -0.73392 | 1.5859
| -3.0534 | 0.64916 | 1.446 | 11.8352
| 1.3547 | -2.6336 | -1.124 | 10.0345
| 0.31188 | -0.57604 | -1.226 | 1.9322
| 2.01 | -0.69562 | -1.2995 | 6.2126
| 0.56958 | 0.2593 | -0.2418 | 0.45013
| 1.3237 | -0.67127 | -0.75117 | 2.767
| -2.1374 | 0.57059 | -0.40477 | 5.0578
| -0.5762 | -0.012598 | -1.1848 | 1.736
| 0.62413 | 0.78076 | -0.62189 | 1.3859
| 0.81152 | -1.2919 | 0.69909 | 2.8162
| 0.70587 | 0.19378 | 0.64143 | 0.94723
| 0.325 | -0.18148 | 0.25014 | 0.20113
| -0.041402 | -1.0388 | 1.0367 | 2.1555
| -0.070083 | -0.65607 | 1.5423 | 2.814
| 0.7687 | -1.1111 | -0.082669 | 1.8323
| -1.0031 | -0.84069 | 0.29023 | 1.7971
| -0.99959 | -1.3885 | 0.38259 | 3.0735
| 1.8793 | -0.95829 | -0.14722 | 4.4718
| -1.6229 | -0.766 | 0.68573 | 3.6909
| 1.2776 | 1.917 | -0.97152 | 6.251
| 0.35971 | -0.78019 | 0.3331 | 0.84904
| -1.9979 | -0.55483 | 1.8232 | 7.6235
| -0.076011 | 1.3975 | 0.93165 | 2.8268
| 1.9797 | 1.1327 | -1.281 | 6.8433
| 0.55006 | -0.75125 | 0.27121 | 0.9405
| 0.11685 | 1.1009 | -0.26644 | 1.2967
| 0.86574 | 1.9451 | -2.2145 | 9.4369
| 0.087427 | 0.20489 | 0.059819 | 0.053203
| -2.3891 | -0.32842 | -0.41627 | 5.9891
| -2.1976 | -1.2111 | -0.30151 | 6.3871
| 0.84354 | -0.21121 | 0.82198 | 1.4318
| 1.5037 | -1.2441 | -0.22513 | 3.8595
| -0.58674 | -0.51351 | -0.069826 | 0.61284
| 0.41596 | -1.4928 | 0.42514 | 2.5823
| -2.5779 | 1.1495 | -0.74481 | 8.5214
| 0.0044816 | 1.9704 | -0.53917 | 4.1734
| -0.29014 | 0.55619 | 0.37816 | 0.53653
| 0.026951 | -0.00027347 | -0.39945 | 0.16029
| -0.24447 | 1.3945 | 1.0438 | 3.0941
| -0.16885 | -0.89736 | 0.69193 | 1.3125
| -0.46636 | 0.27549 | 0.77199 | 0.88935
| -1.1241 | -0.14525 | -0.10093 | 1.295
| -0.79892 | -0.32487 | 0.56308 | 1.0609
| -0.55897 | -0.19617 | -0.98575 | 1.3226
| 1.8521 | 0.079152 | -0.59081 | 3.7857
| 0.74111 | -1.881 | 0.74947 | 4.6492
| -0.3877 | 0.90619 | -0.66997 | 1.4203
| 1.449 | -0.13452 | -0.58095 | 2.4553
| 0.95131 | 0.87016 | -0.0054704 | 1.6622
| 1.6795 | -0.97041 | -0.67101 | 4.2126
| 0.24644 | -0.6286 | 1.6637 | 3.2238
| -0.31644 | -1.0579 | -1.4384 | 3.2884
| -1.6505 | 2.1813 | -1.4685 | 9.6388
| 0.34485 | 0.61591 | -0.46995 | 0.71913
| -1.7974 | -1.1066 | -1.5341 | 6.8086
| -2.062 | -0.011555 | -0.22678 | 4.3033
| -1.019 | 1.4757 | 0.05976 | 3.2197
| -0.33573 | 1.2522 | 1.0987 | 2.888
| -1.192 | -0.47241 | -0.10922 | 1.656
| 0.15764 | -0.58101 | -0.15122 | 0.38529
| 0.82897 | -0.71611 | 0.34814 | 1.3212
| -1.0256 | 0.55248 | -1.6428 | 4.0559
| -0.57579 | 1.6247 | 1.7381 | 5.9921
| -0.49857 | 0.90329 | -0.68389 | 1.5322
| 0.59151 | -0.56221 | 0.97059 | 1.608
| -0.78711 | -1.8673 | 0.89995 | 4.9162
| 0.72447 | 0.39993 | -0.91098 | 1.5147
| -0.82204 | 0.61392 | 1.4146 | 3.0538
| -0.6199 | -0.32336 | 0.50959 | 0.74851
| -0.65172 | 0.63013 | 1.8116 | 4.1036
| -1.3591 | -0.80445 | 0.14034 | 2.514
| -1.2796 | -0.85435 | 0.72026 | 2.886
| -0.38959 | 1.5523 | -1.1444 | 3.8711
| 0.0033337 | -0.73262 | 1.4453 | 2.6256
| -0.62982 | 0.22653 | -0.2352 | 0.50331
| -0.6023 | 0.5955 | -1.2889 | 2.3786
| -0.36693 | 0.6013 | 0.1948 | 0.53414
| -0.021556 | -1.7272 | 0.75276 | 3.5505
| 0.010919 | -0.67605 | 1.014 | 1.4853
| -0.33463 | 1.3369 | 0.27772 | 1.9763
| -1.6968 | 0.94642 | 1.1392 | 5.0726
| -0.21007 | -1.1307 | -0.19769 | 1.3616
| -0.45496 | 1.4177 | 1.1748 | 3.5971
| 0.30022 | 0.89207 | -0.48122 | 1.1175
| 0.050563 | -0.019416 | 0.68246 | 0.46868
| 0.55099 | -0.77245 | -0.23599 | 0.95596
| 0.42245 | 0.33028 | -0.54146 | 0.58073
| -0.172 | -0.94338 | -0.30831 | 1.0146
| 0.17409 | -1.8428 | -0.69045 | 3.9031
| 0.16874 | 0.2027 | 0.20602 | 0.11201
| 0.22547 | 0.64025 | -0.95759 | 1.3777
| 0.70398 | 0.93612 | -0.66543 | 1.8147
| 0.4726 | 0.68115 | -2.3853 | 6.3767
| -0.46503 | -1.2106 | -0.5391 | 1.9725
| 1.2401 | -0.48209 | 0.85733 | 2.5052
| -1.3762 | -0.7519 | 0.80576 | 3.1084
| -1.5963 | 0.25599 | -0.60584 | 2.9808
| -2.338 | 0.85341 | -0.45375 | 6.4003
| -0.8978 | 0.26083 | -0.37575 | 1.0153
| -0.10144 | 0.50887 | -0.89169 | 1.0643
| 0.53038 | -0.14896 | 1.2973 | 1.9864
| -0.44506 | 2.1295 | -0.01347 | 4.733
| 0.90368 | -0.6057 | -0.39502 | 1.3396
| -1.7103 | 0.7188 | 0.45791 | 3.6516
| -1.4891 | -0.60174 | 0.69151 | 3.0577
| 0.7505 | -1.3706 | 0.036617 | 2.4432
| 0.72928 | -0.30188 | 0.51208 | 0.8852
| -0.45812 | 0.84457 | 0.060933 | 0.92689
| 0.97691 | 0.39172 | -1.624 | 3.7452
| -0.96762 | 1.4233 | 0.11309 | 2.9748
| 0.39368 | 0.083078 | 0.60324 | 0.52579
| 0.65234 | 0.88112 | 1.2486 | 2.761
| 0.26851 | -1.3465 | -0.9958 | 2.8768
| -0.37806 | 1.0775 | -0.54371 | 1.5994
| -0.16913 | 0.12539 | -1.2693 | 1.6554
| -1.1812 | -0.72791 | 0.30265 | 2.0167
| -1.3447 | 0.062207 | -1.2046 | 3.2631
| -2.0036 | -0.52789 | 0.82804 | 4.9789
| 0.60617 | 0.27166 | -2.5569 | 6.9787
| 1.8966 | -1.2033 | -1.0727 | 6.1959
| -0.26208 | -0.35247 | -0.47131 | 0.41505
| -1.8913 | 1.4872 | -2.5734 | 12.4111
| -0.66406 | -0.95492 | 0.2083 | 1.3962
| -0.53023 | -0.40907 | -0.91431 | 1.2845
| 0.026258 | -0.77659 | -0.27857 | 0.68138
| 0.046682 | 0.24249 | 1.3942 | 2.0048
| -1.3441 | -1.2788 | -1.0198 | 4.482
| -0.10845 | -0.77568 | 1.0131 | 1.6398
| -1.4279 | -2.217 | -0.96091 | 7.8772
| 0.11219 | 1.1205 | 0.01794 | 1.2684
| 0.81511 | 0.072274 | -2.1839 | 5.439
| 0.79129 | -0.95641 | -0.086334 | 1.5483
| 1.9359 | 0.82494 | 1.0136 | 5.4555
| -0.28606 | 0.35954 | 0.84251 | 0.92092
| 0.70176 | 2.0153 | 0.025602 | 4.5544
| 0.3851 | -1.9901 | -0.35789 | 4.2368
| 0.18878 | 1.1147 | 0.21015 | 1.3222
| 0.2969 | -0.29164 | 0.87028 | 0.93059
| -0.97865 | 0.86514 | 0.25811 | 1.7728
| 0.411 | 0.66285 | -2.0527 | 4.8217
| -0.39214 | -0.48245 | 0.39186 | 0.54009
| 1.4271 | 1.2123 | -2.1147 | 7.9784
| 1.3255 | 0.59891 | 0.17851 | 2.1475
| 0.063482 | -0.5583 | 0.41276 | 0.4861
| 0.12337 | -0.31738 | 0.056205 | 0.11911
| -1.1462 | 0.76081 | -0.18932 | 1.9285
| 0.11782 | -1.8233 | 1.0717 | 4.4869
| 0.39397 | -2.0068 | -1.9419 | 7.9537
| -0.85829 | -1.1424 | 0.68346 | 2.5088
| 0.35179 | 0.51447 | 0.77177 | 0.98407
| -0.53424 | -0.16598 | 0.056725 | 0.31617
| -0.68157 | -0.2257 | -2.2742 | 5.6874
| 0.81729 | 0.75394 | -0.16151 | 1.2625
| 0.37988 | 1.8908 | -0.52137 | 3.9914
| -0.24266 | -0.053209 | -0.0382 | 0.063176
| -0.91748 | -1.3075 | 0.29067 | 2.6359
| 0.37225 | -2.0605 | -0.34656 | 4.5043
| -0.16799 | -1.7052 | 1.0282 | 3.9933
| -1.6873 | -0.69333 | 0.80735 | 3.9795
| -0.19807 | -0.85648 | -1.8377 | 4.1499
| 1.0316 | 0.23538 | 0.96386 | 2.0486
| 0.95149 | -0.65178 | 0.53556 | 1.617
| 1.3374 | -0.25084 | 0.28217 | 1.9311
| -0.24578 | -0.044371 | -1.3422 | 1.8639
| 0.64256 | 0.094135 | 0.55209 | 0.72655
| 0.083004 | 0.88694 | 0.40562 | 0.95808
| 1.0169 | 1.1163 | 3.1936 | 12.4791
| 1.9367 | 1.2599 | -1.9605 | 9.1818
| -0.21249 | 0.32141 | 0.29502 | 0.2355
| -0.53724 | -1.7487 | -0.22207 | 3.3958
| 0.62469 | -0.2253 | 2.0485 | 4.6374
| -0.19724 | -0.61407 | -0.53043 | 0.69735
| 0.85521 | -0.74761 | 0.12146 | 1.3051
| -0.36628 | -0.65771 | -0.87148 | 1.3262
| -0.96971 | 0.83983 | -0.55713 | 1.956
| 0.27307 | -0.11743 | 0.96012 | 1.0102
| -0.79231 | -0.11184 | -0.014111 | 0.64046
| 0.76932 | -0.27325 | -0.10161 | 0.67684
| 2.7539 | -1.7629 | -0.62441 | 11.0818
| -0.16905 | 0.63405 | -2.6064 | 7.224
| -1.0882 | 0.4963 | -0.68149 | 1.895
| -0.59289 | 0.24928 | 0.1867 | 0.44851
| -1.2314 | -1.1902 | -1.4765 | 5.1129
| 1.7463 | -0.42119 | -0.34244 | 3.3443
| -0.077411 | -1.7976 | -0.53844 | 3.5272
| -0.16138 | 0.59616 | 1.278 | 2.0146
| 0.43738 | 0.31564 | -1.0305 | 1.3528
| 0.70809 | 1.5903 | -0.77205 | 3.6265
| 1.1234 | -0.0025756 | 0.34612 | 1.3819
| 1.6105 | -0.53621 | -0.72465 | 3.4064
| -0.25217 | -0.11861 | 0.034196 | 0.078829
| -1.5237 | 0.58677 | -0.84479 | 3.3797
| 0.68359 | -1.0502 | -1.6722 | 4.3666
| 0.17423 | 1.5753 | 0.81526 | 3.1766
| 1.5312 | 1.2366 | 0.2648 | 3.9438
| -1.2741 | 1.0115 | -0.39874 | 2.8055
| 0.76763 | 0.80047 | 0.47567 | 1.4563
| -1.3641 | -1.4901 | -1.6459 | 6.7904
| -0.28083 | 0.44732 | -0.22148 | 0.32801
| 1.5432 | -0.83484 | 1.1085 | 4.3072
| 0.033432 | 0.71835 | 0.66487 | 0.9592
| 0.30046 | -0.87811 | -1.4189 | 2.8746
| 1.7265 | 1.4627 | 0.87948 | 5.8939
| 0.65825 | -0.33366 | 1.0543 | 1.6561
| -0.094377 | 0.30303 | 0.30728 | 0.19515
| -0.5622 | -0.69687 | 1.1971 | 2.2347
| -1.2627 | 1.0395 | 0.92304 | 3.527
| 0.46579 | -0.62664 | -0.52835 | 0.88879
| -0.73569 | 1.047 | 0.50147 | 1.8889
| -1.0605 | 0.80588 | 0.75666 | 2.3466
| 0.057254 | -0.048201 | 0.20763 | 0.048711
| 1.0722 | -1.5327 | -0.27065 | 3.5721
| 0.84722 | -2.1217 | 2.0629 | 9.4752
| -0.51089 | 1.1633 | 0.84364 | 2.3261
| -0.66648 | 1.3041 | 1.8995 | 5.753
| -1.0013 | -1.5292 | -2.3413 | 8.8232
| 1.3443 | 0.075703 | -0.82777 | 2.498
| -1.2391 | -0.058124 | 0.95155 | 2.4441
| -0.88632 | 1.2151 | -1.086 | 3.4413
| 0.035863 | -0.55988 | 0.28723 | 0.39725
| -0.52945 | 0.20683 | -1.4375 | 2.3896
| -0.22613 | -0.76158 | 1.3251 | 2.387
| 0.33941 | 0.6667 | -0.95069 | 1.4635
| -0.092294 | 0.94403 | 0.096395 | 0.90901
| 1.0619 | -0.91574 | 4.2633 | 20.1422
| -0.20956 | 0.68737 | 2.0775 | 4.8323
| -0.27341 | 1.1085 | 0.056915 | 1.3067
| 1.2251 | -1.2477 | 1.4121 | 5.0517
| 0.94382 | 0.80091 | 0.77736 | 2.1365
| 0.23082 | -0.14842 | -0.0047889 | 0.075331
| 0.88803 | 0.35937 | 0.62355 | 1.3066
| -0.38653 | 1.1481 | 1.2435 | 3.0137
| -0.36754 | 0.20288 | 2.0507 | 4.3815
| 1.2593 | 0.33765 | 0.50747 | 1.9574
| 2.063 | -0.74589 | -0.25539 | 4.8776
| 0.48045 | -0.17425 | -2.3959 | 6.0017
| -0.42349 | -0.33385 | 1.8453 | 3.6958
| 0.73693 | -1.4436 | 0.32531 | 2.7329
| -0.77921 | 0.44205 | -0.68044 | 1.2656
| 1.4386 | -0.097203 | -0.477 | 2.3065
| 1.4653 | -1.2553 | 0.50607 | 3.9789
| 0.44051 | -0.45733 | -0.0086468 | 0.40328
| -0.052913 | -0.80996 | 0.43335 | 0.84663
| -1.4173 | -0.55773 | -0.70923 | 2.8229
| 0.78051 | -1.3145 | -0.61802 | 2.7192
| 0.083006 | 0.91497 | 0.25597 | 0.90959
| 0.95392 | -0.12219 | 1.0702 | 2.0703
| -0.6594 | 0.31246 | -1.3386 | 2.3244
| 0.057621 | 1.2227 | 1.1338 | 2.7839
| -1.0812 | -0.36563 | -0.51013 | 1.5629
| -1.2183 | 1.3479 | 1.1303 | 4.5787
| 0.81692 | 0.33978 | 0.27052 | 0.85598
| 1.2361 | -0.25572 | -0.051344 | 1.5961
| 0.96652 | 1.4109 | 0.86456 | 3.6723
| 1.5553 | -0.81367 | 0.14594 | 3.1022
| 0.44544 | 1.1285 | -1.324 | 3.2248
| 0.96783 | -0.083186 | -0.0093831 | 0.9437
| -0.93782 | -0.76188 | -0.63614 | 1.8646
| -1.3085 | 0.094012 | -1.7592 | 4.8158
| -0.62307 | 2.999 | -1.483 | 11.5815
| 1.1275 | -0.0087235 | 1.7295 | 4.2624
| 3.011 | -0.77074 | 0.58685 | 10.0047
| 1.0629 | -0.93491 | -0.10386 | 2.0147
| -0.49491 | 1.892 | -0.034152 | 3.8257
| 0.077481 | -0.32055 | -1.1125 | 1.3464
| 0.34279 | 1.3545 | 0.9866 | 2.9255
| 1.4247 | -0.1859 | -1.0059 | 3.0763
| 0.70707 | -0.36027 | 0.064632 | 0.63391
| 0.69292 | -0.38939 | 1.3363 | 2.4174
| -0.76277 | 0.46633 | -0.54382 | 1.095
| -0.54617 | -0.50235 | 0.77064 | 1.1445
| -1.9503 | 0.087106 | -0.28575 | 3.8929
| 1.0626 | -0.93093 | -0.11047 | 2.0079
| -1.1176 | 0.64877 | -0.74073 | 2.2186
| 1.0136 | 0.59765 | 2.2885 | 6.6219
| 1.5931 | 1.5598 | -1.638 | 7.6541
| -0.91061 | -0.60822 | 1.3823 | 3.11
| -1.6361 | -0.11439 | 0.029654 | 2.6907
| 1.5614 | 1.1333 | 1.3166 | 5.4556
| 2.448 | -0.38192 | 1.9433 | 9.9151
| 1.4707 | 0.45299 | 1.6244 | 5.0066
| 2.2902 | 0.78401 | -0.52469 | 6.1351
| 0.61993 | -1.2257 | 0.40014 | 2.0467
| -0.67965 | -1.7325 | 0.25162 | 3.5268
| 0.37028 | -0.30223 | 0.19575 | 0.26677
| 1.3484 | -1.8783 | -0.41697 | 5.52
| 0.70749 | -0.6736 | 1.1454 | 2.2663
| 1.3339 | -2.3797 | 1.1393 | 8.7404
| -0.51825 | -1.6912 | 1.2488 | 4.688
| 0.32496 | -0.24262 | 0.41435 | 0.33615
| 0.2826 | 2.0929 | -1.1724 | 5.8349
| -0.48717 | 0.054245 | -0.40312 | 0.40279
| -0.14272 | -0.17904 | 0.30731 | 0.14686
| -1.2136 | -0.76647 | 1.3372 | 3.8486
| 0.17193 | 0.26329 | 0.93131 | 0.96621
| -0.53589 | 0.15488 | -0.95754 | 1.228
| -0.30554 | 1.7012 | -0.19061 | 3.0237
| 0.98184 | 1.9473 | -0.20941 | 4.7997
| -0.69045 | -1.5402 | -1.9015 | 6.4648
| 1.187 | 0.71446 | -0.080004 | 1.9258
| 0.95974 | -1.1683 | -0.49231 | 2.5283
| 0.34784 | -0.72565 | -0.40409 | 0.81085
| -0.37307 | -0.42785 | -1.1949 | 1.7499
| -0.073821 | 0.72174 | 1.098 | 1.7319
| -2.4454 | -0.57312 | 0.46447 | 6.5241
| 0.24746 | -0.02972 | 0.93741 | 0.94086
| 1.825 | 0.14385 | 0.86641 | 4.1019
| 0.27261 | 0.23677 | 1.1588 | 1.4731
| 0.86864 | 0.64139 | 1.9061 | 4.7991
| 1.3937 | -0.19781 | 1.4843 | 4.1845
| -1.3707 | 0.35294 | 0.89248 | 2.7999
| -0.69123 | 0.0023368 | 0.71663 | 0.99136
| 0.56118 | 0.34583 | 0.81048 | 1.0914
| 0.01466 | 0.30093 | 0.86179 | 0.83346
| 0.76192 | 0.21779 | -1.1408 | 1.9295
| -0.50063 | -0.86531 | -0.5534 | 1.3056
| -0.7035 | 0.66399 | -0.95471 | 1.8473
| 1.5553 | -0.15749 | 0.50603 | 2.6997
| 0.19927 | 1.0524 | -0.48101 | 1.3786
| -1.3208 | 0.32643 | -0.164 | 1.878
| 0.23235 | 1.3095 | -0.20436 | 1.8105
| 1.107 | 0.12427 | 1.6787 | 4.059
| -1.3483 | 1.021 | 2.3073 | 8.1841
| -0.44211 | -0.096756 | -0.2581 | 0.27144
| 0.038683 | -0.099045 | 0.89707 | 0.81605
| -0.46937 | 1.8074 | 0.55278 | 3.7925
| -0.26474 | 0.18093 | 0.83178 | 0.79469
| -0.4661 | 2.4348 | -0.44957 | 6.3476
| -0.055382 | -0.83456 | -0.064071 | 0.70367
| 0.054314 | 0.72474 | 1.1764 | 1.9121
| 0.096244 | -0.02461 | 1.1307 | 1.2884
| 1.3779 | 0.23284 | 0.53438 | 2.2383
| -0.13061 | -2.2074 | 0.033669 | 4.8907
| -0.72943 | 0.16519 | 1.7494 | 3.6198
| -0.12345 | 0.68129 | 0.71494 | 0.99054
| 0.67878 | -1.4211 | 1.0982 | 3.6863
| 0.3368 | -0.29929 | 2.2734 | 5.3712
| 0.68452 | 0.88947 | 0.22919 | 1.3123
| -0.95393 | 0.38643 | 2.1943 | 5.8741
| -0.57897 | -0.28956 | 0.2311 | 0.47245
| 0.23602 | -1.2441 | -0.087492 | 1.6111
| -0.21089 | 0.099438 | -1.2436 | 1.601
| 2.245 | 1.3382 | -1.0679 | 7.971
| 0.2757 | -0.098238 | -0.77372 | 0.6843
| 0.80885 | 0.60562 | 0.34983 | 1.1434
| -0.39943 | 0.58359 | -0.65595 | 0.93039
| 0.39571 | 0.068654 | 0.10569 | 0.17247
| 0.74946 | 0.95084 | 0.15252 | 1.4891
| 1.0643 | 0.55258 | -1.3415 | 3.2377
| -1.3984 | -0.92532 | -0.15037 | 2.8343
| -0.71295 | -0.13201 | 1.1053 | 1.7473
| -0.56781 | -0.61649 | -1.5246 | 3.0267
| 0.78519 | 0.36533 | 1.1439 | 2.0586
| 0.99465 | 1.4211 | 1.4326 | 5.061
| -0.5416 | 0.62429 | -0.11844 | 0.69709
| 1.3107 | -0.71564 | -0.081419 | 2.2366
| -0.064474 | -0.67459 | 1.0463 | 1.554
| -0.37376 | 0.007198 | 2.2753 | 5.3169
| 0.32338 | 0.48686 | -0.16782 | 0.36977
| 0.32548 | 0.5538 | -1.0553 | 1.5264
| -0.06901 | 1.4706 | -0.2687 | 2.2397
| -1.8976 | -0.3612 | 0.87895 | 4.5038
| -0.47981 | 2.6104 | -0.20644 | 7.0872
| 1.0673 | 1.2626 | -0.094796 | 2.7422
| -0.42477 | 0.31067 | -1.6822 | 3.1067
| 0.21999 | 0.49926 | -0.13521 | 0.31593
| 0.20505 | -1.5008 | 0.12201 | 2.3094
| -1.127 | 1.3645 | -0.064721 | 3.1361
| -1.6193 | -0.97558 | -0.23039 | 3.627
| -0.50487 | -1.0415 | -1.3941 | 3.2831
| 0.78271 | 0.75376 | -1.3427 | 2.9835
| -1.1096 | -0.33047 | -1.2054 | 2.7935
| 0.51985 | 0.79569 | 1.0318 | 1.968
| 0.15098 | -0.1606 | -1.0628 | 1.1782
| -1.6729 | 0.47549 | 0.37197 | 3.1631
| -1.955 | 0.57313 | 0.48638 | 4.3872
| 0.77308 | -0.66682 | 0.56921 | 1.3663
| 1.7701 | -1.7393 | 1.7204 | 9.118
| 0.073798 | 0.71051 | -0.90851 | 1.3357
| 1.3604 | -0.77068 | 0.42956 | 2.6292
| -0.68437 | 0.60803 | 0.61202 | 1.2126
| -1.1129 | 2.0502 | 0.12706 | 5.458
| 0.80325 | -0.04895 | 0.0062835 | 0.64764
| -1.2819 | -1.188 | 1.2801 | 4.6933
| 1.0485 | -0.029367 | -0.98627 | 2.0729
| 1.7113 | 0.12558 | -0.64885 | 3.3655
| -2.0765 | 0.29586 | 1.0698 | 5.5437
| 0.86058 | 0.82327 | -1.7885 | 4.617
| 0.70238 | 0.30512 | 1.2565 | 2.1652
| 0.39542 | -0.54411 | 0.77899 | 1.0592
| -0.19045 | 1.8034 | 0.83641 | 3.9879
| 0.54258 | 0.68289 | 0.76433 | 1.3449
| -0.49812 | 0.42278 | 1.4234 | 2.4529
| 1.3102 | 0.17425 | -1.1624 | 3.0981
| -0.29448 | 0.74157 | 1.456 | 2.7567
| -0.92765 | -0.7809 | 1.0762 | 2.6285
| 0.58926 | -0.67182 | -0.47857 | 1.0276
| -0.22959 | -0.4791 | 0.44306 | 0.47855
| -0.39464 | 0.95032 | -0.79724 | 1.6944
| -0.12062 | 0.78376 | 0.18266 | 0.66219
| -1.2409 | -0.72718 | -0.14644 | 2.09
| -0.40257 | -0.26706 | -1.0347 | 1.3041
| -0.44888 | 0.80082 | 0.653 | 1.2692
| -0.31629 | 1.7629 | -1.928 | 6.9252
| -0.16922 | 0.58912 | -0.51421 | 0.64011
| -0.87158 | 0.0072772 | -0.066448 | 0.76412
| 0.20129 | 0.71911 | -2.91 | 9.026
| 0.33453 | 0.71936 | 0.45254 | 0.83419
| 1.0267 | -1.3437 | 0.046772 | 2.8617
| 0.23406 | 2.1064 | -0.81156 | 5.1503
| -0.90644 | 1.8821 | 0.28015 | 4.4424
| -1.7626 | 0.84157 | -1.1135 | 5.0549
| 0.41655 | -0.05957 | -0.34634 | 0.29702
| 0.0078492 | 0.0098993 | -0.01565 | 0.00040454
| -1.3151 | 1.2587 | -0.63836 | 3.7212
| -0.059887 | -0.14571 | -0.24188 | 0.083321
| -0.68238 | 0.29332 | -1.3679 | 2.4227
| 0.72783 | 0.72428 | -0.23019 | 1.1073
| -0.47966 | -0.34287 | -2.0653 | 4.6131
| 0.51611 | 0.76159 | -0.48921 | 1.0857
| -0.068466 | 0.44844 | -1.7001 | 3.0962
| 1.8479 | -0.8631 | 0.42815 | 4.3432
| 0.90582 | 1.0673 | 0.2749 | 2.0351
| 0.48967 | -0.5413 | -0.07835 | 0.53893
| 0.15684 | 0.25425 | 0.0757 | 0.094972
| 0.11297 | -1.2712 | -2.5532 | 8.1478
| -0.42648 | -0.30222 | -0.38183 | 0.41902
| 0.61789 | -0.25876 | -0.74738 | 1.0073
| 1.2319 | 0.72013 | -0.32333 | 2.1408
| 0.239 | -0.67702 | 0.99881 | 1.5131
| 0.54578 | -0.44488 | 1.807 | 3.7612
| -1.3329 | -0.27101 | 1.5013 | 4.104
| 1.4385 | 0.39625 | 1.9025 | 5.8456
| 1.7807 | 0.50594 | -0.39281 | 3.5812
| -0.85628 | -0.18642 | 1.6051 | 3.3442
| -1.4471 | 0.87468 | -0.60073 | 3.22
| 1.6236 | 0.18689 | -0.84739 | 3.3889
| 0.3753 | 0.14663 | -0.2623 | 0.23115
| -0.38113 | 2.0792 | -0.45489 | 4.6751
| -1.9675 | -0.24665 | -0.59501 | 4.2861
| 0.029212 | 0.76604 | -0.49241 | 0.83014
| 1.2652 | 1.7076 | -0.40909 | 4.6841
| 0.56483 | -0.26237 | -0.56533 | 0.70747
| -2.0381 | -0.15966 | -1.0174 | 5.2146
| 0.39121 | -0.75423 | 1.5918 | 3.2557
| 0.028661 | -0.85757 | -0.088448 | 0.74407
| 0.23109 | 0.77188 | 2.1795 | 5.3993
| -0.60426 | -0.42922 | 0.068075 | 0.55399
| 1.336 | -0.8639 | 1.1223 | 3.7909
| -0.89013 | 1.4887 | 0.78823 | 3.6299
| 0.3631 | -0.29594 | 1.3386 | 2.0112
| -0.58416 | 0.62239 | -0.43209 | 0.91531
| 1.5793 | 0.24616 | 0.99644 | 3.5478
| -0.058943 | 1.5359 | -0.29549 | 2.4497
| -0.04925 | -1.1666 | 0.37286 | 1.5025
| 1.6362 | -0.026619 | 0.21434 | 2.7237
| 0.17685 | 1.3201 | -1.3096 | 3.4888
| 0.73689 | 1.4346 | 0.10202 | 2.6114
| 0.47056 | 0.37447 | 0.53477 | 0.64763
| -1.1895 | 0.48196 | 0.067011 | 1.6517
| -2.1797 | -0.74244 | 0.5078 | 5.5603
| 1.0177 | 0.78171 | -0.58004 | 1.9833
| -0.29304 | 1.4696 | -0.90973 | 3.0732
| 0.16841 | 0.39675 | 0.40154 | 0.34701
| 0.020514 | 0.94608 | -0.6533 | 1.3223
| 0.76295 | -0.30139 | 0.60131 | 1.0345
| 0.13384 | 0.2427 | -2.0359 | 4.2216
| 0.44582 | 1.4851 | -0.93394 | 3.2766
| -1.3605 | -0.36034 | 0.38463 | 2.1288
| -1.6658 | 1.1367 | -0.98451 | 5.0361
| -0.29858 | 0.13566 | -0.57288 | 0.43575
| 0.10277 | 0.74807 | 2.0423 | 4.7411
| -1.2073 | 1.1748 | 1.0504 | 3.9411
| 0.06723 | -2.4149 | -0.063147 | 5.8404
| 1.3265 | 0.22859 | 2.0547 | 6.0337
| -1.7014 | 0.077374 | -2.122 | 7.4039
| -0.21854 | -0.73388 | 0.84486 | 1.3001
| 1.235 | 0.41783 | -1.2036 | 3.1484
| -0.4456 | -1.3482 | -0.86902 | 2.7714
| -0.50012 | -0.10087 | 0.79361 | 0.89011
| 0.35476 | 0.49347 | 0.20545 | 0.41158
| 0.27843 | 0.29452 | -1.5973 | 2.7157
| -0.38176 | -0.75114 | 2.6894 | 7.9427
| 1.7182 | -0.92245 | 0.28037 | 3.8818
| 0.12553 | -0.22494 | -0.76717 | 0.65491
| -0.39096 | -1.3141 | -0.25406 | 1.9444
| -0.36226 | -0.2055 | 0.44425 | 0.37081
| -1.3184 | 2.4009 | -0.38992 | 7.6545
| 0.157 | 0.69116 | -0.96511 | 1.4338
| -0.93744 | -0.84715 | 0.74059 | 2.1449
| 1.0126 | -0.71468 | 0.38183 | 1.682
| -0.52876 | 0.075792 | 0.33723 | 0.39905
| -0.59488 | 1.0205 | 0.80923 | 2.0502
| -0.81295 | 0.48606 | -0.94491 | 1.79
| -1.0849 | -0.84873 | -0.36772 | 2.0327
| 1.0105 | -0.22885 | -1.3814 | 2.9817
| 0.39432 | 0.19994 | -0.21309 | 0.24087
| 1.3213 | -0.17794 | 0.23366 | 1.8322
| 0.46594 | 0.81677 | -0.50863 | 1.1429
| -0.6915 | -0.82779 | 0.064356 | 1.1676
| 1.542 | -1.826 | -1.5847 | 8.2235
| 0.41329 | 0.48244 | -0.34633 | 0.5235
| -0.94148 | 0.44421 | 0.34049 | 1.1996
| 1.4203 | -0.26823 | 0.23704 | 2.1455
| 1.356 | -1.4179 | -0.71443 | 4.3598
| -0.81182 | 0.19223 | 0.61767 | 1.0775
| -0.36423 | 0.007283 | -0.70666 | 0.63208
| -0.55185 | 2.2015 | -1.664 | 7.9198
| -0.73612 | -1.3114 | 0.96873 | 3.2001
| -1.3336 | -0.49828 | 1.3961 | 3.9758
| -0.67915 | 1.7244 | 0.97477 | 4.3849
| 0.45254 | 0.10099 | -0.92715 | 1.0746
| -0.88081 | 0.92561 | -0.75517 | 2.2029
| 0.92121 | -1.7743 | -0.28296 | 4.0769
| 1.7113 | -0.41086 | 0.50923 | 3.3566
| 1.1289 | -0.10317 | -0.36964 | 1.4217
| -0.98128 | -0.16608 | 2.1904 | 5.7883
| -0.20802 | 0.87238 | -0.42072 | 0.98132
| -0.79967 | 0.12547 | -0.31465 | 0.75422
| -0.45872 | 0.26445 | -0.11892 | 0.2945
| -0.71238 | 0.95659 | -2.212 | 6.3156
| -2.4382 | 0.27719 | 1.0383 | 7.0996
| -0.45091 | 1.4814 | -0.61548 | 2.7766
| 0.17158 | -0.045635 | 0.16904 | 0.060097
| -0.91014 | -1.0446 | -0.4732 | 2.1435
| 0.81124 | 1.3589 | -1.4269 | 4.5409
| 1.3651 | -0.81818 | -0.63567 | 2.9369
| -0.41107 | 0.40044 | -0.17357 | 0.35946
| 1.2873 | -0.87062 | -0.90956 | 3.2424
| -1.4982 | -1.0196 | -0.45516 | 3.4914
| -1.2197 | -0.93254 | -1.9726 | 6.2485
| 0.046903 | 0.034896 | -0.54748 | 0.30315
| -0.90157 | 0.2805 | 1.2822 | 2.5356
| 0.90486 | -0.43871 | -0.079057 | 1.0175
| 0.30345 | -0.61066 | 0.44378 | 0.66193
| -0.2655 | 0.73232 | -0.07088 | 0.6118
| -0.19966 | 1.4659 | 0.093072 | 2.1974
| -1.1017 | 0.5625 | 0.58671 | 1.8743
| -1.4223 | 0.43672 | 0.82739 | 2.8982
| -0.33975 | -0.83488 | -1.4745 | 2.9866
| -1.0307 | -0.13632 | 0.60464 | 1.4466
| -0.5277 | 0.15678 | 0.03673 | 0.30439
| -2.2127 | 1.6194 | -1.4393 | 9.5902
| 0.59825 | -1.0995 | 0.3347 | 1.6789
| -0.092432 | -0.46463 | 0.64964 | 0.64645
| -0.65967 | 1.6228 | -0.27071 | 3.142
| 1.0064 | 1.6546 | -0.15047 | 3.7734
| -0.72566 | 0.72933 | 1.3181 | 2.796
| -1.8111 | 0.047559 | 1.4122 | 5.2765
| -0.61051 | 1.6649 | -0.76683 | 3.7326
| 0.52188 | 0.12252 | -0.13504 | 0.30561
| 1.1018 | 0.69138 | 1.0259 | 2.7443
| -0.083672 | 0.60128 | -0.052007 | 0.37125
| -0.88146 | -0.64481 | -0.31047 | 1.2891
| 0.79077 | 0.70862 | -0.23593 | 1.1831
| -0.72847 | 1.006 | -0.12775 | 1.559
| -0.72838 | -0.019896 | -1.0757 | 1.688
| 1.4578 | 0.26527 | -0.50206 | 2.4476
| -0.15176 | -0.071446 | 1.7746 | 3.1773
| 1.046 | 0.14041 | -0.54877 | 1.4151
| 0.817 | 1.0267 | -0.09122 | 1.7299
| -0.28755 | 0.29227 | 0.18067 | 0.20075
| -1.6619 | 1.544 | -1.2321 | 6.6638
| -0.66856 | 1.3172 | 1.2064 | 3.6373
| 1.3032 | 0.7761 | 0.18107 | 2.3335
| 0.81466 | 0.50114 | 1.1869 | 2.3236
| 1.4253 | -0.90605 | -0.2832 | 2.9326
| 1.4882 | -0.69187 | -0.039835 | 2.695
| -0.31481 | 0.59525 | 0.11684 | 0.46709
| -0.75553 | 0.35681 | -0.041339 | 0.69985
| -0.49266 | 0.35918 | -0.91678 | 1.2122
| 0.039044 | -0.20213 | -0.32474 | 0.14784
| -1.4905 | 1.4303 | 1.3389 | 6.0603
| 0.46802 | -0.80024 | 0.41191 | 1.0291
| -1.0376 | 1.1186 | -0.5972 | 2.6846
| -0.59738 | 1.6725 | 1.0997 | 4.3633
| -1.7788 | -0.69722 | 0.42591 | 3.8316
| -1.984 | 2.9256 | 0.7235 | 13.0188
| -0.74498 | 1.4685 | 0.43222 | 2.8983
| -0.50361 | -0.91498 | 0.50944 | 1.3503
| -0.231 | 0.7692 | 0.089539 | 0.65304
| -0.5087 | -0.46842 | -0.52328 | 0.75201
| -0.11234 | 0.43085 | -0.82403 | 0.87728
| -0.84064 | 0.0090307 | 2.3509 | 6.2333
| 0.82939 | -0.16014 | 0.91978 | 1.5595
| 0.28554 | -1.1638 | -0.18609 | 1.4705
| 0.52383 | 1.4274 | 2.3618 | 7.8899
| -0.54136 | -0.9443 | -0.60723 | 1.5535
| -0.2316 | -0.72174 | -0.32462 | 0.67993
| -0.93724 | -2.0622 | 0.068671 | 5.1358
| 0.28501 | 0.60444 | -0.43976 | 0.63997
| -0.83387 | -1.6063 | -0.71903 | 3.7924
| -0.71204 | -1.0848 | 0.98976 | 2.6634
| 0.29912 | -0.27147 | -0.36598 | 0.29711
| 0.15115 | -1.4053 | 0.54829 | 2.2984
| -0.96274 | -0.47103 | -0.39751 | 1.3068
| 0.78513 | -1.1597 | 0.65577 | 2.3913
| -1.2478 | -0.32189 | 0.64666 | 2.0788
| 0.65191 | 0.77348 | -0.18557 | 1.0577
| 0.77164 | -0.95433 | -0.33983 | 1.6217
| 0.36803 | -0.98806 | -0.12844 | 1.1282
| 1.5715 | 0.71192 | -0.53226 | 3.2596
| -0.44096 | 2.0133 | -1.2857 | 5.901
| -0.36867 | 2.213 | -1.6263 | 7.6782
| -0.60974 | 1.9748 | -0.16919 | 4.3003
| 0.65155 | 0.737 | -0.81259 | 1.628
| -0.37913 | -0.48143 | 0.59786 | 0.73295
| -0.51911 | -0.78488 | -0.3862 | 1.0347
| 1.1027 | -2.1613 | 2.7242 | 13.3084
| 0.64879 | 0.84476 | -0.53901 | 1.4251
| -0.32341 | -0.62938 | 1.3315 | 2.2735
| 1.3789 | 0.8636 | 0.89432 | 3.4469
| 0.00037348 | -0.6511 | 0.42515 | 0.60468
| 0.79793 | -1.586 | -1.1035 | 4.3697
| -0.25022 | 0.085311 | -0.074298 | 0.075408
| 0.96563 | 0.59191 | -0.81843 | 1.9526
| 0.79883 | 0.43863 | 0.1694 | 0.85922
| 0.8729 | 0.29475 | 1.0337 | 1.9173
| -0.090708 | 1.0316 | 0.12532 | 1.0882
| -1.3576 | 0.49441 | 1.3227 | 3.8369
| 2.1767 | 0.63933 | 0.3129 | 5.2447
| -0.15641 | -0.66312 | -0.63131 | 0.86274
| 1.094 | 1.4675 | -0.51595 | 3.6164
| -0.62973 | -0.73524 | -0.3631 | 1.069
| -1.2541 | 1.0788 | -0.57804 | 3.0707
| -0.4085 | -0.79492 | 0.23517 | 0.85408
| 1.072 | -0.58365 | -0.14792 | 1.5118
| 0.65702 | 1.9536 | 0.29419 | 4.3347
| 0.49227 | 0.46617 | 0.16366 | 0.48643
| 1.7305 | 0.13883 | -0.95474 | 3.9255
| -2.4738 | 0.19763 | 1.1888 | 7.5719
| -0.93754 | 0.065799 | -1.6912 | 3.7433
| 0.35294 | 1.9122 | 1.6273 | 6.4292
| 0.14811 | 1.1061 | -0.57219 | 1.5728
| 0.004587 | 1.4047 | -0.10331 | 1.984
| 1.1266 | -0.50322 | 0.31012 | 1.6186
| -1.285 | -1.299 | -0.76886 | 3.9298
| -1.0126 | -1.5999 | 1.2894 | 5.2475
| 0.41847 | -0.59286 | 0.97796 | 1.483
| -0.17928 | 0.88435 | -0.36625 | 0.94837
| 0.39918 | 1.744 | 0.2675 | 3.2726
| -0.61345 | -1.3628 | 1.7043 | 5.1381
| 0.031524 | 1.2281 | 0.70242 | 2.0027
| 1.2852 | 0.46296 | -1.2465 | 3.4198
| -1.7295 | 0.85086 | -0.20012 | 3.7553
| 1.1058 | 1.3999 | 0.68054 | 3.6456
| -2.5 | 1.0545 | -0.20454 | 7.404
| -0.37641 | -2.135 | -0.70299 | 5.1942
| 0.2126 | -0.20526 | -0.047739 | 0.08961
| -2.4254 | -0.80524 | -1.1117 | 7.7671
| -0.52827 | -0.10904 | 1.1369 | 1.5835
| -1.8107 | 0.10178 | 0.0052764 | 3.289
| -1.8961 | -1.2166 | 0.15341 | 5.0987
| -1.6673 | -0.0024416 | -0.27103 | 2.8533
| -0.5995 | -1.6918 | 0.63881 | 3.6297
| -0.59143 | 0.91684 | 0.12832 | 1.2068
| 0.11313 | 1.7905 | 0.060628 | 3.2222
| 0.11235 | 0.40412 | -0.037137 | 0.17731
| 0.42937 | -1.7105 | 0.11603 | 3.1235
| 0.77583 | 0.20796 | 0.53127 | 0.92741
| 0.72296 | -0.49802 | -0.48104 | 1.0021
| 0.10773 | -0.68375 | 1.0985 | 1.6858
| -0.44096 | 0.10559 | 0.72581 | 0.73239
| 0.46196 | 1.6942 | 0.6284 | 3.4786
| 1.3327 | -0.32332 | 0.10509 | 1.8918
| -0.5403 | -0.54147 | -2.5637 | 7.1576
| 0.81033 | -0.68702 | -1.1107 | 2.3623
| 0.55684 | -1.1779 | -1.35 | 3.5198
| -1.0255 | 1.7958 | 1.2139 | 5.7501
| 1.1221 | -0.97081 | -0.38013 | 2.346
| 1.3489 | -1.4508 | -0.088054 | 3.9324
| 0.18974 | -0.7492 | 0.25814 | 0.66394
| -0.2172 | 0.69058 | -0.57214 | 0.85142
| -0.35054 | 0.50687 | -0.10881 | 0.39163
| 0.96294 | 0.83962 | -0.043605 | 1.6341
| 1.0266 | -0.82026 | 0.7784 | 2.3327
| 0.91072 | 0.62285 | -0.34123 | 1.3338
| 0.32518 | -0.49576 | -0.10315 | 0.36216
| 0.40977 | -1.164 | -0.39274 | 1.677
| -0.10191 | -0.83837 | 0.61238 | 1.0883
| -1.8155 | 0.27798 | 0.98693 | 4.3474
| -0.54961 | 1.0927 | -1.2495 | 3.0573
| -0.050426 | -0.10558 | 0.80546 | 0.66245
| -0.52647 | -0.27422 | -1.1594 | 1.6965
| 0.26445 | -0.5765 | 1.1259 | 1.6699
| 0.93942 | -1.4549 | -0.29516 | 3.0863
| 0.41734 | 1.7376 | 0.58544 | 3.5361
| -0.49102 | 0.69926 | -0.50532 | 0.98541
| 0.044643 | -1.6351 | 0.1588 | 2.7007
| -0.8334 | -0.15336 | -0.22024 | 0.76658
| 0.0069656 | 0.25944 | 1.4187 | 2.08
| 0.40863 | 0.61001 | 0.074002 | 0.54457
| 0.76565 | 0.10071 | 0.75385 | 1.1647
| -0.16506 | -0.26331 | 1.8379 | 3.4743
| -0.037097 | 0.27887 | 0.34614 | 0.19895
| 1.0508 | 0.29444 | 0.60279 | 1.5543
| -0.78791 | 0.17151 | 1.3206 | 2.3942
| 0.96773 | -0.68736 | -0.42324 | 1.5881
| -0.18795 | 1.326 | -1.2231 | 3.2896
| 0.65737 | -0.72415 | 2.5864 | 7.6459
| 0.76078 | 0.54105 | -0.66004 | 1.3072
| -0.46484 | 1.8934 | -0.39171 | 3.9546
| -1.5264 | -1.307 | 0.72008 | 4.5567
| -0.30782 | -0.94917 | 0.3545 | 1.1214
| -0.54686 | -0.60665 | 0.11799 | 0.68099
| 0.14156 | -1.4352 | 1.6384 | 4.7642
| 0.51966 | -0.62307 | -2.4052 | 6.4434
| 0.15925 | -2.5953 | 0.14268 | 6.7813
| -0.73801 | -0.057327 | -1.8539 | 3.9849
| 1.6843 | 1.8275 | -0.91747 | 7.0183
| -0.036877 | 0.094134 | -1.8998 | 3.6196
| 0.58332 | 1.3883 | 0.069411 | 2.2724
| -0.17656 | 1.0494 | 0.4559 | 1.3402
| -0.00751 | 0.095581 | -0.59152 | 0.35908
| -0.16496 | -0.53815 | -2.5588 | 6.8642
| 0.0043831 | -0.81933 | 0.068556 | 0.67603
| 0.9423 | 0.27198 | -0.51894 | 1.2312
| -0.05882 | 1.5659 | 0.32185 | 2.559
| -0.86573 | 0.064109 | -1.5725 | 3.2263
| -0.2117 | -1.4725 | 0.44878 | 2.4144
| -0.0088148 | -0.52095 | -0.11775 | 0.28534
| 1.4779 | 0.10932 | -0.1461 | 2.2176
| 1.4018 | -1.5796 | -1.3361 | 6.2453
| -0.22516 | 0.9392 | 1.1733 | 2.3095
| -0.043055 | 0.13567 | 0.68584 | 0.49064
| -0.88476 | -0.84146 | -0.2304 | 1.5439
| 1.6532 | 1.6646 | -2.2033 | 10.3586
| 1.1792 | 1.5864 | 0.57049 | 4.2327
| 1.6918 | 1.3063 | -0.024037 | 4.5694
| -1.0555 | 1.6195 | -0.98934 | 4.7159
| 0.02045 | -0.18292 | -1.4024 | 2.0005
| -1.5521 | 0.13153 | -1.5587 | 4.8558
| -0.39976 | 0.75111 | 0.20562 | 0.76625
| 0.99858 | -0.18972 | 0.586 | 1.3766
| 1.1161 | 0.15171 | 0.64747 | 1.688
| 0.083753 | 0.44312 | -1.4737 | 2.3751
| 0.7781 | 0.96586 | -0.79644 | 2.1726
| 0.99996 | 0.066649 | 2.4274 | 6.8965
| -0.63765 | -0.40278 | -2.2403 | 5.5879
| 0.11246 | -1.1804 | 1.8524 | 4.8372
| 1.0122 | -0.36178 | -0.69244 | 1.6349
| -0.13554 | -0.2419 | -0.48604 | 0.31312
| 0.76691 | 0.23394 | -0.87906 | 1.4156
| 0.21698 | -1.6598 | 0.59125 | 3.1517
| 1.2254 | -0.045803 | 1.9879 | 5.4556
| 0.40633 | 0.58375 | -0.77934 | 1.1132
| 1.5433 | 0.32912 | 1.763 | 5.5983
| -2.2586 | 0.20285 | -1.3804 | 7.0478
| 0.22419 | -0.73639 | 1.5389 | 2.9609
| -0.55637 | 0.15245 | -2.061 | 4.5804
| 0.66483 | -1.5293 | -1.2879 | 4.4396
| -0.047147 | -1.0523 | 0.01927 | 1.11
| 0.99359 | -0.19589 | -1.1253 | 2.292
| -0.80927 | 0.27141 | -0.18324 | 0.76216
| -0.63085 | -0.82212 | 0.30994 | 1.1699
| 0.65631 | -0.55715 | 0.66129 | 1.1785
| -0.5551 | 0.72231 | -0.42013 | 1.0064
| 0.090455 | -2.0856 | 0.40028 | 4.5179
| -0.70068 | 0.88238 | -0.15404 | 1.2933
| 1.1883 | -0.91695 | 1.2665 | 3.857
| 0.25021 | -2.872 | -0.22996 | 8.364
| -0.6735 | 0.12837 | 1.6834 | 3.3041
| -0.59837 | 1.6282 | -1.4768 | 5.19
| -2.9679 | -1.4136 | 0.65478 | 11.2354
| 0.59349 | -1.043 | -0.0046934 | 1.4402
| -1.3229 | 0.9232 | -0.19309 | 2.6395
| -0.93487 | 0.8332 | -0.89793 | 2.3745
| -0.89356 | -0.22672 | -0.83051 | 1.5396
| 1.4067 | 0.86767 | -0.83419 | 3.4276
| -0.68655 | 1.3077 | 0.90317 | 2.9971
| 0.85737 | -0.29734 | 0.95013 | 1.7262
| -0.7242 | -1.3471 | 1.1101 | 3.5715
| -1.083 | 1.1575 | 0.82006 | 3.1851
| 3.0718 | 1.979 | -0.18255 | 13.3859
| -1.5221 | 0.028709 | -0.97578 | 3.2699
| -1.1293 | 0.25102 | 1.7909 | 4.5456
| 1.2135 | -0.79229 | -0.28776 | 2.1832
| 0.88792 | 0.59964 | -0.71576 | 1.6603
| -0.33536 | -0.20563 | 1.7302 | 3.1485
| -0.78744 | 0.73756 | -2.4097 | 6.9707
| 0.5815 | 1.2818 | 1.4104 | 3.9704
| 0.75927 | -0.58472 | -0.52923 | 1.1985
| 0.22107 | 0.14747 | 0.47007 | 0.29158
| -1.3857 | 1.1277 | -0.13246 | 3.2097
| 2.2227 | 0.13108 | -0.63008 | 5.3544
| 0.27267 | 0.44178 | 0.53131 | 0.55181
| 1.1575 | -0.10906 | 0.81128 | 2.0098
| 0.17229 | -0.75484 | 1.0336 | 1.6679
| -0.78905 | -1.38 | 2.8844 | 10.8466
| 0.45852 | -1.2959 | -1.4546 | 4.0054
| -1.6723 | -0.923 | -0.9326 | 4.5183
| -0.42069 | 2.0969 | 0.042157 | 4.5759
| 1.5923 | 0.38222 | 0.085753 | 2.6889
| -2.1352 | -1.7372 | -0.21997 | 7.6251
| -0.66528 | 1.4832 | 1.2285 | 4.1518
| 1.9809 | 0.49215 | -1.6257 | 6.8092
| 1.528 | 1.3798 | 0.24936 | 4.3009
| -0.23356 | -1.1191 | 0.28122 | 1.3861
| 0.11273 | 0.71815 | -0.076646 | 0.53432
| 0.13578 | -0.4711 | 0.59316 | 0.59222
| -1.9024 | 1.5704 | 0.34378 | 6.2035
| -0.50623 | -0.84374 | -0.41789 | 1.1428
| 1.8445 | 0.31484 | -0.9609 | 4.4248
| 0.065773 | 0.022103 | -0.96052 | 0.92741
| 1.6788 | 1.0689 | -0.10357 | 3.9717
| 0.70045 | -0.82993 | -0.40051 | 1.3398
| -0.032225 | -1.4604 | 1.6846 | 4.9715
| -1.5387 | 1.3084 | -1.3251 | 5.8354
| -1.1906 | 0.19339 | 0.90046 | 2.2658
| 0.10071 | 0.72147 | 0.11745 | 0.54445
| 0.061942 | -0.37558 | 0.71331 | 0.6537
| -1.0209 | 0.26488 | -0.59341 | 1.4645
| 0.0055303 | 0.045755 | -0.59458 | 0.35565
| 1.8205 | -0.39064 | 1.2551 | 5.0422
| -1.0706 | -0.082508 | -0.5134 | 1.4166
| 0.88072 | 0.92481 | 0.62543 | 2.0221
| -1.8037 | -0.63814 | 2.1648 | 8.3467
| 0.17404 | -0.26458 | 0.91677 | 0.94076
| -0.53927 | 3.3647 | -2.7964 | 19.4318
| -0.67982 | 0.059878 | 0.56221 | 0.78183
| -0.33004 | -0.26137 | 1.0804 | 1.3445
| 1.4698 | -0.028929 | -1.0859 | 3.3404
| 0.67152 | -1.0284 | 0.01616 | 1.5087
| 1.295 | 0.63176 | 0.58229 | 2.4152
| -2.7255 | -1.4069 | 0.81332 | 10.0693
| -0.51465 | -1.0938 | -0.55483 | 1.769
| -0.55486 | -0.49445 | 0.81449 | 1.2157
| -0.13944 | 2.4678 | -0.58532 | 6.4522
| 0.72024 | -0.71934 | -1.3368 | 2.8232
| -0.18666 | 1.4727 | -0.032949 | 2.2047
| -0.77331 | -1.1234 | -0.84971 | 2.5821
| -0.66851 | -1.3374 | 1.5672 | 4.6919
| -0.0052125 | 0.86402 | 0.39898 | 0.90574
| -1.04 | -1.3742 | -0.54649 | 3.2686
| -0.8979 | -0.40915 | 0.016861 | 0.97392
| 0.10689 | 0.98869 | -0.17766 | 1.0205
| -0.10462 | -0.79006 | 0.71799 | 1.1507
| -0.081741 | -0.68907 | 1.2604 | 2.0701
| 0.79736 | 0.59544 | 0.1657 | 1.0178
| 0.37941 | 0.3367 | 0.29669 | 0.34534
| 0.44664 | -0.027722 | 1.8283 | 3.5428
| -1.6397 | -1.1656 | -0.42662 | 4.229
| -0.89216 | 0.0010475 | 0.51783 | 1.0641
| -1.8077 | -1.6592 | 0.51408 | 6.2853
| 1.8639 | -1.0778 | -0.26333 | 4.7051
| 0.73464 | -1.1271 | -0.16042 | 1.8357
| -0.9986 | -0.44844 | -1.5316 | 3.5441
| 1.128 | -1.0631 | 1.4143 | 4.4029
| -1.2711 | 0.27343 | 0.080484 | 1.6968
| -0.17703 | 0.83888 | -1.8314 | 4.089
| 1.2267 | 0.24847 | -0.64551 | 1.9833
| -1.9106 | 1.83 | -0.21325 | 7.0449
| -1.231 | -1.4354 | 0.19481 | 3.6136
| 0.24006 | 0.36149 | 0.51825 | 0.45689
| -0.61652 | 0.0047505 | -0.69818 | 0.86757
| 0.22263 | 0.70299 | 0.070833 | 0.54878
| -0.10862 | 0.43396 | -1.0448 | 1.2918
| 0.11112 | -0.94477 | -0.59181 | 1.2552
| -1.1922 | -1.6882 | -0.72116 | 4.7914
| 1.1285 | 1.8503 | -0.58157 | 5.0353
| 0.052005 | -1.6878 | 0.19085 | 2.8877
| -0.42148 | -0.68884 | -1.0774 | 1.8128
| 1.8827 | -0.96233 | 1.2749 | 6.096
| 0.40201 | -0.3166 | -0.55915 | 0.5745
| 1.6812 | -0.10695 | 0.035985 | 2.8393
| -0.23108 | -0.33091 | 0.47165 | 0.38535
| 0.38739 | -1.7559 | -1.0763 | 4.3916
| 0.2035 | 0.36383 | 0.39467 | 0.32954
| 0.56874 | 0.82539 | 0.66337 | 1.4448
| -1.0459 | 0.48996 | 0.21735 | 1.3812
| -0.046602 | 0.69203 | -0.44503 | 0.67913
| -0.24636 | -0.15552 | 1.1749 | 1.4652
| -0.092079 | -1.6344 | -0.56244 | 2.9961
| -0.79151 | 2.0673 | 1.5888 | 7.4246
| 0.21076 | -0.061981 | 0.37927 | 0.19211
| -0.78856 | -0.13313 | -0.40999 | 0.80764
| 1.1172 | 0.30368 | -0.28818 | 1.4233
| -1.554 | 0.068901 | 2.5524 | 8.9343
| -0.67476 | -0.77214 | -0.78306 | 1.6647
| -0.13167 | 0.85952 | -0.22131 | 0.80509
| -1.2633 | 0.0014145 | 0.69225 | 2.0751
| 1.4396 | -0.38605 | -0.76897 | 2.8129
| -0.059655 | 0.31927 | 0.0081569 | 0.10556
| -0.74849 | 1.0381 | -1.6022 | 4.2049
| 0.42898 | -0.67929 | 0.57728 | 0.97871
| -0.2804 | -0.35738 | 0.2045 | 0.24816
| 0.77255 | 2.5523 | -1.4607 | 9.2449
| 0.64297 | -0.57865 | 0.17865 | 0.78016
| -0.12166 | 0.43025 | 0.37313 | 0.33914
| -0.98437 | 0.67102 | 0.9807 | 2.381
| -0.21671 | 0.71237 | 0.25441 | 0.61916
| -0.32836 | -2.5549 | -1.825 | 9.9659
| 0.68212 | -0.3617 | 1.8559 | 4.0406
| -1.0148 | -1.3834 | 0.86482 | 3.6917
| 0.38247 | -0.8197 | 0.31248 | 0.91583
| 0.54906 | -2.3404 | 2.0908 | 10.1505
| 1.5014 | 0.56527 | -0.49848 | 2.8223
| -0.30941 | 0.90146 | -1.0789 | 2.0723
| -0.6062 | -0.69398 | 1.0474 | 1.9461
| -0.41727 | 0.63927 | -0.28596 | 0.66456
| 1.099 | -0.50046 | 1.1794 | 2.8492
| -0.94459 | 0.74956 | -0.34507 | 1.5732
| -0.090646 | -1.4831 | -2.6308 | 9.1292
| -0.11199 | -0.2123 | 0.10467 | 0.06857
| 1.1236 | -0.40134 | -0.48769 | 1.6613
| 0.18199 | -0.87275 | 1.6333 | 3.4624
| 0.34403 | -0.67354 | 1.0891 | 1.7581
| 0.28392 | 1.7889 | -0.12111 | 3.2953
| -0.1135 | -0.28395 | -0.60477 | 0.45926
| -0.39966 | -0.68884 | -0.73714 | 1.1776
| -2.1972 | 0.38901 | -2.2833 | 10.1925
| 0.5624 | -0.30086 | 1.662 | 3.1689
| -0.6907 | 0.51446 | 1.3694 | 2.617
| 0.64364 | -1.5121 | 0.0072577 | 2.7007
| -1.1809 | -0.74948 | 0.71649 | 2.4695
| -0.091105 | 0.66333 | -1.2017 | 1.8924
| -0.56892 | -0.79537 | -1.3769 | 2.852
| -0.63015 | 0.99439 | -1.5935 | 3.9251
| 2.4389 | -0.13622 | -0.26767 | 6.0383
| 0.40107 | 1.0898 | 0.62761 | 1.7425
| -1.6272 | 1.7933 | 0.29234 | 5.9489
| 0.41713 | -0.69604 | 3.2685 | 11.3415
| -0.49058 | 1.2037 | 1.5821 | 4.1926
| -0.39556 | -0.17192 | -0.27382 | 0.261
| 1.0136 | 0.76347 | -0.071545 | 1.6153
| -0.26013 | 1.9795 | -0.068898 | 3.9908
| 0.10275 | -0.58468 | -1.1846 | 1.7557
| 0.73112 | 0.19606 | 0.26179 | 0.64151
| -0.51404 | 0.37353 | 0.93206 | 1.2725
| 0.58203 | -0.28701 | -1.0877 | 1.6042
| -1.0489 | -0.33204 | -1.5925 | 3.7463
| 0.30564 | -0.52578 | -0.7165 | 0.88323
| 0.50111 | 2.0143 | -1.2104 | 5.7735
| -0.51074 | -1.5918 | 0.22751 | 2.8465
| 0.19147 | -1.4404 | 0.68555 | 2.5815
| 0.10139 | -1.4281 | -1.3457 | 3.8608
| -0.86919 | 0.32834 | 1.0287 | 1.9216
| -0.53738 | 0.83662 | 0.27456 | 1.0641
| 1.3071 | 0.87952 | 1.9931 | 6.4544
| 0.82003 | -2.487 | 1.0387 | 7.9366
| 2.5925 | 0.36334 | -1.8795 | 10.3858
| 0.67535 | -0.40009 | -0.30719 | 0.71053
| -0.051156 | 0.15873 | 0.79324 | 0.65705
| 0.00022122 | 0.92765 | 0.010047 | 0.86063
| -0.33577 | -0.17138 | 0.98917 | 1.1206
| -0.35329 | 0.056501 | 1.2373 | 1.659
| -0.60762 | 0.4948 | -1.7123 | 3.546
| -0.13376 | 0.99096 | -0.18065 | 1.0325
| 0.033218 | 0.41511 | 1.4076 | 2.1549
| 0.43152 | 2.0733 | 0.58029 | 4.8214
| -0.59312 | 0.90635 | -0.61302 | 1.5491
| -1.916 | -0.49496 | -1.5073 | 6.1881
| -0.62403 | -2.2417 | 0.32052 | 5.5175
| 0.64627 | -0.70959 | -0.47019 | 1.1423
| 0.67326 | 0.69551 | -0.082707 | 0.94385
| -1.4915 | -0.058107 | 0.31087 | 2.3246
| 0.7884 | 0.45313 | 0.9855 | 1.7981
| 0.16665 | 0.26989 | -1.5612 | 2.5379
| -0.16012 | 1.6544 | -0.87044 | 3.5202
| 1.5123 | -0.5916 | 0.68317 | 3.1038
| -1.1963 | 1.6423 | 2.484 | 10.2987
| -0.49764 | 0.42137 | -0.30325 | 0.51716
| 0.38917 | -0.63715 | 0.1646 | 0.5845
| 0.22325 | -1.917 | -0.59424 | 4.078
| 0.35654 | 1.3449 | -1.9654 | 5.7985
| -0.17719 | 1.4836 | 0.5879 | 2.5782
| -0.1795 | 1.1419 | 1.1516 | 2.6623
| -1.1446 | 1.7261 | -1.584 | 6.7986
| -0.97721 | -1.6326 | 1.2839 | 5.2687
| -0.45852 | 1.0529 | 0.013175 | 1.319
| -0.13717 | -1.0356 | -0.79146 | 1.7177
| 0.94819 | -0.051574 | 1.1715 | 2.2742
| -0.71141 | -0.3046 | -0.09521 | 0.60794
| -1.4648 | 0.27715 | 0.4135 | 2.3934
| 1.6595 | 0.40893 | 2.2265 | 7.8785
| 0.51689 | -1.2094 | -0.46139 | 1.9428
| 0.78166 | 0.68871 | -0.0021814 | 1.0853
| -0.93006 | 0.061381 | 0.97181 | 1.8132
| 1.1588 | -0.82985 | 0.052028 | 2.0341
| -0.77513 | -2.3499 | -0.60371 | 6.4874
| -1.2313 | -0.51903 | -1.8073 | 5.0517
| 1.1367 | -2.4238 | -0.53598 | 7.4541
| 0.19633 | 0.11355 | -1.7511 | 3.1177
| -2.6898 | -0.36172 | 0.48354 | 7.5995
| -1.1864 | -0.92219 | 1.7927 | 5.4717
| -0.48942 | -0.73837 | 0.91626 | 1.6242
| -0.89867 | -1.276 | -0.13151 | 2.4532
| 0.67238 | 0.031484 | 0.65739 | 0.88525
| -0.65707 | -2.0992 | -0.88829 | 5.6273
| -1.638 | -0.39718 | -2.2589 | 7.9435
| 0.74421 | 0.48942 | 0.52161 | 1.0655
| 0.56089 | 0.46045 | 0.24488 | 0.58659
| -0.54911 | 1.6042 | -1.3974 | 4.8278
| 0.78658 | 1.8867 | -1.2866 | 5.8339
| -0.49235 | -1.5046 | -0.67856 | 2.9668
| -0.36791 | 1.3149 | -1.0577 | 2.983
| -0.90849 | 0.36388 | -0.7842 | 1.5727
| 0.37275 | -0.2772 | 1.3509 | 2.0407
| 0.82536 | 0.11771 | -1.2977 | 2.3791
| 0.0099057 | 0.31936 | -0.11171 | 0.11456
| -0.46587 | -0.087472 | -1.6929 | 3.0905
| 0.76256 | 0.88394 | -0.52913 | 1.6428
| 1.0728 | -0.66834 | 0.55928 | 1.9104
| -0.59335 | 1.7558 | 1.1866 | 4.8428
| 0.75369 | -0.55614 | 0.31704 | 0.97786
| -0.76523 | -0.16061 | -0.48774 | 0.84925
| -0.05998 | 0.037727 | 0.44236 | 0.20071
| 1.063 | 1.5297 | -0.67955 | 3.9318
| -0.38994 | 1.6951 | -1.1487 | 4.3449
| -2.3802 | -0.13808 | 1.3321 | 7.4588
| -1.4386 | -1.3994 | -0.18195 | 4.0608
| 1.0278 | 0.45317 | -0.55023 | 1.5645
| 0.52451 | 0.9097 | 1.6163 | 3.7152
| -0.087762 | -1.5931 | 1.0882 | 3.73
| 0.33708 | -1.053 | 2.0792 | 5.5455
| -1.1211 | -0.44058 | -0.89521 | 2.2523
| -0.94221 | 1.3023 | -0.58529 | 2.9264
| 0.16841 | 1.1515 | -0.092022 | 1.3629
| 0.90688 | -0.50343 | -0.96138 | 2.0001
| 0.91186 | 2.1219 | 0.69679 | 5.8195
| -1.4776 | -2.3533 | -0.26094 | 7.7896
| -0.24858 | 1.4597 | -0.56425 | 2.5109
| 0.16031 | 1.0979 | -0.20109 | 1.2715
| -1.0111 | -1.1 | -1.0538 | 3.3427
| 0.88518 | 1.6379 | -0.96768 | 4.4026
| -0.18048 | 1.2321 | 0.56711 | 1.8723
| 0.020392 | -0.32149 | 0.86732 | 0.85602
| -1.7639 | -0.81993 | -0.32212 | 3.8874
| -0.82317 | -1.077 | 0.26152 | 1.906
| -0.6435 | 0.40301 | 0.42487 | 0.75702
| 0.35037 | 0.72794 | 0.8681 | 1.4063
| -0.91654 | -0.88355 | 0.29326 | 1.7067
| -0.14683 | 0.70469 | 0.87896 | 1.2907
| 0.44094 | 0.6478 | -1.1165 | 1.8606
| -1.1383 | -0.029022 | -0.00037401 | 1.2965
| 0.27301 | -0.7778 | 0.49657 | 0.92608
| -1.1726 | 0.45143 | -0.32546 | 1.6847
| 0.20395 | 0.57416 | -1.1593 | 1.7153
| 0.37441 | 0.094879 | -1.3321 | 1.9237
| -0.45706 | 2.0509 | -0.809 | 5.0695
| -0.1477 | 2.3446 | -0.056217 | 5.5222
| -1.2698 | -0.93926 | -0.29943 | 2.5841
| 0.15505 | 1.7232 | 0.67774 | 3.4529
| -0.11109 | -0.49955 | 0.59625 | 0.61741
| -0.22811 | 0.12826 | -0.50518 | 0.32369
| 1.9099 | 1.0976 | 1.5225 | 7.1706
| 0.10725 | 0.33051 | 1.8269 | 3.4582
| -0.19723 | -0.13961 | 1.0651 | 1.1929
| 1.2975 | -1.9074 | 0.82878 | 6.0083
| 0.072358 | -0.51916 | 0.12295 | 0.28988
| 1.5194 | 0.96233 | -0.34675 | 3.3548
| -0.90999 | 1.1637 | -0.055355 | 2.1854
| -0.51495 | -0.39739 | 0.36651 | 0.55743
| 0.33692 | -0.45803 | 0.52851 | 0.60263
| -1.4867 | 1.2218 | 2.0307 | 7.827
| 0.17754 | 1.8117 | 0.27145 | 3.3875
| 1.4867 | -0.47102 | 1.1777 | 3.8192
| 0.82491 | 0.73206 | 0.80502 | 1.8644
| 0.58049 | -0.20877 | 0.52043 | 0.6514
| 0.62752 | 0.36012 | -0.64792 | 0.94326
| 0.56182 | -0.43192 | 0.12523 | 0.51787
| 0.15375 | 0.57508 | -1.2738 | 1.9769
| -1.6455 | 0.46598 | -0.33825 | 3.0391
| 3.2266 | -1.0838 | 0.98976 | 12.5651
| -0.10308 | 1.2528 | -0.19341 | 1.6175
| 0.49037 | -0.41355 | -2.3964 | 6.1543
| 0.57349 | -0.22612 | -0.24566 | 0.44038
| -0.030927 | 0.14511 | -0.072706 | 0.027298
| 0.62935 | -1.5172 | 0.30344 | 2.7901
| -0.71269 | -0.17361 | 0.97782 | 1.4942
| -0.54352 | 0.53217 | 0.042566 | 0.58043
| -0.06928 | -1.1855 | 0.66675 | 1.8549
| -0.55053 | -1.1102 | 0.31501 | 1.635
| 0.27957 | 0.93686 | -0.97614 | 1.9087
| 0.5185 | 0.6178 | 0.25731 | 0.71673
| -1.3392 | -0.30451 | 1.7537 | 4.9619
| 0.27132 | 1.7047 | -2.2398 | 7.9963
| 0.26582 | 0.7482 | 1.5808 | 3.1294
| -0.44514 | 0.62969 | -1.4235 | 2.621
| -1.0045 | -0.55843 | 0.42763 | 1.5037
| -0.28982 | -0.68297 | 2.0059 | 4.5742
| -0.44575 | 0.57255 | 0.36504 | 0.65976
| 1.1307 | 0.46405 | 0.87157 | 2.2534
| -1.7535 | 0.58362 | 0.54893 | 3.7167
| -1.2437 | 1.2494 | -0.87773 | 3.8782
| -0.16695 | -1.5534 | 1.3245 | 4.1952
| 0.62672 | -0.87764 | -0.49673 | 1.4098
| 0.84355 | -0.66353 | -0.96411 | 2.0814
| -0.32656 | -0.28881 | 0.35652 | 0.31716
| 0.048219 | 1.0223 | 0.13354 | 1.0653
| -1.037 | 0.35041 | -0.6079 | 1.5676
| 0.43791 | 0.60599 | 0.36936 | 0.69541
| -0.14619 | 1.0684 | -1.0908 | 2.3525
| -1.2627 | 2.8033 | -0.08167 | 9.4597
| -0.91569 | 0.92481 | 0.3309 | 1.8033
| 2.6916 | 1.9984 | -1.1804 | 12.6315
| 1.9498 | 1.2232 | 0.24621 | 5.3586
| 0.61429 | -1.0811 | -1.6778 | 4.3612
| -0.16916 | 0.89318 | 0.9815 | 1.7897
| -1.5343 | 0.55783 | 0.77168 | 3.2606
| 0.1136 | -1.6803 | 0.52036 | 3.1072
| 0.685 | 0.23199 | 0.79102 | 1.1488
| 2.0085 | -0.69047 | -0.93049 | 5.3765
| -0.78098 | 0.61616 | -0.44876 | 1.191
| -0.13067 | 2.2348 | 0.12302 | 5.0264
| -1.511 | 0.9574 | 1.1265 | 4.4687
| -0.94517 | -1.1735 | -1.1385 | 3.5665
| 0.26099 | 1.9547 | -0.26365 | 3.9583
| -0.64695 | -1.1552 | -0.74448 | 2.3073
| -0.14458 | 0.23609 | -0.32661 | 0.18332
| -1.423 | -0.079612 | 1.6847 | 4.8696
| -0.38488 | 0.33148 | 0.4491 | 0.45971
| -0.1328 | -0.11689 | -1.4782 | 2.2164
| 0.58298 | -0.73338 | -0.10477 | 0.88869
| -0.14949 | -0.38523 | -1.9818 | 4.0983
| 0.133 | 0.58204 | -2.1986 | 5.1902
| -1.6238 | 0.95629 | 0.39056 | 3.7038
| 1.6948 | -2.4326 | 0.11953 | 8.8043
| -1.0893 | 1.0474 | 0.054713 | 2.2865
| -0.19206 | -0.071197 | -1.6864 | 2.8858
| 0.11766 | 0.31947 | -0.13102 | 0.13307
| -0.81943 | -1.1039 | -0.8422 | 2.5993
| 1.5821 | 0.21002 | -0.20736 | 2.5901
| 0.06355 | 2.0908 | -0.88077 | 5.1513
| 2.4549 | 1.5308 | -0.074588 | 8.3755
| -1.4974 | 0.26259 | -0.78718 | 2.9309
| -0.0069869 | -1.0541 | -0.62847 | 1.5062
| 2.2416 | 0.36577 | -1.5133 | 7.4487
| -1.19 | -2.3971 | 0.15045 | 7.1848
| 0.87264 | -0.95581 | 0.42075 | 1.8521
| 0.85552 | -0.89523 | -0.053971 | 1.5363
| -0.38144 | -0.9995 | 1.7432 | 4.1832
| -1.0138 | 2.7382 | 0.069996 | 8.5303
| -0.64024 | -0.29955 | -0.25532 | 0.56482
| -0.076251 | -0.64964 | 1.2647 | 2.0274
| 0.39897 | -0.85495 | 0.68693 | 1.362
| -0.21803 | -0.79482 | 0.13423 | 0.69729
| 0.41891 | -0.68937 | -0.58473 | 0.99263
| 0.55606 | 0.034064 | -0.019967 | 0.31076
| 0.54763 | -1.0534 | 0.48406 | 1.6438
| 0.61985 | -0.72562 | 1.1608 | 2.2583
| 0.66806 | -0.08226 | 0.89739 | 1.2584
| -0.93548 | 0.32759 | -1.023 | 2.029
| 0.32485 | -0.46274 | 1.9599 | 4.1608
| 0.18864 | -1.262 | -0.37045 | 1.7656
| 0.31555 | -0.33867 | -0.034585 | 0.21547
| -0.89489 | 1.2116 | 1.7728 | 5.4119
| 0.66251 | -0.098232 | -0.8466 | 1.1653
| -0.77358 | 1.5664 | 1.0102 | 4.0727
| -0.013683 | 1.6872 | 1.2012 | 4.2898
| -0.24352 | -0.4199 | -0.66017 | 0.67143
| -0.43173 | 1.2168 | 0.26582 | 1.7377
| 1.9589 | 0.1421 | 0.77541 | 4.4589
| -0.10323 | -0.44672 | -0.86006 | 0.94992
| -2.8786 | -0.41684 | -1.6971 | 11.3401
| -0.098404 | 0.55708 | 0.02536 | 0.32067
| -0.39079 | -0.96815 | 0.22916 | 1.1426
| 0.28465 | -0.48374 | -0.52535 | 0.59102
| -0.11183 | 0.28966 | -1.3089 | 1.8096
| 0.67027 | 0.087583 | 0.66104 | 0.89391
| -0.053717 | 0.16907 | -1.1179 | 1.2812
| 0.82364 | -2.2207 | -0.95194 | 6.5162
| 0.19014 | -1.2084 | -0.92353 | 2.3492
| 0.33705 | 0.80651 | -2.3082 | 6.092
| 0.91352 | 0.39987 | 0.50227 | 1.2467
| 0.18328 | -0.94961 | 0.069943 | 0.94024
| 0.86859 | -0.050748 | 1.0261 | 1.81
| 0.66595 | -0.64066 | 0.73778 | 1.3983
| 1.7551 | -0.40657 | 0.36251 | 3.377
| -1.7316 | -0.90714 | -1.3528 | 5.6514
| 1.4504 | 0.25997 | -0.43801 | 2.3632
| -0.73635 | 0.96993 | 0.41129 | 1.6521
| 0.24153 | 1.9424 | -1.2507 | 5.3956
| 0.29318 | -1.2153 | -0.055051 | 1.566
| 0.33485 | -0.59356 | -0.35493 | 0.59042
| -2.3114 | 0.92832 | 0.12787 | 6.2207
| -1.0026 | 0.25311 | 1.9113 | 4.7224
| -0.18216 | 0.85 | -1.0391 | 1.8355
| -0.064627 | -0.55364 | 0.73164 | 0.84598
| 0.46545 | 1.729 | 1.4877 | 5.4194
| -0.34856 | -0.94189 | -0.19148 | 1.0453
| -1.464 | -1.3675 | -1.2419 | 5.5556
| 1.1001 | 0.86574 | -0.74655 | 2.5171
| -0.96168 | -0.68556 | -0.81327 | 2.0562
| -0.36127 | -0.13645 | 0.17026 | 0.17813
| 0.42952 | 1.4201 | -1.2705 | 3.8154
| 0.20664 | 0.091134 | -0.20836 | 0.094421
| 0.39298 | -1.0343 | 0.93046 | 2.0899
| 0.44874 | -0.15217 | 1.0318 | 1.2891
| 0.34022 | 0.43606 | 1.2506 | 1.8699
| 1.3529 | 0.90664 | 1.0555 | 3.7665
| -0.96437 | -0.69292 | -1.3782 | 3.3095
| 1.7172 | -0.12075 | 0.49502 | 3.2085
| 0.0061592 | 0.6343 | -0.13318 | 0.42011
| -0.62101 | 1.1714 | -0.94123 | 2.6437
| -0.59579 | -0.97042 | 0.92439 | 2.1512
| -0.61839 | -1.0185 | 1.2652 | 3.0203
| -0.41838 | 0.58872 | 1.4121 | 2.5157
| -0.98215 | -0.68306 | 0.3262 | 1.5376
| -2.0542 | 0.48591 | 0.16586 | 4.4832
| 0.83738 | -1.614 | -0.40104 | 3.467
| 0.10171 | 0.24723 | -1.21 | 1.5355
| -0.11213 | 0.59712 | -1.2827 | 2.0144
| 0.6272 | 0.54609 | -0.74022 | 1.2395
| -0.75016 | -0.68318 | 0.3489 | 1.1512
| -0.34352 | -0.68098 | 0.075223 | 0.5874
| 0.96043 | 0.85195 | 0.54856 | 1.9492
| 0.12204 | -0.70818 | -0.70312 | 1.0108
| -0.066147 | -0.10817 | 1.3197 | 1.7577
| 0.87459 | -0.3259 | -0.14927 | 0.8934
| -1.7573 | 0.21473 | -0.52415 | 3.4089
| -0.042564 | -0.80204 | -2.4009 | 6.4096
| -0.01675 | 0.27065 | 0.66494 | 0.51568
| -0.44519 | 0.60188 | -0.6319 | 0.95975
| 0.62183 | 0.64782 | -1.1324 | 2.0887
| 0.93737 | -1.1063 | -0.020797 | 2.1029
| -0.60511 | -1.2709 | -1.3351 | 3.764
| -0.62203 | 0.57676 | 0.25704 | 0.78564
| -1.5436 | -0.041066 | -1.1974 | 3.8184
| 0.42738 | -0.42775 | -0.66455 | 0.80725
| 0.98612 | 0.99166 | 0.83103 | 2.6464
| 1.2433 | 0.80576 | -0.07992 | 2.2015
| -1.3306 | -1.5125 | 0.13163 | 4.0754
| 1.0479 | -0.68034 | -1.1457 | 2.8735
| 0.58992 | -0.36976 | -0.075779 | 0.49047
| -1.0638 | -0.83592 | -1.0005 | 2.8315
| -1.3229 | -0.59191 | 0.81598 | 2.7662
| 2.1229 | 1.8255 | -0.12293 | 7.8543
| 0.57449 | -0.31437 | 0.24042 | 0.48667
| 0.85266 | 0.26407 | 0.26209 | 0.86545
| -0.16834 | -0.8613 | -0.69922 | 1.2591
| -0.94223 | -0.80292 | 1.0973 | 2.7365
| 0.0022414 | -0.65527 | 0.24509 | 0.48945
| -0.34818 | -1.7407 | 0.11696 | 3.1651
| 0.056154 | 0.41531 | -0.15712 | 0.20032
| 0.20578 | 0.95629 | -0.38517 | 1.1052
| -1.2308 | 0.63843 | 1.0311 | 2.9856
| 0.544 | 1.1601 | -0.735 | 2.1819
| -0.068357 | -1.1253 | 1.6898 | 4.1264
| -0.49654 | -0.92941 | -1.6658 | 3.8854
| -0.44604 | 0.013536 | -0.18878 | 0.23477
| -0.61664 | -0.41253 | -0.70283 | 1.0444
| 0.20606 | -0.21136 | 1.0552 | 1.2007
| -1.0018 | 1.3339 | -0.98922 | 3.7613
| 2.6565 | 0.096776 | 0.28943 | 7.1504
| -0.71688 | 1.5402 | 1.3512 | 4.7116
| -0.39844 | 1.3355 | -0.65465 | 2.3709
| -1.7009 | 0.28277 | -0.80065 | 3.614
| 0.060388 | -1.0087 | 1.6968 | 3.9001
| 0.034456 | 0.83952 | 0.47623 | 0.93278
| -0.77639 | 0.60252 | 0.99743 | 1.9607
| 0.43723 | -0.09103 | -0.63449 | 0.60203
| -1.2155 | 0.022157 | -0.19081 | 1.5142
| 0.20959 | -0.23655 | 0.29368 | 0.18613
| -1.8795 | 0.98784 | -1.9254 | 8.2155
| 2.9361 | 0.27073 | -0.40729 | 8.8599
| -0.1633 | -0.28912 | 0.0099389 | 0.11036
| -0.11316 | 0.27796 | 0.74299 | 0.64211
| 0.3718 | -0.44419 | -0.88516 | 1.1191
| -0.5988 | -0.33309 | 2.0409 | 4.6348
| 0.097148 | 1.3895 | 1.1816 | 3.3364
| 1.8826 | -0.18803 | -0.73593 | 4.1211
| -1.5493 | -1.7365 | -0.80534 | 6.064
| 0.2423 | 1.1067 | 0.64832 | 1.7038
| -0.13797 | -3.4191 | 1.3882 | 13.6359
| -0.12311 | 0.49214 | 1.1218 | 1.5158
| -1.2133 | 0.31025 | 1.6419 | 4.2642
| -0.98401 | 0.21602 | -1.1836 | 2.4159
| 0.053552 | -1.8984 | 2.4927 | 9.8202
| 0.93725 | -0.57478 | 0.25899 | 1.2759
| -0.89993 | -1.1304 | -0.070137 | 2.0925
| -0.17233 | 1.0484 | -0.52902 | 1.4087
| -0.67511 | 0.28439 | 0.13026 | 0.55362
| 1.7606 | 1.153 | 1.1884 | 5.8415
| -0.47354 | -0.23911 | 0.34742 | 0.40211
| -1.7539 | 2.6251 | 0.28031 | 10.0456
| -1.4241 | -0.39984 | -0.93581 | 3.0636
| -2.313 | 0.51602 | -0.067088 | 5.6207
| 0.5001 | -0.9418 | -1.0458 | 2.2307
| -0.21686 | -1.568 | -0.061117 | 2.5092
| 1.1042 | 0.5754 | -1.4963 | 3.7892
| 0.86889 | -1.6305 | -1.9205 | 7.102
| 2.6032 | 1.2447 | 0.65905 | 8.7604
| -0.054878 | -0.13415 | 0.87003 | 0.77796
| 0.27815 | 0.45913 | -0.6689 | 0.7356
| 2.0848 | -0.065056 | 1.9441 | 8.1299
| 0.62039 | 0.014468 | 0.51614 | 0.65149
| 0.58618 | 0.51376 | 0.46264 | 0.82158
| -1.7067 | -0.78342 | -0.70108 | 4.0182
| 0.94162 | -0.67079 | -1.6747 | 4.1411
| 1.4983 | -1.5507 | 0.24431 | 4.7094
| 0.64205 | -0.52619 | -0.5555 | 0.99769
| -2.1468 | -0.55888 | -0.39751 | 5.0791
| -1.0543 | 1.0745 | -1.0301 | 3.3274
| -0.96784 | 2.5266 | 0.53903 | 7.6108
| 1.3228 | 1.3212 | 2.4205 | 9.3542
| 0.083494 | 0.45246 | 0.18335 | 0.24531
| 0.069015 | 0.83886 | 1.6374 | 3.3894
| -1.6355 | 1.0251 | -1.5142 | 6.0188
| -1.9358 | -1.4805 | -0.602 | 6.3016
| 1.0714 | 0.19421 | -0.66106 | 1.6226
| 0.6679 | -1.0887 | -0.036678 | 1.6327
| 0.0090939 | 1.3094 | 2.6953 | 8.9792
| 0.48852 | 0.34315 | -0.016633 | 0.35668
| -0.047943 | 0.20954 | -0.17102 | 0.075452
| -0.45206 | 0.30112 | -0.90356 | 1.1115
| -0.85753 | 2.9985 | 0.43334 | 9.9142
| 1.5315 | -1.1341 | -0.8286 | 4.3182
| -0.25025 | -0.34562 | 0.19149 | 0.21875
| 1.3575 | 1.3584 | -0.90566 | 4.5082
| -2.0593 | -2.3056 | -0.49349 | 9.8
| -0.25007 | 0.49292 | 0.27576 | 0.38154
| -1.1482 | -0.44409 | 0.80411 | 2.1621
| 0.05696 | -0.77641 | -1.0469 | 1.702
| 1.5203 | 1.9678 | 1.0193 | 7.2224
| 1.1302 | 0.053643 | -0.70124 | 1.7721
| 1.1185 | 0.53604 | -1.2052 | 2.9908
| -2.3495 | -0.68572 | 0.36286 | 6.1219
| -0.4335 | 0.66269 | 0.027083 | 0.62781
| -1.1373 | 0.18573 | 0.17489 | 1.3585
| 0.68427 | 0.96459 | 0.5604 | 1.7127
| -0.060691 | 0.2913 | -0.29828 | 0.17751
| -1.6091 | 0.090466 | -0.2767 | 2.6741
| -0.37595 | -0.40024 | -0.98906 | 1.2798
| -0.86099 | 0.21492 | 1.3784 | 2.6876
| -1.0193 | -0.016703 | -0.16686 | 1.0671
| 0.39791 | -0.48573 | 1.715 | 3.3356
| 0.77064 | 0.89367 | 0.30971 | 1.4884
| 1.6719 | 0.58075 | 0.40398 | 3.2957
| -0.57304 | 1.458 | 0.62583 | 2.8458
| -2.5951 | -0.73279 | 1.6559 | 10.0136
| 0.52449 | 0.012655 | -0.35122 | 0.39861
| -0.4043 | -0.21022 | -1.3217 | 1.9546
| 0.93714 | 1.0836 | -0.38171 | 2.1982
| -0.52114 | -1.5839 | -1.6353 | 5.4545
| -0.086281 | -0.13384 | 1.8443 | 3.4267
| 2.1263 | 0.26435 | -0.55464 | 4.8985
| -1.2355 | -0.60394 | -0.024392 | 1.8918
| -1.351 | -0.42894 | -0.27615 | 2.0856
| 0.16869 | -0.43577 | 0.63101 | 0.61653
| -0.73147 | 1.3962 | 0.68525 | 2.9539
| 1.4511 | -0.45804 | 0.35228 | 2.4395
| 1.9754 | 0.92126 | -0.18091 | 4.7838
| 0.54214 | 1.1811 | 1.4044 | 3.6614
| 0.40549 | -0.27326 | 0.15781 | 0.264
| 1.7798 | 1.0901 | 0.77677 | 4.9594
| 1.0763 | -2.0379 | 0.12215 | 5.3265
| -0.58682 | -0.90208 | 0.29162 | 1.2431
| -0.94541 | -2.0746 | 0.67583 | 5.6547
| 1.5706 | -0.86268 | -0.24007 | 3.2687
| -0.14499 | -0.28184 | -1.8828 | 3.6453
| 1.6973 | -0.018025 | 0.90808 | 3.7059
| 1.2538 | 1.7783 | 0.23896 | 4.7914
| -0.024143 | -0.17841 | -0.18284 | 0.065843
| -1.3874 | 2.1561 | -0.12792 | 6.59
| -0.49363 | 1.1119 | 1.291 | 3.1466
| 0.33401 | -0.53257 | -0.55408 | 0.7022
| -1.2826 | -1.3787 | -0.21362 | 3.5913
| -0.16459 | 0.67818 | 0.69133 | 0.96496
| -0.48046 | 0.10691 | -2.3446 | 5.7396
| -2.2273 | 0.19125 | 1.5587 | 7.4269
| 0.98935 | -0.46022 | -1.4307 | 3.2375
| -1.408 | -0.65548 | 1.1738 | 3.79
| 0.12226 | 1.3443 | 0.50579 | 2.0779
| 0.60576 | -0.067537 | 2.0619 | 4.6228
| -0.41016 | -0.73225 | -0.45241 | 0.9091
| -0.15838 | -0.87583 | 0.23205 | 0.84601
| 0.74062 | 0.21779 | 0.21573 | 0.64248
| -1.325 | 0.041042 | 1.2616 | 3.3489
| 0.64085 | 0.52961 | -1.4276 | 2.7293
| 0.48102 | -1.1736 | 0.73585 | 2.1501
| 0.68161 | -0.20336 | 0.91431 | 1.3419
| -0.42963 | -0.1909 | 0.36838 | 0.35672
| -0.15595 | 1.7593 | 0.18649 | 3.1543
| -1.298 | 2.1014 | 0.24186 | 6.1593
| 1.9865 | 0.41905 | 0.38034 | 4.2665
| 0.74847 | 2.0259 | 0.52591 | 4.9412
| 1.4391 | 0.84559 | -1.0047 | 3.7957
| -1.1336 | -1.0531 | 0.47168 | 2.6165
| -0.95588 | -0.54376 | 0.054312 | 1.2123
| 0.43586 | -1.1883 | 0.10931 | 1.6139
| 1.3463 | 0.040073 | 0.83217 | 2.5067
| -1.0941 | 0.050603 | -1.4124 | 3.1945
| -0.32543 | -0.26537 | 1.1138 | 1.417
| 0.51018 | -0.2698 | -2.5693 | 6.9346
| -0.36643 | 0.47546 | -1.6177 | 2.9772
| 0.52886 | -0.25717 | 0.18094 | 0.37857
| -0.50611 | -1.9537 | -0.062648 | 4.0772
| 0.79953 | -0.28318 | 0.445 | 0.91747
| -0.69209 | 0.30233 | 1.0249 | 1.6208
| 1.0756 | -1.6803 | -0.018465 | 3.9807
| -0.71142 | 0.72889 | -0.54068 | 1.3297
| 0.36929 | 0.77909 | -0.20818 | 0.7867
| 0.91769 | 0.94936 | 0.58908 | 2.0904
| 0.4203 | -0.32281 | -0.34835 | 0.40221
| 0.62754 | -0.46998 | 0.17188 | 0.64423
| -2.2942 | 0.44587 | -0.34457 | 5.581
| 2.2195 | 0.47684 | -1.0198 | 6.1935
| 0.26382 | -1.3343 | -1.5644 | 4.2973
| -0.2348 | 0.36052 | -0.59375 | 0.53765
| -1.4214 | 0.44029 | -1.492 | 4.4404
| -1.0415 | 0.25268 | 0.51373 | 1.4124
| 1.1077 | -0.81986 | 0.2845 | 1.9802
| -1.0813 | 0.33745 | 2.4008 | 7.0471
| 1.4474 | -0.37317 | -1.9535 | 6.0503
| -0.67001 | 0.32064 | -0.34802 | 0.67284
| 0.57096 | 0.78246 | -0.47805 | 1.1668
| -0.24227 | 0.24729 | 0.12694 | 0.13596
| 2.3468 | 2.7618 | -0.016319 | 13.1349
| 0.11711 | 0.75497 | -1.3701 | 2.4609
| -0.21337 | 0.062455 | 0.31951 | 0.15151
| -0.39661 | -0.16865 | 1.079 | 1.3499
| 0.13113 | 1.3665 | -0.30081 | 1.975
| 0.33816 | -0.83096 | -0.75171 | 1.3699
| 2.484 | -1.0259 | -0.45708 | 7.4318
| 0.9533 | 1.598 | 0.79001 | 4.0866
| 0.41328 | 1.5009 | -0.10028 | 2.4335
| 0.80466 | -0.061197 | 0.96145 | 1.5756
| -0.34926 | 1.1962 | 0.35605 | 1.6796
| 0.11984 | 1.0796 | 0.45687 | 1.3886
| -0.30603 | 1.0427 | 0.86141 | 1.9229
| 0.1671 | -0.4962 | -0.15965 | 0.29962
| 1.3512 | -0.94854 | 0.39328 | 2.8801
| 0.0029695 | 0.20241 | -1.5385 | 2.4079
| -0.59481 | -0.76769 | -1.4718 | 3.1093
| 0.71784 | 0.38422 | 0.29986 | 0.75283
| -1.1523 | -0.49543 | 0.31047 | 1.6696
| 1.1422 | 1.0554 | 0.66646 | 2.8627
| -0.84152 | 6.5224e-06 | -0.12143 | 0.72289
| 0.33778 | -0.65519 | 1.5274 | 2.8763
| 1.2453 | -0.79883 | -1.3623 | 4.0447
| -0.8434 | 1.1511 | -0.76334 | 2.6191
| -1.4528 | 0.54785 | -0.097842 | 2.4205
| -1.0637 | 0.061095 | -0.07306 | 1.1405
| -0.011255 | -0.97742 | 0.22711 | 1.007
| 0.22167 | 1.7423 | 0.31444 | 3.1835
| -0.39012 | -0.49328 | -0.20632 | 0.43808
| 0.98688 | -0.97518 | 1.5801 | 4.4216
| 1.8234 | -0.53359 | -0.057778 | 3.6128
| 1.3646 | 1.1381 | 1.242 | 4.7001
| 0.068981 | -0.77907 | -0.064901 | 0.61592
| -0.32545 | -0.48277 | -0.85959 | 1.0779
| 0.16138 | 0.52831 | -2.1304 | 4.8437
| -0.81701 | 1.2524 | -0.40089 | 2.3967
| -2.2506 | 1.7853 | -1.8457 | 11.6591
| -0.53671 | 1.3478 | 1.2384 | 3.6382
| -1.0812 | -0.0023827 | 1.1286 | 2.4427
| -2.2844 | 0.98983 | 0.081042 | 6.2046
| -0.014974 | 0.28237 | -0.71635 | 0.59311
| 0.50932 | 0.079092 | 0.68644 | 0.73687
| 0.70428 | 0.93339 | 1.6761 | 4.1766
| 0.20481 | -0.43932 | -0.39521 | 0.39114
| -0.73511 | -1.3683 | 0.95162 | 3.3183
| 0.66467 | 0.94949 | -1.9334 | 5.0813
| -1.0698 | 1.5706 | 0.83957 | 4.3162
| 0.40989 | -0.34096 | 2.5627 | 6.8516
| 0.9078 | 0.46504 | -0.97063 | 1.9825
| -2.5895 | -0.3413 | 0.23402 | 6.8768
| -0.89746 | -0.31928 | 0.16592 | 0.93491
| 0.31668 | 0.76891 | -0.66533 | 1.1342
| 0.88241 | -0.60919 | -1.0864 | 2.3301
| -1.6136 | -0.67676 | -0.83847 | 3.7648
| 0.37311 | 0.19046 | -1.3104 | 1.8927
| 0.75651 | 1.0654 | -0.22829 | 1.7594
| 1.4314 | -0.65056 | -1.2026 | 3.9184
| -0.47364 | -1.7476 | 0.59012 | 3.6265
| 0.57992 | 1.911 | 0.15231 | 4.0113
| -2.1666 | -1.8972 | 0.37129 | 8.4313
| 1.32 | 1.1095 | 0.2153 | 3.0198
| -0.020718 | 0.76306 | 0.2187 | 0.63052
| 0.083705 | 1.4526 | -0.89331 | 2.9151
| -0.72602 | 0.57302 | -1.4936 | 3.0863
| -2.464 | 1.4013 | 0.19319 | 8.0721
| 0.33483 | -1.474 | -0.57494 | 2.6153
| -0.37378 | 0.65976 | -0.34975 | 0.69733
| 0.19141 | 1.1457 | -1.0934 | 2.5447
| -1.3364 | -0.25737 | 0.65242 | 2.2778
| 0.91117 | -0.90406 | 0.54647 | 1.9462
| -1.996 | 0.14456 | -0.33304 | 4.1157
| -1.0765 | -0.98114 | -1.3578 | 3.9651
| -0.043446 | -1.1908 | 1.4971 | 3.6612
| 0.9082 | -0.46012 | 0.080642 | 1.043
| -0.90374 | -0.98929 | 1.3243 | 3.5493
| -1.2645 | -0.36122 | -0.2593 | 1.7966
| -1.897 | -0.56473 | -0.43694 | 4.1084
| 1.2469 | 0.22321 | 0.1339 | 1.6224
| 1.3862 | -1.9978 | -0.85234 | 6.6394
| 0.37679 | -0.89462 | -1.3563 | 2.7817
| 1.8705 | 1.1665 | -0.7285 | 5.3902
| -2.186 | -0.47143 | 0.30701 | 5.0949
| -1.1911 | 0.95167 | -0.89705 | 3.1291
| -0.063001 | 0.87498 | -1.5613 | 3.2072
| 1.2283 | -1.5355 | 0.2438 | 3.9259
| 0.37249 | 0.42112 | 1.4478 | 2.4122
| -1.1224 | 0.82648 | -0.26254 | 2.0118
| -1.4051 | -0.10693 | -0.56761 | 2.308
| -1.0931 | 1.9562 | 1.1584 | 6.3634
| -1.6895 | 0.13025 | -1.4668 | 5.0229
| -0.49791 | 2.7811 | 0.64154 | 8.3939
| -1.6458 | 0.98188 | 1.0985 | 4.8792
| 0.29145 | -0.10957 | 0.77841 | 0.70287
| -0.055031 | -0.13096 | 0.64118 | 0.43129
| -1.014 | -1.1125 | -0.3934 | 2.4206
| 0.27134 | 0.69096 | -0.58509 | 0.89338
| -0.68629 | 0.24047 | -0.68053 | 0.99194
| 0.88537 | -1.8297 | 0.12663 | 4.1479
| 2.6597 | -1.1359 | -1.034 | 9.4335
| 0.55829 | -0.9429 | 1.3398 | 2.9958
| -0.5533 | -0.69655 | 0.065039 | 0.79555
| 0.6506 | 0.57641 | 0.13407 | 0.7735
| 0.23319 | -0.64427 | -2.0348 | 4.6097
| -0.0050015 | 0.073536 | 2.3806 | 5.6728
| 1.0644 | -0.30706 | -1.8025 | 4.4763
| 0.13614 | -1.4231 | -2.0029 | 6.0552
| -0.098184 | -1.0057 | 1.3581 | 2.8655
| 2.4765 | 0.59924 | -0.67864 | 6.9527
| 0.38557 | 0.96341 | -0.36251 | 1.2082
| -1.2023 | -1.773 | -0.3277 | 4.6964
| -0.28792 | -1.0975 | -1.9985 | 5.2814
| 0.91265 | -0.07963 | 0.064572 | 0.84344
| -0.074822 | 1.3235 | 0.82286 | 2.4342
| 0.27824 | 1.2957 | -0.17201 | 1.7859
| -0.87124 | 1.0514 | 1.2589 | 3.4492
| -1.2683 | 1.1823 | 0.99792 | 4.0024
| -0.26823 | 1.6179 | -1.2089 | 4.1511
| -0.15749 | -0.95304 | 1.0842 | 2.1087
| -0.53367 | 0.45098 | -0.81532 | 1.1529
| -0.27385 | -0.017627 | 0.97329 | 1.0226
| 0.36024 | -1.7966 | 1.0743 | 4.5116
| 0.10526 | 1.4114 | 0.24424 | 2.0627
| 0.83069 | -0.0054129 | 1.0469 | 1.786
| 0.47005 | 0.50938 | 0.00054108 | 0.48042
| 0.64273 | 0.52427 | 1.0605 | 1.8127
| -0.70425 | 0.29013 | 1.3759 | 2.4733
| -0.19798 | -0.54999 | -0.99203 | 1.3258
| -0.1857 | 1.2271 | 0.092481 | 1.5488
| 0.85295 | 0.19229 | 0.090264 | 0.77264
| -1.8323 | 0.3257 | -0.5749 | 3.7941
| -1.1305 | -1.2912 | 0.9813 | 3.9083
| 1.0513 | 0.50498 | 0.53991 | 1.6518
| 0.71219 | -2.3018 | -0.052666 | 5.8084
| 0.61006 | -1.9144 | -0.016163 | 4.0373
| -0.29261 | 0.18052 | 1.5648 | 2.5669
| 0.027999 | 0.4381 | -0.97919 | 1.1515
| -0.41075 | 1.2757 | 1.9593 | 5.6349
| 1.1036 | 1.568 | -0.19734 | 3.7157
| -0.3652 | -0.47588 | -0.53268 | 0.64358
| 0.96793 | 0.161 | 2.3335 | 6.408
| 0.74176 | 0.50766 | 1.2436 | 2.3544
| 0.69813 | 1.828 | 0.89304 | 4.6265
| -1.2278 | 0.035075 | 0.83364 | 2.2037
| 0.36869 | 0.62438 | -0.48133 | 0.75746
| -0.62802 | 0.24496 | 1.0903 | 1.6431
| -0.73194 | 0.94956 | 0.14253 | 1.4577
| 2.5357 | 0.37917 | -0.80521 | 7.2221
| 1.4444 | -1.7066 | 2.6853 | 12.2096
| 1.5797 | 0.51959 | -1.4133 | 4.7628
| -1.089 | -0.33365 | -0.51825 | 1.5657
| -0.12757 | 0.51202 | 0.6664 | 0.72253
| 1.1747 | 0.69836 | 0.11695 | 1.8812
| 0.072374 | 3.2682 | -0.68008 | 11.1491
| -1.9163 | -1.8121 | -1.2147 | 8.4314
| -0.73832 | 0.25136 | -0.16645 | 0.636
| 1.9729 | -1.5493 | 2.0074 | 10.3219
| -0.61002 | -1.9646 | 1.0948 | 5.4304
| 0.28062 | -1.1068 | -0.74174 | 1.854
| 0.83965 | 1.5882 | -0.50366 | 3.481
| 0.17862 | 0.55498 | 0.41967 | 0.51603
| -0.72954 | -0.43305 | 0.76293 | 1.3018
| 1.3548 | -0.56099 | 0.79249 | 2.7783
| -0.22495 | -0.14603 | -0.069233 | 0.076722
| -0.70744 | 0.50534 | -0.9352 | 1.6304
| 0.33526 | -0.71799 | 0.75369 | 1.1959
| -1.3364 | -0.14286 | 0.30777 | 1.9011
| -0.15155 | -1.3592 | -0.1495 | 1.8927
| -0.014617 | -0.3107 | 0.80088 | 0.73815
| 1.5658 | 0.96208 | 0.27397 | 3.4524
| -1.022 | 0.47093 | 0.98865 | 2.2436
| -0.48163 | 2.3074 | -0.24539 | 5.6164
| -0.66749 | -0.31875 | -0.47872 | 0.77632
| 0.5094 | -0.23727 | 0.26956 | 0.38845
| -0.42885 | -1.631 | -1.3824 | 4.7552
| 1.0707 | 1.8539 | 0.5911 | 4.9326
| 0.70074 | -2.4876 | -0.41877 | 6.8543
| 0.13834 | 0.25637 | 1.4125 | 2.08
| -0.45162 | -0.13562 | -0.74981 | 0.78457
| -0.29745 | -0.86145 | -0.88968 | 1.6221
| 1.6827 | 0.14984 | 1.0427 | 3.941
| 0.69095 | -0.041268 | -0.16016 | 0.50477
| -0.94897 | 1.3832 | 1.7084 | 5.7323
| -0.99075 | 0.34015 | -0.087684 | 1.105
| 0.74332 | 1.0583 | 1.291 | 3.339
| -0.80494 | -0.025736 | 0.5055 | 0.90412
| 0.9509 | 0.16841 | -0.76708 | 1.521
| 0.31201 | -0.15006 | 0.40693 | 0.28546
| 0.16139 | -0.30873 | -0.79604 | 0.75504
| -0.36313 | -0.052839 | 0.58813 | 0.48055
| -0.21679 | -1.2887 | -0.15977 | 1.7334
| -0.32081 | -0.21961 | -0.59373 | 0.50366
| 0.43086 | -0.71781 | -0.91918 | 1.5458
| -0.59308 | 1.316 | -0.58226 | 2.4226
| -0.7492 | 0.72917 | -0.27391 | 1.168
| -1.3466 | -0.18503 | 1.8064 | 5.1105
| 0.90033 | -0.59975 | 0.42944 | 1.3547
| 0.2923 | 1.4764 | -0.015414 | 2.2655
| 1.2229 | 0.14169 | 0.65241 | 1.9413
| -0.52747 | 0.30027 | -0.081197 | 0.37498
| -1.6465 | 2.1476 | 0.96593 | 8.2561
| -0.27745 | -0.35415 | -0.16449 | 0.22946
| -0.52003 | -1.1847 | -2.8919 | 10.0369
| 0.6011 | -0.053019 | 1.363 | 2.2219
| -0.26085 | 1.8916 | -0.14031 | 3.6657
| -0.31758 | 0.0061578 | -0.009159 | 0.10098
| -0.88478 | 0.19464 | -2.1147 | 5.2929
| -0.2971 | -0.2947 | 0.18362 | 0.20883
| 0.19325 | -1.7329 | -0.096549 | 3.0495
| 0.64472 | 0.79976 | -1.4027 | 3.0228
| -0.75816 | -0.51025 | 0.0007695 | 0.83517
| -0.43245 | 0.019762 | -0.75662 | 0.75988
| -0.25483 | 0.10278 | 0.47049 | 0.29686
| -0.16351 | 2.1092 | -0.01209 | 4.4754
| 0.17471 | -0.73515 | 1.2969 | 2.2529
| -0.43627 | 1.2651 | -1.1679 | 3.1549
| 1.3315 | -1.157 | 0.66235 | 3.5501
| -1.3695 | 0.67784 | 0.58325 | 2.6753
| 2.0622 | 0.094783 | -0.098438 | 4.2715
| -0.17872 | 0.16482 | -0.34126 | 0.17556
| 1.2655 | -0.3589 | 0.37796 | 1.8731
| -0.88712 | 0.049409 | 0.65879 | 1.2234
| -1.413 | 0.90744 | -0.18119 | 2.8529
| 0.64519 | -0.21257 | 0.43288 | 0.64884
| 0.22965 | 1.3655 | -0.87122 | 2.6764
| 0.73256 | 0.060087 | -2.7667 | 8.1949
| 0.89871 | -0.90985 | -1.2316 | 3.1523
| 0.021509 | -0.97544 | 0.22511 | 1.0026
| 0.15597 | 0.18184 | -1.2791 | 1.6934
| 0.4326 | -0.9712 | -0.88996 | 1.9224
| 1.1936 | 0.51206 | 1.2901 | 3.3513
| -0.8379 | -0.3424 | 0.78241 | 1.4315
| -0.30713 | 0.58044 | -0.96204 | 1.3568
| 1.4553 | -1.1174 | -1.0367 | 4.4413
| 0.30377 | -0.24758 | -1.9436 | 3.9311
| 0.40249 | -1.5465 | -0.010088 | 2.5537
| 0.2677 | 0.64598 | 0.80283 | 1.1335
| -0.81921 | 1.7438 | -0.80469 | 4.3596
| -0.30421 | -1.3683 | 1.4413 | 4.0422
| -0.79342 | 1.2432 | -1.0159 | 3.2073
| 1.2585 | 0.89532 | -1.2966 | 4.0666
| -1.7658 | -0.98157 | 0.75741 | 4.6553
| -0.2025 | 0.84873 | -0.82653 | 1.4445
| -0.38084 | 1.7584 | 1.6175 | 5.8532
| -0.026819 | -0.38346 | -0.022778 | 0.14828
| -2.0111 | -0.060662 | -1.2299 | 5.561
| -0.10609 | -0.67128 | -0.56078 | 0.77635
| 0.49838 | -1.6618 | 0.44276 | 3.2059
| -1.0229 | 0.045961 | -0.35625 | 1.1754
| -0.2588 | 1.5856 | -0.59051 | 2.9298
| -0.85364 | -1.1097 | -0.81592 | 2.6258
| 0.668 | -0.43875 | 0.20735 | 0.68173
| 0.39038 | -1.3899 | 0.22616 | 2.1353
| -0.3915 | 1.3688 | -0.3108 | 2.1236
| -0.15786 | -0.96606 | 2.2908 | 6.2059
| 1.6828 | -1.4112 | -1.1165 | 6.0698
| -1.4485 | -0.39488 | 0.22992 | 2.3068
| -1.304 | -0.68703 | 0.14728 | 2.1942
| 0.22053 | 1.3923 | 0.23664 | 2.043
| 0.086176 | 1.3984 | -0.93779 | 2.8423
| -0.09544 | 1.958 | 0.088691 | 3.8509
| -0.31229 | 0.32043 | -0.53922 | 0.49095
| 0.74587 | -0.42479 | 1.0379 | 1.8139
| -1.139 | 0.48324 | -0.29763 | 1.6195
| -0.12569 | 0.19592 | -0.97583 | 1.0064
| 0.14704 | -2.0211 | 0.46242 | 4.3201
| -1.1827 | -0.72865 | 0.058481 | 1.9332
| -0.78931 | -0.26781 | 1.0168 | 1.7285
| 0.75022 | -0.13437 | -0.83424 | 1.2768
| -0.64743 | -0.37686 | -0.063358 | 0.56521
| -0.0068065 | 0.38186 | -0.75323 | 0.71322
| 0.23604 | 2.3834 | 0.08348 | 5.7434
| -0.49815 | 1.0583 | -0.23832 | 1.425
| -0.58897 | -0.87624 | 0.95803 | 2.0325
| 2.108 | -0.092785 | -0.94598 | 5.3472
| 0.93277 | 1.5025 | -1.2473 | 4.6833
| 0.77627 | 0.22906 | -0.6042 | 1.0201
| 0.88698 | 2.0899 | -0.2511 | 5.2175
| -0.81547 | 0.50005 | -1.929 | 4.6362
| -0.22659 | -0.87147 | -0.63971 | 1.22
| -0.148 | 2.1208 | -1.2313 | 6.0359
| 1.6052 | -0.39817 | 0.95489 | 3.6469
| 0.40901 | -1.0216 | -0.9408 | 2.096
| 0.13828 | -0.62335 | 0.73479 | 0.9476
| 1.16 | -0.80453 | 1.0962 | 3.1946
| 2.3768 | -0.89533 | -1.6255 | 9.093
| -1.3928 | 1.3412 | 0.46831 | 3.9582
| -0.45404 | -0.077189 | 0.034695 | 0.21332
| 0.67211 | 0.057927 | -1.0073 | 1.4697
| 1.0723 | 0.94725 | 0.76168 | 2.6273
| -0.47674 | 0.38487 | 0.42053 | 0.55226
| 0.025671 | -0.24407 | 0.47821 | 0.28891
| -0.43361 | 0.66314 | -0.25161 | 0.69108
| -0.72177 | 0.29935 | -0.54281 | 0.90521
| 0.19065 | 0.45907 | -0.19198 | 0.28395
| 0.84795 | 0.33705 | -0.93423 | 1.7054
| 0.70857 | -0.59359 | -0.26334 | 0.92377
| 0.23052 | 0.68761 | 0.84953 | 1.2477
| 1.4047 | 0.20563 | 1.9665 | 5.8825
| 0.2371 | -1.0191 | 1.1303 | 2.3724
| -0.97727 | -1.4298 | 0.060124 | 3.0029
| -1.2491 | 0.72105 | -0.083692 | 2.0871
| 0.60157 | -0.32855 | 1.1567 | 1.8079
| 2.0268 | -0.37002 | 0.68052 | 4.7079
| 0.68637 | -0.52565 | -0.80875 | 1.4015
| 2.1897 | -2.2929 | 1.0352 | 11.1239
| 0.062764 | -1.5689 | -0.4808 | 2.6967
| 0.85936 | -1.9174 | -0.047973 | 4.4172
| -0.3478 | -0.51423 | 0.022753 | 0.38591
| 0.042733 | -0.13143 | 0.47997 | 0.24948
| 0.50659 | -0.69159 | 0.88554 | 1.5191
| 1.141 | -1.0769 | 0.65305 | 2.888
| -1.134 | -0.56101 | 0.38221 | 1.7469
| 0.25737 | -0.18227 | -0.1706 | 0.12857
| 0.76243 | -0.40164 | 1.4458 | 2.8329
| 0.53891 | 0.70393 | -1.0584 | 1.9061
| 0.73018 | 0.35991 | 2.2214 | 5.5974
| -0.70527 | -1.383 | -0.059568 | 2.4137
| 0.029347 | 0.95324 | 1.1798 | 2.3014
| 0.053189 | 1.0728 | -1.1536 | 2.4845
| 0.063241 | -0.72169 | -1.2609 | 2.1146
| -0.55375 | -0.20972 | 0.94036 | 1.2349
| 0.28291 | 0.16189 | 0.94812 | 1.0052
| -0.62857 | -0.68278 | -0.95993 | 1.7827
| -0.59766 | 0.32744 | -0.27264 | 0.53874
| 1.824 | 0.66245 | 0.31988 | 3.8681
| -0.021468 | 1.5973 | 0.95131 | 3.4567
| -1.1895 | 0.92561 | 0.67137 | 2.7225
| 0.81057 | -1.0897 | 0.99575 | 2.836
| -1.6674 | -0.4318 | -1.1147 | 4.2091
| -1.208 | 1.5566 | 1.3335 | 5.6604
| -0.91202 | -0.83474 | 0.8461 | 2.2445
| -0.11309 | -0.43374 | 0.6699 | 0.64969
| -0.8799 | 0.080914 | -1.7979 | 4.0131
| 0.4857 | 1.4305 | 0.54344 | 2.5776
| -1.2044 | 1.4491 | -0.18109 | 3.5834
| -1.2926 | 1.2009 | 0.35722 | 3.2406
| -0.084922 | -0.85455 | -1.5381 | 3.1032
| 0.096278 | 1.1881 | -0.48148 | 1.6526
| -0.21318 | -0.83956 | 1.1057 | 1.973
| 0.70427 | 0.48327 | -0.034112 | 0.73072
| -0.55077 | -0.52318 | -0.016301 | 0.57733
| 0.77208 | 1.1338 | -0.053028 | 1.8845
| -0.81106 | -0.68099 | -0.97897 | 2.0799
| 1.5656 | -1.3514 | -0.58363 | 4.6178
| 2.8707 | -1.1221 | -0.12004 | 9.5142
| -0.34818 | 1.1328 | -1.1858 | 2.8105
| -0.63368 | 0.054314 | -1.5168 | 2.7051
| 1.0803 | 1.0077 | -0.71068 | 2.6875
| 0.30018 | -1.0981 | 1.2128 | 2.7668
| 0.52094 | -1.7326 | -0.70527 | 3.7706
| 0.45964 | 0.6442 | -1.3151 | 2.3557
| -2.0177 | 0.44554 | -0.51201 | 4.532
| -0.93612 | -0.093316 | 0.9031 | 1.7006
| 1.1148 | -0.12758 | -0.73059 | 1.7927
| -1.6074 | 0.015781 | 0.19998 | 2.624
| -1.7746 | -0.314 | 0.96619 | 4.1814
| 1.0077 | -0.49438 | -1.1786 | 2.6488
| 2.1337 | -0.29333 | -0.21493 | 4.6848
| -0.50477 | -1.4758 | 0.32272 | 2.537
| -0.052393 | 1.5286 | 1.2084 | 3.7995
| 1.5083 | 0.74162 | 0.10307 | 2.8356
| -1.9677 | 1.7068 | -0.84346 | 7.4963
| 0.36639 | -0.93785 | -0.62395 | 1.4031
| 0.5798 | -0.098758 | -1.0607 | 1.4709
| -1.4444 | 0.76181 | 0.61393 | 3.0436
| -1.5284 | 0.055558 | 2.2726 | 7.5039
| 0.27368 | 1.0756 | 0.29364 | 1.3179
| -0.40191 | 1.3262 | -2.588 | 8.6179
| -0.51944 | -2.0124 | -0.15853 | 4.3445
| -0.029448 | 0.78191 | -0.23324 | 0.66665
| -0.30593 | 1.0971 | -0.90481 | 2.1159
| -0.20477 | 1.8214 | -0.70616 | 3.8579
| -0.19112 | -1.0492 | 0.71738 | 1.652
| -0.75532 | 0.81366 | -1.7672 | 4.3555
| 1.7812 | -0.32795 | 0.38059 | 3.4251
| -0.082485 | 2.4864 | -0.35607 | 6.3159
| -2.0091 | 0.33766 | 0.32422 | 4.2554
| -0.25946 | 0.34916 | -0.45673 | 0.39783
| -0.10902 | 0.23656 | 0.05738 | 0.071136
| 0.44845 | 0.27041 | 0.22006 | 0.32266
| 0.3902 | -1.7318 | 1.47 | 5.3124
| -0.29592 | 0.58731 | -2.3701 | 6.0497
| -0.24156 | 2.2554 | 0.35257 | 5.2697
| -1.7918 | 0.73558 | 0.39736 | 3.9097
| 1.0983 | -0.66924 | 1.0803 | 2.8212
| 2.6157 | 2.1641 | -0.44704 | 11.7252
| 0.78742 | 1.2811 | 0.65064 | 2.6847
| 1.0378 | -0.27141 | -0.32004 | 1.2532
| -0.023999 | 1.8815 | 0.073587 | 3.5462
| 0.076728 | -0.00064245 | 0.18026 | 0.038381
| -0.41097 | -0.48559 | 2.5863 | 7.0937
| 0.3038 | 0.10817 | -0.63838 | 0.51152
| 0.3398 | -1.5344 | 0.94323 | 3.3595
| 1.1983 | -0.083653 | -0.50652 | 1.6994
| -0.40539 | 0.83826 | -1.2184 | 2.3515
| 0.66368 | 0.82475 | -0.77974 | 1.7287
| -0.76819 | 1.1192 | 0.11031 | 1.8549
| 1.2 | -0.50118 | -0.80356 | 2.3369
| -0.38013 | -0.58249 | 0.39543 | 0.64015
| -0.2397 | 0.16482 | -0.55434 | 0.39191
| -0.60813 | -0.4681 | 1.8234 | 3.9138
| 0.81395 | -0.62558 | -0.58428 | 1.3953
| -1.1891 | -0.33139 | -1.5544 | 3.9401
| 1.3983 | 1.1727 | 0.021532 | 3.3308
| -0.39597 | -0.41642 | 1.0258 | 1.3825
| 0.30548 | -0.3724 | 0.36597 | 0.36593
| -0.47095 | 1.2255 | -0.36922 | 1.8601
| 2.1722 | 0.93008 | -1.079 | 6.7477
| -1.3389 | -3.1593 | 0.43346 | 11.9618
| -1.0943 | 0.026629 | 0.099869 | 1.2081
| -1.2756 | -0.87969 | -0.045188 | 2.403
| -1.0834 | -0.78692 | -0.95695 | 2.7088
| 0.57316 | -0.17378 | 0.022036 | 0.3592
| -0.17157 | 1.4554 | 1.6432 | 4.8476
| 0.05346 | 0.54947 | -0.70097 | 0.79613
| 0.84077 | 0.9273 | -0.29325 | 1.6528
| -0.49564 | -0.61485 | 0.37809 | 0.76666
| 0.55362 | -1.3897 | -1.0165 | 3.2711
| 1.8105 | -0.82194 | -0.20992 | 3.9975
| 0.091819 | 0.60081 | -1.4188 | 2.3824
| 0.16514 | 0.062631 | -1.867 | 3.5168
| 1.4046 | -0.34507 | -0.97017 | 3.0333
| 1.3656 | -0.11544 | -2.3335 | 7.3234
| 0.17788 | -0.0079562 | -0.56138 | 0.34686
| 0.11857 | -1.1599 | -0.94207 | 2.2468
| 1.6828 | 0.80307 | 0.23197 | 3.5304
| -1.8799 | -1.4429 | -0.69486 | 6.0989
| -0.93372 | -0.16952 | 0.08648 | 0.90805
| -1.417 | 0.34154 | 0.032307 | 2.1256
| -1.8579 | 0.3083 | -0.85235 | 4.2733
| -0.31264 | -1.1372 | -1.5937 | 3.9309
| 0.40919 | -0.7247 | 0.82044 | 1.3658
| -0.45236 | -2.0755 | -0.89979 | 5.3221
| -0.07645 | 0.76908 | 0.064009 | 0.60143
| -1.0517 | 3.0729 | 1.2769 | 12.1795
| -0.85836 | -0.34709 | 1.759 | 3.9513
| -0.19881 | 0.87245 | -0.71511 | 1.3121
| 0.119 | -1.4658 | -0.79842 | 2.8002
| 0.72246 | -0.20896 | 1.1434 | 1.8729
| 1.0041 | -0.9906 | 0.62041 | 2.3745
| 2.4562 | -0.11188 | 0.89434 | 6.8452
| 0.19747 | -0.094547 | 1.0316 | 1.112
| -0.4622 | 0.67734 | -0.064243 | 0.67655
| -2.7326 | 0.46564 | -0.50865 | 7.9425
| 1.2412 | 0.70306 | 0.38475 | 2.1828
| 1.3617 | -1.3827 | -0.72454 | 4.2912
| -0.67323 | -1.3168 | -1.7355 | 5.1992
| -0.99844 | -0.90362 | -1.493 | 4.0424
| -0.075695 | 1.0923 | 2.1711 | 5.9125
| 1.0263 | 0.95608 | -0.72576 | 2.4941
| -2.0086 | 1.393 | 0.49589 | 6.2209
| 0.28872 | -0.35129 | -0.39848 | 0.36555
| -0.38117 | -1.1462 | 1.2892 | 3.1211
| 1.3218 | 1.2836 | -0.60649 | 3.7628
| 0.25592 | 0.27478 | 0.50983 | 0.40092
| -1.5061 | 1.2739 | 1.4568 | 6.0136
| -0.82222 | -1.9453 | -0.22331 | 4.5101
| 0.36727 | -1.8129 | 0.97791 | 4.3779
| 0.36004 | 1.6449 | -0.11981 | 2.8498
| -1.2042 | -0.069132 | -0.47221 | 1.678
| -1.2674 | 1.1478 | 0.49863 | 3.1725
| 1.0137 | -0.77912 | -0.75029 | 2.1977
| 0.077087 | 0.56553 | 0.044442 | 0.32774
| -0.17194 | -1.3494 | -0.92918 | 2.7138
| 0.22772 | 0.37365 | -1.4212 | 2.2113
| -1.7206 | -0.37005 | 0.55951 | 3.4103
| 0.52593 | -1.0086 | 0.90364 | 2.1105
| -0.94526 | -1.3329 | -1.218 | 4.1536
| -1.0165 | -2.2259 | 2.3744 | 11.6258
| 1.4793 | -0.18454 | 1.0763 | 3.381
| 0.23095 | 1.3211 | 0.75353 | 2.3665
| 0.040376 | -1.5771 | 0.34247 | 2.6062
| -0.042813 | -0.73973 | 1.0511 | 1.6538
| 0.63965 | 0.62401 | -0.026511 | 0.79924
| -0.76261 | 0.9014 | -1.3384 | 3.1855
| -0.50455 | -0.52507 | 0.69831 | 1.0179
| -1.0802 | 0.10478 | 0.064346 | 1.1819
| -1.6204 | 0.67966 | -1.3227 | 4.8371
| -0.011134 | 0.33407 | 0.19545 | 0.14993
| -0.56267 | -0.43156 | -0.27873 | 0.58054
| 0.31973 | 0.2329 | 1.1615 | 1.5056
| 0.65911 | 0.082334 | 1.124 | 1.7045
| 1.626 | -1.5349 | -0.76675 | 5.5878
| 0.095631 | 1.0302 | -1.4299 | 3.1153
| 2.1921 | 0.27155 | -0.59005 | 5.2271
| 1.106 | -1.0621 | -1.6461 | 5.061
| 1.0165 | 0.83663 | 0.68469 | 2.202
| 0.28983 | -0.35748 | 0.17836 | 0.2436
| -1.2036 | 0.5425 | -1.2815 | 3.385
| -1.9213 | 2.1635 | 0.96145 | 9.2965
| -0.43004 | 0.78684 | -0.9087 | 1.6298
| 0.19441 | -0.69742 | -1.4419 | 2.6033
| -0.10474 | -0.38215 | -0.72259 | 0.67916
| 0.28895 | 0.33861 | -0.61653 | 0.57825
| 0.69219 | -0.41486 | -1.2392 | 2.1869
| -0.30447 | -0.36499 | -0.12185 | 0.24077
| -1.056 | 0.059234 | -0.2438 | 1.1781
| 1.1038 | -0.50277 | 1.8118 | 4.7538
| 1.1856 | -0.49255 | 2.4606 | 7.7029
| -0.056424 | -0.59741 | 0.82124 | 1.0345
| 0.087444 | -0.29973 | -1.1425 | 1.4027
| -1.2407 | 0.92825 | 0.79234 | 3.0288
| -1.0662 | 1.5039 | 0.27661 | 3.475
| -0.50434 | 0.18393 | 0.76173 | 0.86842
| -0.086848 | 1.5914 | 1.0438 | 3.6296
| -0.2062 | -0.75369 | 0.83416 | 1.3064
| 0.21509 | 0.20235 | 1.0436 | 1.1763
| 0.4708 | -1.6728 | 1.0922 | 4.2129
| 0.40799 | -1.4727 | -0.20209 | 2.3761
| -0.77362 | 1.7968 | 0.2105 | 3.8715
| 0.36052 | -1.5004 | 0.031571 | 2.3822
| -0.39948 | -0.88325 | -0.4898 | 1.1796
| -0.32812 | -0.39922 | -1.1509 | 1.5915
| -0.98576 | -0.77215 | -0.22001 | 1.6163
| -1.0494 | -0.14263 | -2.2115 | 6.0126
| 0.39269 | -0.38678 | -1.5503 | 2.7072
| -0.67618 | -0.39056 | 0.031632 | 0.61076
| 0.15065 | 1.7212 | -0.23668 | 3.0411
| 0.98821 | 1.473 | 0.68636 | 3.6173
| 1.6591 | -0.17949 | 1.1258 | 4.0521
| 0.12611 | -1.0821 | -1.1836 | 2.5878
| 0.89726 | 0.94748 | 1.0199 | 2.7429
| -0.73461 | 0.29355 | -1.4675 | 2.7793
| 0.12712 | -0.34824 | 1.5574 | 2.563
| 0.02656 | -0.40126 | 0.64447 | 0.57706
| 1.6937 | 1.6259 | -1.199 | 6.9496
| -0.26941 | -0.1372 | 0.011382 | 0.091534
| 1.045 | -1.4003 | -1.6146 | 5.6596
| -1.6218 | -0.66938 | 1.1328 | 4.3616
| -0.33194 | 1.0907 | -0.4231 | 1.4788
| 1.0785 | -0.69861 | -1.2378 | 3.1833
| -1.0315 | 0.1906 | -0.86093 | 1.8416
| 0.48278 | 0.038972 | 0.81448 | 0.89797
| 1.2753 | 0.041545 | -1.445 | 3.7162
| -0.48938 | 0.28583 | -0.92678 | 1.1801
| 1.281 | -0.098425 | 1.677 | 4.4628
| -0.45818 | 1.3071 | 1.324 | 3.6713
| 1.9785 | 0.39383 | 2.5015 | 10.3271
| -0.84296 | -2.1577 | -1.4599 | 7.4976
| -0.045469 | 0.11074 | 0.1396 | 0.033819
| 0.12894 | 0.6124 | -0.30391 | 0.48403
| 0.48606 | -0.83026 | 1.2352 | 2.4513
| -1.0662 | 0.063128 | -1.5484 | 3.5383
| 0.82233 | 0.093464 | 0.50205 | 0.93702
| 0.15501 | -0.069733 | 0.14942 | 0.051217
| 0.47998 | 0.69421 | -0.11256 | 0.72498
| -1.84 | -0.63177 | -0.37447 | 3.925
| -0.47933 | 0.074316 | -0.45622 | 0.44341
| 1.9198 | 2.0542 | 1.7136 | 10.8417
| -1.195 | 2.9619 | -0.48874 | 10.4398
| 0.67434 | 0.84406 | -0.76686 | 1.7552
| -1.5898 | 0.40414 | 0.14017 | 2.7106
| 0.48658 | 0.40923 | -0.61768 | 0.78575
| -2.3943 | -1.7844 | 1.1329 | 10.2
| 0.44464 | -0.12572 | -0.75127 | 0.77792
| -0.053786 | -0.51057 | 0.075832 | 0.26933
| -0.27481 | 0.45244 | 2.2744 | 5.453
| 0.83845 | -0.60533 | 0.65986 | 1.5048
| -0.39197 | -0.26213 | -0.6023 | 0.58512
| -0.70472 | -1.8729 | -0.8605 | 4.7449
| 1.149 | -1.3894 | -0.80803 | 3.9034
| -1.4542 | 0.10968 | 0.60377 | 2.4913
| 0.90825 | -0.57651 | -0.39977 | 1.3171
| -1.8519 | -1.9521 | 0.88047 | 8.0154
| 0.55996 | 0.012542 | 0.67826 | 0.77375
| -1.9894 | 1.6437 | 0.77265 | 7.2562
| -0.2862 | -0.57396 | 0.28531 | 0.49275
| 0.57008 | -1.3857 | 1.2875 | 3.9029
| -1.0833 | -0.67215 | -1.2768 | 3.2556
| 0.37133 | 0.90623 | -0.18397 | 0.99298
| 0.44131 | -1.5062 | -1.2226 | 3.9582
| 0.14951 | -2.5234 | 0.2883 | 6.4731
| 0.1174 | -0.58492 | -0.52069 | 0.62703
| -0.86236 | -1.1606 | -0.05075 | 2.0933
| -0.48073 | 1.0852 | 0.75162 | 1.9737
| 1.4108 | -1.3933 | 0.22602 | 3.9826
| 0.98441 | 0.19328 | 1.0441 | 2.0966
| -1.1978 | 1.07 | -1.179 | 3.9696
| 0.55796 | -1.2229 | -0.47937 | 2.0365
| 0.76071 | 1.6118 | -0.482 | 3.4089
| 1.4874 | 1.7563 | -1.4339 | 7.3529
| -0.91855 | 0.9419 | 0.8452 | 2.4453
| -1.4298 | -0.62903 | 0.79109 | 3.0659
| -1.1891 | -0.26812 | 0.33767 | 1.5998
| -0.4306 | 1.0579 | -0.4363 | 1.4949
| 0.2903 | -0.83616 | 0.22061 | 0.8321
| 1.0074 | 1.1593 | -0.0121 | 2.3589
| 0.18315 | -0.63511 | 0.35173 | 0.56062
| 0.55399 | 0.119 | 0.45934 | 0.53207
| 0.45414 | -0.15434 | -0.70974 | 0.73379
| -0.37354 | -0.35243 | -1.2504 | 1.8273
| 0.74694 | -0.85018 | 1.0262 | 2.3337
| 0.31212 | 0.19131 | 0.41213 | 0.30387
| -1.731 | 2.1304 | -0.48802 | 7.773
| -0.42772 | -0.82912 | -1.4774 | 3.0531
| -0.063717 | 1.2586 | 2.255 | 6.6729
| 1.3071 | 0.46808 | 1.0189 | 2.9658
| 1.0509 | 0.17253 | 0.32813 | 1.2418
| 0.92231 | 0.29944 | -1.4549 | 3.057
| 0.0298 | 0.82093 | 0.23742 | 0.73118
| 0.38721 | 0.077022 | -2.0628 | 4.4108
| 1.0069 | 0.68461 | -0.57514 | 1.8133
| 1.3039 | -0.91658 | -0.35598 | 2.6669
| -0.11098 | 0.54148 | -0.4589 | 0.5161
| 0.90246 | 0.4348 | -0.019716 | 1.0039
| -0.66114 | -0.46755 | 0.83615 | 1.3549
| -0.60391 | 0.3309 | -0.77061 | 1.068
| 0.3535 | 0.1133 | 0.42404 | 0.3176
| 1.6112 | -0.029259 | -1.9198 | 6.2823
| -0.54197 | 0.069104 | 1.82 | 3.6111
| 0.42882 | -0.19776 | 0.9631 | 1.1506
| 1.0128 | 0.5815 | 0.076483 | 1.3698
| -1.3301 | 0.1928 | 0.037321 | 1.8079
| -0.15672 | -0.52305 | -0.17949 | 0.33036
| -0.86565 | 0.52441 | 0.10309 | 1.035
| -0.47111 | 0.66519 | -0.77119 | 1.2592
| 1.6833 | -1.0618 | -0.79267 | 4.5892
| 0.67252 | -0.82599 | -0.11327 | 1.1474
| 2.1395 | 0.80559 | 0.39854 | 5.3853
| 1.1786 | -0.20043 | -0.46989 | 1.65
| 1.0887 | -0.32149 | -1.2935 | 2.9617
| 0.031367 | -0.82608 | -0.31857 | 0.78488
| -0.6453 | 0.88341 | -0.5836 | 1.5374
| -0.56357 | -0.80522 | 1.3333 | 2.7436
| -0.57026 | -0.99601 | -0.50141 | 1.5687
| 0.55141 | 0.043498 | 0.89124 | 1.1002
| 0.98039 | 0.11543 | -0.78926 | 1.5974
| 0.69808 | 0.073968 | -1.9433 | 4.2693
| 0.59507 | -1.2805 | 0.13876 | 2.0131
| 0.5417 | 1.5202 | 0.72101 | 3.1242
| 0.52358 | 0.81837 | 1.1756 | 2.326
| -0.50011 | 0.41297 | 0.9907 | 1.4021
| 0.80349 | -0.40479 | -0.68189 | 1.2744
| 0.03178 | 0.10958 | 0.41593 | 0.18602
| -0.052682 | 2.2229 | 0.61261 | 5.3196
| -0.91405 | 0.50085 | 0.73625 | 1.6284
| 1.2063 | -0.96887 | 2.7527 | 9.9711
| -0.3135 | 0.42541 | -0.36397 | 0.41172
| -0.56948 | -1.5642 | 0.38418 | 2.9186
| 0.54365 | 1.1422 | 0.56612 | 1.9206
| -0.38354 | -1.6773 | -1.3304 | 4.7303
| -0.40167 | -1.2997 | 0.40292 | 2.013
| 1.324 | -1.3016 | 1.6626 | 6.2113
| 0.053672 | -1.3689 | 1.9824 | 5.8066
| -0.37613 | -1.0044 | -0.79009 | 1.7745
| 1.0236 | 0.14398 | 2.9866 | 9.9884
| -1.3393 | 1.5567 | 0.84575 | 4.9322
| 1.5807 | 2.2172 | 1.3299 | 9.1834
| 0.27209 | 0.96205 | 1.027 | 2.0543
| -0.91826 | 0.28168 | 0.048997 | 0.92494
| -0.63085 | 0.47826 | 0.60488 | 0.99259
| 0.81277 | 0.21532 | 0.73999 | 1.2545
| 0.70605 | -0.95178 | -2.5728 | 8.0236
| -0.40199 | 1.2383 | 1.132 | 2.9765
| 0.6703 | -0.074115 | -0.036039 | 0.4561
| -1.6503 | 0.16291 | -1.4906 | 4.9719
| 1.8441 | 1.1789 | 1.7758 | 7.9439
| -1.4198 | -0.15527 | -1.0334 | 3.1078
| -1.1118 | 1.4568 | -0.1422 | 3.3788
| -0.23636 | -0.22669 | -0.63835 | 0.51474
| 0.22947 | -0.39537 | -1.3104 | 1.9261
| 0.69905 | -0.97017 | -0.097596 | 1.4394
| 0.21799 | -0.054083 | 0.372 | 0.18883
| -0.0026889 | 0.031914 | 0.46012 | 0.21274
| 0.18589 | -0.052986 | -0.09645 | 0.046664
| 0.045444 | 0.39767 | 0.00061682 | 0.16021
| -1.3533 | -0.8029 | -0.11907 | 2.4902
| -0.74913 | -0.092737 | -0.46386 | 0.78496
| 1.3537 | 0.38298 | -0.045895 | 1.9812
| 1.5586 | -0.69049 | -0.36894 | 3.0421
| 0.23212 | 0.54061 | 0.86991 | 1.1029
| -0.49316 | -0.76367 | -0.27582 | 0.90247
| -0.071579 | 0.39071 | -0.42575 | 0.33905
| -0.77732 | 0.18249 | 0.82591 | 1.3197
| -0.91303 | -1.3741 | -0.38965 | 2.8736
| -0.29909 | 0.55581 | -0.89813 | 1.205
| 0.46033 | -0.95044 | 1.8918 | 4.694
| -0.88081 | 0.40824 | 1.2881 | 2.6017
| -0.48355 | 0.21726 | -0.063261 | 0.28503
| -1.252 | -0.43252 | -0.61055 | 2.1275
| 0.38423 | -0.60121 | 1.8388 | 3.8902
| -0.99586 | 2.1601 | 0.75697 | 6.2309
| 2.4532 | -0.7944 | 1.3051 | 8.3524
| -1.3507 | 0.71993 | -1.4637 | 4.485
| -0.14712 | 0.59756 | -1.596 | 2.9259
| 1.8334 | 1.6605 | 0.13325 | 6.1364
| -1.0991 | -0.40048 | -0.39532 | 1.5246
| -0.45733 | -1.2733 | -1.2479 | 3.3877
| 2.0135 | 0.59359 | -0.95245 | 5.3137
| -2.1987 | 2.2232 | -0.8192 | 10.448
| -0.73251 | -0.018342 | 1.0986 | 1.7438
| 1.2058 | -0.85462 | 1.6356 | 4.8595
| -0.95912 | 0.15715 | -0.40949 | 1.1123
| -0.38977 | -1.1347 | 0.18595 | 1.4741
| -1.5451 | 0.84177 | 0.17905 | 3.1281
| -0.19156 | -1.8582 | 0.085134 | 3.4969
| -0.40324 | 1.8176 | -0.23817 | 3.5228
| -1.1066 | -0.1164 | 1.1744 | 2.6174
| 0.50269 | 1.9684 | 0.3834 | 4.2744
| 0.37756 | -0.28966 | 0.22524 | 0.27719
| -0.29956 | 0.16897 | -0.29985 | 0.2082
| 0.40682 | 1.5375 | 1.274 | 4.1524
| 1.4704 | -0.043594 | 2.3519 | 7.6951
| -0.28219 | 0.63972 | 0.08294 | 0.49575
| -2.2153 | 0.45325 | 0.72137 | 5.6335
| 0.35158 | -0.0022498 | -1.3744 | 2.0126
| -1.6135 | 0.7956 | -1.4049 | 5.2101
| -0.43024 | 0.11614 | 0.91119 | 1.0289
| 0.87167 | -0.63721 | -0.013337 | 1.166
| -1.1647 | -2.0047 | 0.0058949 | 5.3754
| 0.24875 | -0.28411 | 0.38033 | 0.28724
| -0.20577 | -0.026647 | 0.87718 | 0.81249
| 2.2249 | 2.616 | -0.27118 | 11.8671
| -3.358 | 0.79964 | 1.2266 | 13.4199
| -0.21363 | 0.10348 | -0.12649 | 0.072344
| -1.6344 | 1.0418 | -0.02748 | 3.7572
| 1.0025 | -1.5086 | 0.65957 | 3.7159
| 0.83253 | -1.0075 | 2.52 | 8.0584
| 1.7767 | -0.36226 | -0.93535 | 4.1629
| -0.75325 | -0.37961 | 0.52068 | 0.9826
| -0.41267 | -1.3475 | -0.28287 | 2.0662
| 1.9498 | 0.56033 | 1.9197 | 7.8007
| 0.56428 | 1.8621 | -0.86645 | 4.5366
| -0.71316 | 1.3979 | 1.6117 | 5.0602
| -0.33402 | 0.40104 | 2.5618 | 6.8352
| -0.93947 | 0.3195 | 0.84416 | 1.6973
| -0.84173 | 1.5937 | 0.72765 | 3.7778
| 0.069646 | -0.30307 | -0.79554 | 0.72958
| 0.71306 | -1.1956 | -0.68431 | 2.4061
| -1.1454 | 0.5252 | -1.1792 | 2.9783
| 0.32861 | -0.93547 | 2.3925 | 6.7069
| 0.8735 | 0.36021 | -1.894 | 4.4801
| 0.49062 | 0.086817 | 1.2672 | 1.854
| 0.55566 | 1.946 | -0.06882 | 4.1005
| 0.37895 | 0.90288 | -0.42713 | 1.1412
| 1.047 | 2.0898 | -0.50806 | 5.7215
| -1.8286 | -0.93275 | -0.62809 | 4.6082
| 0.024354 | -0.13858 | 0.95941 | 0.94027
| -0.36559 | 1.153 | -1.4523 | 3.5722
| -0.74711 | 1.6251 | -0.76855 | 3.7898
| -0.021498 | 0.98317 | -0.19799 | 1.0063
| -1.1893 | -0.16191 | 1.698 | 4.3238
| 2.2637 | -0.44607 | 0.139 | 5.3425
| -0.65535 | 1.3712 | 0.36063 | 2.4396
| -0.089592 | -1.0462 | 0.28789 | 1.1854
| -0.93911 | -0.84652 | -0.34277 | 1.716
| 1.0878 | -0.016749 | -1.4391 | 3.2547
| 0.005328 | 0.24671 | -0.47182 | 0.2835
| 0.74268 | 0.9851 | 0.16757 | 1.5501
| 1.2291 | -0.36551 | 0.92285 | 2.496
| -0.46911 | -1.4404 | -0.67122 | 2.7453
| -0.12924 | 0.86154 | -2.8468 | 8.8632
| 0.9276 | 0.6407 | 0.40879 | 1.438
| 0.97583 | -1.0081 | 0.75709 | 2.5417
| -0.65099 | -1.6157 | -0.52871 | 3.3139
| 0.92483 | -0.3479 | -1.3972 | 2.9286
| 0.74595 | 0.93738 | -0.48924 | 1.6745
| 1.2546 | -1.8609 | -0.249 | 5.0989
| -1.3677 | 0.0040867 | -0.41819 | 2.0454
| -1.0528 | 1.5903 | 0.74639 | 4.1945
| 0.78377 | 0.44002 | 0.044801 | 0.80992
| -0.5935 | -0.7065 | -0.080955 | 0.85795
| 0.78497 | -1.1224 | 0.17796 | 1.9075
| 1.5401 | 1.2598 | -0.32259 | 4.063
| 1.2851 | -0.96931 | -0.040535 | 2.5927
| 1.0875 | 0.81993 | -1.2249 | 3.3554
| 0.32803 | -0.35708 | -0.62898 | 0.63072
| -0.17541 | -0.79768 | -0.78775 | 1.2876
| 2.5598 | 1.3072 | 0.12195 | 8.2761
| 1.5201 | -0.47581 | -0.58122 | 2.8748
| 0.60082 | 1.2922 | 0.49098 | 2.2718
| 1.2223 | -0.90263 | 0.77273 | 2.9059
| 0.54519 | 1.3851 | 0.39628 | 2.3728
| 1.1859 | 0.68575 | -1.3498 | 3.6985
| 0.40406 | 0.50136 | -0.33947 | 0.52986
| -0.31087 | 0.28842 | -0.22742 | 0.23155
| -0.73714 | -0.17599 | -0.87039 | 1.3319
| -1.4324 | -0.95005 | 1.1066 | 4.1788
| 0.23768 | -0.19348 | 0.016837 | 0.094212
| -0.40714 | 0.44463 | -0.61299 | 0.73921
| 0.49297 | 0.37425 | -0.45917 | 0.59392
| 0.98952 | 1.784 | -1.107 | 5.3873
| 0.089083 | -0.58687 | -0.69009 | 0.82857
| 1.6697 | -1.3327 | 0.28216 | 4.6436
| -0.28057 | 0.58153 | 0.69896 | 0.90544
| -0.85904 | 1.4641 | 0.83026 | 3.5708
| -1.0743 | 0.66312 | -1.29 | 3.2579
| -1.0927 | 0.18676 | -0.73895 | 1.7749
| -0.20616 | 0.48564 | 0.040578 | 0.28
| 0.48548 | 0.2614 | 0.99284 | 1.2898
| -0.74589 | -0.35019 | -0.69141 | 1.157
| -0.39388 | -0.28692 | -1.6177 | 2.8545
| 1.3799 | -0.89507 | 0.32601 | 2.8117
| 0.6599 | 0.72706 | 0.69809 | 1.4514
| 1.3889 | -0.3809 | 1.5204 | 4.3858
| 0.026018 | 0.10184 | -0.12503 | 0.026682
| -0.71336 | -0.51926 | -0.5899 | 1.1265
| -0.19582 | -0.45045 | -0.66543 | 0.68405
| -0.50809 | -1.3471 | -1.5467 | 4.4652
| -1.2486 | -0.077176 | -0.67328 | 2.0183
| 0.82486 | 0.031217 | -0.16755 | 0.70944
| -1.8074 | -0.63156 | 0.9121 | 4.4974
| 0.663 | -1.3189 | -0.52046 | 2.4498
| -0.18143 | 0.98161 | 0.6353 | 1.4001
| -0.44387 | 1.7767 | -1.2909 | 5.0202
| 0.058484 | -0.94683 | 1.7613 | 4.0021
| -0.73204 | 1.5254 | -0.88825 | 3.6516
| -0.52973 | -2.2025 | -0.038947 | 5.1333
| -1.1861 | -0.4532 | 2.1921 | 6.4178
| -0.42157 | 1.4916 | 0.72381 | 2.9265
| -0.98309 | 1.4152 | -0.53077 | 3.251
| -0.52177 | 0.7795 | -0.90216 | 1.6938
| 0.32944 | 0.12659 | -1.2491 | 1.6848
| 0.80067 | 0.68311 | -1.561 | 3.5444
| 0.65587 | 1.2367 | 0.90615 | 2.7808
| -2.0036 | 0.88831 | 0.7557 | 5.3744
| -0.92845 | 0.78214 | 0.038918 | 1.4753
| 0.24538 | -0.94385 | 0.19078 | 0.98746
| 0.38813 | 0.42149 | -1.9526 | 4.1409
| 0.11744 | 0.73674 | 0.61176 | 0.93082
| -0.54261 | 1.0428 | 2.2701 | 6.5353
| 0.040102 | -1.2817 | -0.97917 | 2.6031
| 0.056007 | -1.9238 | -0.39712 | 3.8619
| -1.2967 | -2.2027 | -0.17595 | 6.5641
| 0.71876 | -0.34351 | -0.5264 | 0.91172
| -0.36452 | -0.94584 | 0.4542 | 1.2338
| -0.055471 | 0.63293 | -0.9836 | 1.3711
| -0.80323 | -1.057 | -0.8184 | 2.4322
| 0.92808 | -1.1906 | 1.2951 | 3.9562
| -0.4466 | 0.14179 | 0.17135 | 0.24892
| -0.58942 | -0.37597 | -0.050363 | 0.49131
| 0.3359 | -0.63906 | 1.0454 | 1.614
| -2.838 | 0.63486 | 2.8606 | 16.6402
| 0.03785 | -1.1481 | -0.13935 | 1.339
| -0.67101 | -0.071971 | -0.34396 | 0.57374
| 0.38863 | -1.4305 | -0.95679 | 3.1127
| -0.58357 | -0.3968 | -0.22038 | 0.54657
| -2.4311 | -0.1576 | 0.23159 | 5.9889
| -1.475 | -0.42897 | 1.027 | 3.4143
| 0.09955 | -2.0313 | -1.867 | 7.6221
| -0.23109 | -0.3757 | 0.56474 | 0.51349
| 0.49203 | 1.8503 | -0.49618 | 3.912
| -0.5299 | 0.040432 | -1.6127 | 2.8833
| -1.1873 | 1.1762 | 1.5292 | 5.1317
| 0.82823 | 0.30681 | 2.1336 | 5.3324
| -0.18271 | 0.63765 | 1.7368 | 3.4565
| -0.24459 | -0.55105 | 0.65496 | 0.79246
| -0.014256 | -0.33306 | 0.014956 | 0.11136
| -1.6167 | 1.1635 | -0.24287 | 4.0264
| -0.66296 | -0.80036 | -0.80358 | 1.7258
| 0.68619 | 1.2587 | -0.27265 | 2.1296
| -0.49088 | -1.282 | -0.80404 | 2.5309
| 0.50348 | 2.2863 | 0.49115 | 5.7218
| 1.2116 | 1.8554 | -0.55566 | 5.2192
| -0.51784 | 1.1274 | -1.0355 | 2.6115
| 1.0116 | -0.51484 | 0.27258 | 1.3627
| 0.48726 | 0.91312 | -1.3387 | 2.8632
| -0.60536 | 1.2225 | 1.2885 | 3.5212
| 0.094361 | 1.0335 | -0.62711 | 1.4703
| 0.058613 | -0.746 | 0.74376 | 1.1131
| 0.74995 | -0.01729 | -2.9987 | 9.5549
| -0.20825 | 1.5992 | -1.7811 | 5.7733
| -0.32758 | -0.32388 | 0.90677 | 1.0344
| 0.3596 | 1.9277 | -0.63318 | 4.2462
| 0.15432 | -0.042892 | 1.6835 | 2.8598
| -0.45317 | -1.5603 | 1.9315 | 6.3704
| 0.42542 | 0.035012 | -0.89976 | 0.99177
| -1.0709 | 1.2457 | 0.18017 | 2.7309
| -1.3847 | -0.44288 | 0.6265 | 2.5059
| 0.63488 | -0.78433 | -0.82697 | 1.7021
| 0.71743 | 0.14151 | 0.41049 | 0.70324
| 0.57327 | -1.4029 | -0.30566 | 2.3902
| -0.5183 | -0.52832 | -0.58872 | 0.89436
| 0.50839 | -0.15275 | 0.11736 | 0.29556
| 0.46055 | -0.42091 | -0.21814 | 0.43686
| 1.4118 | -0.26142 | -0.53863 | 2.3516
| 0.51215 | 0.90099 | 0.42919 | 1.2583
| 0.59825 | -0.38181 | 0.61792 | 0.88551
| -1.0389 | -0.26417 | 2.1398 | 5.728
| -0.11181 | 0.043234 | -0.16932 | 0.04304
| 0.5628 | -0.069794 | 0.20269 | 0.3627
| 1.1469 | 0.2656 | -0.57879 | 1.721
| -0.13367 | 0.22151 | -0.25467 | 0.13179
| 1.3553 | -2.529 | 0.19495 | 8.2706
| 0.16285 | -0.70958 | -0.10763 | 0.5416
| -0.79499 | 0.54185 | -0.4546 | 1.1323
| -1.1784 | 1.4771 | 1.2985 | 5.2564
| -0.88889 | -0.36886 | 0.18329 | 0.95979
| -0.41348 | -0.75023 | -0.44137 | 0.92863
| -0.88204 | 1.1359 | -0.56021 | 2.3822
| -0.4033 | -0.50451 | -1.8609 | 3.88
| 0.061568 | -0.39664 | -1.2676 | 1.768
| 1.486 | -0.35642 | 1.0974 | 3.5396
| 1.0289 | -0.078836 | -0.24872 | 1.1267
| 0.80648 | -0.96967 | -0.017619 | 1.591
| 0.37017 | -0.80096 | -0.081273 | 0.78516
| -0.1541 | -0.59838 | -0.77849 | 0.98785
| 0.078562 | 0.97536 | -0.77348 | 1.5558
| -0.93232 | -0.17916 | 0.7312 | 1.436
| -1.9816 | 0.66247 | -0.20121 | 4.4063
| -1.2831 | -1.3308 | -0.084021 | 3.4243
| 1.3849 | 0.76134 | 1.5275 | 4.8308
| -0.43428 | 1.5857 | 0.16963 | 2.7317
| 0.0080097 | 2.1626 | -0.058835 | 4.6805
| 1.6026 | 1.0476 | 0.11074 | 3.678
| -0.55269 | -0.30904 | -0.92691 | 1.2601
| 0.63534 | -0.71812 | -1.0939 | 2.1159
| -1.5985 | -2.2322 | 0.93881 | 8.4191
| -2.656 | -1.9742 | -0.76356 | 11.5351
| -0.36342 | 0.2985 | 0.21132 | 0.26583
| -0.22087 | -0.20585 | -1.1398 | 1.3903
| 0.65339 | 0.51987 | 1.4253 | 2.7286
| -0.24357 | 0.38413 | 1.8234 | 3.5315
| -0.68876 | -0.5539 | -0.079355 | 0.78749
| -0.066291 | 0.081419 | 0.008665 | 0.011099
| 1.3781 | 0.4784 | 0.045533 | 2.13
| -2.4604 | 1.9504 | 0.7197 | 10.376
| 1.0771 | -1.8631 | -3.1681 | 14.6683
| -0.66855 | 0.18038 | 1.7665 | 3.5999
| -0.12508 | -0.44673 | 0.54057 | 0.50743
| 0.82254 | 0.23113 | 1.401 | 2.6927
| -0.24547 | 0.58231 | -1.3595 | 2.2474
| 0.93518 | 0.27394 | 1.1986 | 2.3863
| 0.40918 | 0.84631 | -0.26753 | 0.95525
| 0.65192 | -0.33898 | -1.374 | 2.4279
| -0.82741 | -2.9627 | 0.2613 | 9.5307
| 0.70052 | 1.7926 | -0.11157 | 3.7165
| -0.48432 | -0.11544 | -0.54887 | 0.54915
| -2.2565 | 0.51319 | -0.75283 | 5.9219
| 0.87524 | -0.88237 | 0.22522 | 1.5953
| -1.3548 | 0.58908 | -0.5997 | 2.5422
| 1.1946 | -0.49974 | 1.3384 | 3.4681
| -0.45614 | -0.12128 | -0.51019 | 0.48307
| 0.81198 | 1.0803 | 0.993 | 2.8123
| -0.11852 | -0.49253 | 0.28256 | 0.33647
| -0.31881 | 0.36199 | 1.0661 | 1.3692
| -2.4742 | -1.2682 | 0.062623 | 7.734
| 0.93987 | -0.96274 | 1.328 | 3.5739
| -1.4952 | 0.0072746 | 0.96888 | 3.1743
| -0.65967 | 0.83227 | -1.2012 | 2.5708
| 0.14148 | -0.24093 | 0.49694 | 0.32501
| 0.90208 | 0.084417 | -0.41522 | 0.99328
| 1.2762 | 0.53789 | 2.0999 | 6.3275
| -1.7172 | -0.10127 | 2.4414 | 8.9192
| -1.2506 | 1.0324 | -0.018978 | 2.6302
| 0.72065 | 0.55589 | -2.0593 | 5.0693
| 0.81246 | 0.32099 | -0.61252 | 1.1383
| 0.58334 | 0.20183 | 1.1222 | 1.6403
| -0.86597 | 0.16566 | -0.70495 | 1.2743
| -0.66908 | 0.80605 | 0.91097 | 1.9273
| 1.8491 | -0.88989 | 0.34438 | 4.3297
| -0.25358 | 0.71299 | -0.21456 | 0.61869
| -0.01895 | -0.45116 | 0.07992 | 0.21029
| 0.34714 | -0.39813 | -1.5481 | 2.6755
| -1.2086 | 0.38731 | 0.16148 | 1.6368
| -0.10836 | 0.46857 | 1.5533 | 2.644
| -0.59589 | 1.0652 | -1.6818 | 4.3184
| 1.1569 | -1.2802 | 0.20074 | 3.0175
| -0.099719 | 0.68059 | -0.17677 | 0.50439
| 2.8665 | -0.38871 | -2.307 | 13.6897
| -1.328 | 0.92638 | 0.76987 | 3.2144
| 0.51771 | -0.85209 | 0.76687 | 1.5822
| 0.1234 | -0.62575 | -0.41965 | 0.5829
| 1.1739 | 0.43741 | 1.6716 | 4.3636
| 0.12523 | 1.6678 | 0.84771 | 3.5158
| -0.20469 | -0.045844 | 0.34776 | 0.16493
| -1.0063 | -0.49777 | 0.060734 | 1.2641
| -0.49928 | 0.56514 | 0.87325 | 1.3312
| -0.08443 | 2.1218 | 2.1823 | 9.2715
| 1.0204 | 1.1126 | 0.4606 | 2.4913
| 0.28755 | -0.29634 | 1.4022 | 2.1367
| 0.22061 | 2.1265 | -1.9156 | 8.2404
| -0.36972 | 0.099362 | 0.2542 | 0.21118
| -1.5892 | -0.73977 | -0.066452 | 3.0773
| 0.071404 | 0.50792 | 0.8195 | 0.93466
| 0.13452 | -0.14616 | -1.8164 | 3.3388
| -1.1282 | -0.92874 | -1.1053 | 3.3571
| 1.8786 | 0.19725 | -1.0848 | 4.745
| 0.38037 | 1.2577 | 0.27295 | 1.801
| 0.60967 | 2.5652 | 1.4642 | 9.0957
| -0.97287 | -2.0875 | 0.15177 | 5.3272
| -0.25834 | -0.17295 | -0.04023 | 0.098269
| 1.3155 | -1.5468 | -0.97296 | 5.0699
| 0.94784 | -0.56254 | -0.69028 | 1.6913
| -1.5749 | -0.0095154 | 1.5857 | 4.9951
| -0.094548 | -2.2147 | 1.0837 | 6.0883
| -0.60618 | 0.23644 | 0.27912 | 0.50127
| 0.1059 | 0.073779 | 0.67424 | 0.47126
| -0.30088 | 1.5115 | 0.66862 | 2.8223
| 0.23178 | -0.03387 | -1.8803 | 3.5903
| 0.41175 | -0.10334 | -0.096524 | 0.18953
| 1.2856 | -2.0935 | -0.60926 | 6.4068
| -1.0899 | -1.4044 | 1.4356 | 5.2212
| -2.2486 | -0.038393 | 0.04163 | 5.0593
| -1.9266 | 0.35698 | -0.30198 | 3.9306
| 0.24175 | -0.88615 | -0.16009 | 0.86934
| -1.7387 | 0.79418 | -1.7001 | 6.5442
| -0.57638 | -0.43981 | 1.3114 | 2.2455
| 0.383 | 0.086479 | -0.35889 | 0.28297
| -1.5778 | -0.42853 | -0.78904 | 3.2958
| 0.78019 | -0.78724 | -1.9472 | 5.0202
| -0.66329 | -0.96669 | 1.1758 | 2.7568
| -0.0045055 | 0.091496 | -0.52595 | 0.28502
| 0.025513 | 0.46565 | -0.86089 | 0.95862
| -0.23775 | -0.83463 | -1.412 | 2.7467
| -0.28282 | 1.3199 | -0.96049 | 2.7447
| 1.1579 | 1.1024 | -0.71838 | 3.072
| -0.34616 | 0.21374 | -0.071715 | 0.17065
| 2.16 | -0.0207 | 0.37136 | 4.8041
| 0.62115 | 0.60563 | -2.1892 | 5.5452
| 0.55714 | -0.64668 | -0.99277 | 1.7142
| 0.70296 | 1.1974 | 0.70436 | 2.4241
| 2.4158 | 1.6239 | 0.54184 | 8.7667
| -0.51325 | -0.95154 | 0.10576 | 1.18
| -0.49438 | 0.50014 | -0.86793 | 1.2479
| -0.2338 | -0.1185 | -0.58514 | 0.41109
| 0.69476 | 0.52368 | 0.22251 | 0.80644
| -1.9364 | 1.3202 | 0.55795 | 5.8039
| -0.085213 | 0.23347 | 1.3592 | 1.9091
| -0.3165 | -0.2759 | -0.83516 | 0.87379
| -1.6482 | -0.35121 | -2.2376 | 7.8467
| 0.98485 | 1.2293 | 0.90486 | 3.3
| -1.4961 | 1.6523 | -0.91858 | 5.812
| -0.66988 | 0.68584 | -0.44026 | 1.113
| -1.7257 | -2.1911 | -0.5946 | 8.1326
| -1.0724 | -1.0624 | -0.10783 | 2.2903
| 0.62281 | 0.20578 | 0.55139 | 0.73427
| 0.86137 | -0.19895 | -0.47031 | 1.0027
| 0.99783 | -0.31429 | -0.28154 | 1.1737
| 1.0501 | 0.37846 | 0.92028 | 2.0928
| 0.63178 | 0.3396 | -0.64161 | 0.92614
| -0.13344 | 1.1148 | -0.095947 | 1.2697
| 1.4862 | 0.15378 | 0.016655 | 2.2327
| 1.5814 | -0.41393 | 0.098513 | 2.6818
| -0.46638 | -1.0323 | -1.0689 | 2.4258
| 0.21181 | 1.3832 | -0.52791 | 2.2367
| -0.4331 | 0.1476 | 1.0487 | 1.3092
| -0.30289 | 0.26187 | 0.93587 | 1.0362
| 0.392 | 0.53428 | 1.2051 | 1.8913
| -0.0010993 | -1.0462 | 1.0487 | 2.1942
| -0.9217 | -0.45786 | 1.3431 | 2.863
| 1.4663 | -0.69604 | 0.12004 | 2.6489
| 1.0666 | 0.3579 | 0.79634 | 1.8999
| 1.2316 | 0.65342 | 0.59471 | 2.2975
| -0.78066 | 0.24507 | 0.78424 | 1.2845
| 0.15534 | -1.7058 | -0.70405 | 3.4295
| 0.048269 | 0.13965 | 0.034059 | 0.022993
| 0.7273 | 1.6203 | -0.023266 | 3.1548
| 0.89108 | -0.281 | -0.87254 | 1.6343
| 1.1116 | -0.66052 | -0.68447 | 2.1405
| -0.013996 | 0.56747 | 0.49262 | 0.5649
| -0.26295 | -0.21992 | -0.78822 | 0.7388
| 0.19867 | 0.28145 | -0.85587 | 0.85121
| -1.3356 | 0.51855 | 0.66276 | 2.4919
| 0.051853 | -2.1624 | 0.54833 | 4.9793
| 1.1219 | -0.64464 | 0.38904 | 1.8255
| 0.9972 | -0.48896 | 0.56012 | 1.5472
| -0.45733 | -0.59207 | 0.076127 | 0.56549
| -0.0072871 | 0.36853 | 0.32446 | 0.24114
| -0.92126 | 1.4334 | 0.2892 | 2.9869
| -0.026854 | 0.75636 | 0.068018 | 0.57742
| -1.5011 | -0.75565 | 0.62186 | 3.2111
| 0.24957 | -1.0282 | -1.4489 | 3.2187
| 0.53751 | 1.4544 | -0.31691 | 2.5045
| -1.7555 | 0.055508 | 0.2925 | 3.1704
| -2.0611 | 0.065349 | -1.4141 | 6.252
| 1.3764 | -0.81534 | 0.35356 | 2.6844
| 1.0079 | 0.23813 | -0.76598 | 1.6594
| -0.85302 | -1.7949 | -1.075 | 5.1051
| 0.57061 | -0.25025 | -0.36421 | 0.52087
| -0.78761 | -0.95229 | 1.0853 | 2.705
| -0.68149 | 0.50655 | 1.0497 | 1.8229
| 1.1002 | 0.38328 | -3.7311 | 15.2784
| -0.40512 | 0.56706 | -0.074908 | 0.49129
| -1.2023 | -2.2611 | -1.773 | 9.7017
| -0.5699 | -0.77235 | 0.64902 | 1.3425
| -0.77317 | -0.67263 | 1.8809 | 4.5882
| 1.782 | 0.13634 | -1.6627 | 5.9587
| 0.74743 | 0.29553 | -1.256 | 2.2236
| 0.36301 | 0.31693 | 0.60872 | 0.60276
| 3.3944 | -1.3922 | 1.2388 | 14.995
| 0.44645 | -2.3835 | 1.665 | 8.6527
| -0.48295 | 1.5541 | 1.2461 | 4.2013
| 0.82615 | -0.17145 | -0.14214 | 0.73212
| -0.86822 | -0.56715 | -0.48149 | 1.3073
| 1.6631 | -0.60338 | -0.65448 | 3.5584
| -0.4593 | -1.0509 | 1.0477 | 2.413
| 0.28647 | -1.9916 | 0.86235 | 4.792
| -0.87306 | 0.075397 | 1.6817 | 3.5959
| -1.6151 | 0.81386 | 0.40182 | 3.4324
| 0.011386 | -1.9495 | 0.51785 | 4.0687
| 0.63142 | 0.86951 | -0.11385 | 1.1677
| -2.5556 | -0.81348 | -1.0301 | 8.2537
| 0.72256 | 0.64033 | 0.5484 | 1.2329
| -0.52361 | -0.6609 | 1.9987 | 4.7056
| -0.21791 | -1.3332 | 0.77941 | 2.4323
| 0.35666 | 0.52537 | 0.43157 | 0.58947
| 1.0425 | 1.4323 | 0.11521 | 3.1514
| 1.0892 | -1.5906 | -0.32531 | 3.8221
| -0.98535 | 1.2041 | -1.1408 | 3.722
| -1.5883 | 0.80016 | 1.7079 | 6.0798
| -0.1094 | -2.5262 | 0.43828 | 6.5859
| 0.48079 | -0.078962 | -0.49909 | 0.48648
| 0.16734 | 0.77224 | -1.1957 | 2.0541
| 1.359 | -0.96168 | -0.12755 | 2.7879
| -1.2354 | 2.3918 | 0.18271 | 7.2803
| -0.66064 | -0.21516 | 0.69947 | 0.97199
| -0.49411 | -0.19683 | 0.21525 | 0.32922
| 0.34915 | 2.4107 | -0.5792 | 6.2688
| 0.80907 | 1.3805 | -0.45124 | 2.7641
| 0.63664 | 0.43992 | -0.56243 | 0.91516
| 0.18732 | 2.0141 | 0.70466 | 4.5882
| -0.79766 | -1.6874 | 0.24242 | 3.5425
| -0.60833 | -1.2732 | 0.13236 | 2.0086
| 1.5705 | -0.63124 | -0.14032 | 2.8846
| -1.8772 | -1.4823 | -0.22914 | 5.7735
| 0.56999 | 0.97546 | -1.289 | 2.9379
| 0.53568 | -0.086974 | -0.033399 | 0.29564
| 0.36945 | -0.44722 | 0.26309 | 0.40572
| 0.68736 | 0.56919 | 0.40648 | 0.96167
| 0.52313 | -0.74295 | -1.3868 | 2.7488
| -0.10899 | 0.31991 | -1.0036 | 1.1214
| -0.076534 | -0.30643 | -0.59844 | 0.45788
| 2.202 | -0.061783 | 0.47862 | 5.0815
| -0.28719 | 1.1434 | -0.40663 | 1.5552
| -0.45839 | -3.6652 | -0.68977 | 14.1198
| 0.8075 | 0.54045 | -0.53409 | 1.2294
| 2.4594 | 0.027538 | 1.27 | 7.6625
| -1.6968 | 0.88329 | 1.2557 | 5.2362
| 2.1421 | -0.63817 | 0.46674 | 5.2136
| -0.65891 | -0.061797 | -0.21905 | 0.48596
| 0.3829 | -0.33163 | -1.2501 | 1.8193
| -1.6502 | -0.32199 | -0.2907 | 2.9114
| 0.88469 | 0.96193 | 1.2347 | 3.2325
| 1.5711 | -0.26537 | 1.1091 | 3.7689
| -0.97809 | -0.4282 | 1.8911 | 4.7162
| -0.7863 | 0.56056 | 1.621 | 3.5603
| 0.92761 | -1.3909 | 1.199 | 4.2328
| 0.80701 | 0.025336 | 1.9102 | 4.3009
| 1.945 | -1.2631 | 1.8322 | 8.735
| 0.33689 | 1.5593 | -0.955 | 3.457
| -1.8175 | -0.47811 | 0.66757 | 3.9774
| 0.58989 | -1.4597 | -1.6861 | 5.3217
| -0.33385 | -0.95953 | -2.7081 | 8.3659
| -0.46363 | -0.3893 | 0.37721 | 0.50879
| -1.3101 | 0.99355 | 1.0391 | 3.7833
| 1.9566 | 0.19742 | -0.94774 | 4.7656
| -0.0054262 | 0.37926 | 0.37964 | 0.288
| -1.412 | 0.23875 | 0.4066 | 2.2162
| 0.66213 | -0.74535 | -0.74928 | 1.5554
| 0.25074 | -0.033653 | -0.86122 | 0.80571
| 0.42657 | -0.40766 | -0.7299 | 0.8809
| 0.27174 | -1.0364 | 0.20661 | 1.1907
| 0.44467 | -0.082242 | -0.71253 | 0.7122
| -1.2608 | 0.85617 | 1.0367 | 3.3975
| -1.1108 | 0.41025 | -0.81009 | 2.0584
| 0.41059 | -0.74746 | 0.043978 | 0.72921
| 0.18097 | 1.0503 | 0.99608 | 2.1281
| -1.4201 | -0.50666 | 0.37023 | 2.4103
| -2.3067 | -0.18633 | -0.7649 | 5.9406
| 0.48302 | -0.59666 | 0.30378 | 0.68159
| 1.1122 | 1.6476 | 1.3678 | 5.8227
| 0.57712 | -1.0036 | 0.63527 | 1.7438
| 1.0643 | -0.17911 | 0.13896 | 1.1842
| -2.136 | -0.41546 | -0.057284 | 4.7384
| -0.28055 | -0.96662 | 0.44302 | 1.2093
| -0.92543 | -0.49103 | 2.3796 | 6.7598
| -2.5048 | -1.4032 | 0.60783 | 8.6125
| -0.9973 | 0.78245 | 1.3397 | 3.4016
| 1.2947 | 0.73524 | -0.13803 | 2.2358
| -0.48643 | 0.93076 | 0.42054 | 1.2798
| -0.23252 | 0.5432 | -0.98587 | 1.3211
| 0.91194 | 0.90209 | 1.1663 | 3.0056
| 0.34321 | -0.76243 | 0.70602 | 1.1976
| 0.54985 | 1.3992 | 1.3215 | 4.0064
| 0.64183 | -0.48324 | -0.43617 | 0.83571
| 2.1698 | -0.39711 | -0.29992 | 4.9557
| -0.38265 | 0.64946 | 0.71002 | 1.0723
| 0.65846 | 1.7346 | 1.4354 | 5.5027
| 1.4378 | -1.6064 | -0.14307 | 4.6685
| -1.4043 | 1.1944 | 0.16598 | 3.4262
| 0.023605 | -1.1525 | 0.67071 | 1.7786
| -0.87026 | -0.081439 | -0.82989 | 1.4527
| -0.16756 | 0.29415 | 0.56667 | 0.43571
| 0.098453 | -0.36896 | -2.5015 | 6.4033
| 0.010884 | -0.52967 | 0.94382 | 1.1715
| -0.33309 | 1.6713 | 0.059026 | 2.9077
| 0.43382 | -0.42767 | 0.7146 | 0.88176
| 0.23606 | -0.047397 | -0.31664 | 0.15823
| -0.3439 | -0.47301 | 1.6379 | 3.0246
| -1.8542 | -0.56617 | -0.72678 | 4.2867
| -1.8086 | 0.14373 | -2.0227 | 7.3831
| -0.93676 | -0.75372 | -1.0313 | 2.5092
| 0.58511 | -0.59731 | 0.030001 | 0.70002
| -0.1279 | 1.8499 | 0.639 | 3.8468
| -1.8983 | 0.84472 | -0.55877 | 4.6293
| 0.50815 | 0.17138 | -0.94293 | 1.1767
| -0.46555 | -0.32598 | 0.14172 | 0.34308
| -0.40024 | -1.2168 | -1.7709 | 4.7768
| -0.63249 | 0.68742 | -1.3652 | 2.7364
| -1.3886 | 0.11497 | 0.12057 | 1.956
| 0.61719 | -1.0388 | 2.4195 | 7.3141
| -0.66402 | -0.84485 | 0.01056 | 1.1548
| 0.1672 | 0.29394 | -0.0015169 | 0.11436
| -1.7323 | -1.0455 | -0.65482 | 4.5226
| 0.74016 | -0.34785 | 1.0337 | 1.7374
| -0.47482 | 1.5694 | -0.20782 | 2.7318
| 0.60506 | 0.78738 | 0.73979 | 1.5334
| 0.63247 | 1.1991 | -0.59395 | 2.1905
| -0.60047 | 0.763 | 0.64616 | 1.3603
| 0.10453 | -1.1252 | -0.55734 | 1.5876
| 1.2669 | -0.99481 | -0.18938 | 2.6305
| -0.0031666 | -0.86416 | -0.40248 | 0.90877
| -0.17609 | 0.45723 | -1.0518 | 1.3463
| 0.54029 | 0.36946 | 0.26352 | 0.49785
| 2.0824 | 0.70113 | -1.8826 | 8.3723
| -2.4023 | 0.35525 | 0.23183 | 5.9508
| -1.2483 | -0.36175 | -0.2716 | 1.7628
| -0.95169 | -0.047302 | -0.24739 | 0.96915
| 0.22084 | -0.4183 | -0.44403 | 0.4209
| 0.67977 | 0.77 | 2.8988 | 9.4579
| 3.0277 | -0.12954 | -0.30278 | 9.2754
| -0.28238 | -0.44174 | 0.0067435 | 0.27492
| -0.56778 | 1.0411 | 1.088 | 2.59
| 0.62997 | -0.99393 | 0.72779 | 1.9144
| -1.4098 | 1.1791 | -0.69948 | 3.8672
| 0.36167 | -0.42159 | 0.14354 | 0.32915
| -0.69819 | -0.7492 | 0.67401 | 1.5031
| -0.93253 | 1.8424 | 0.0093211 | 4.2641
| 0.87705 | 0.71911 | 0.49349 | 1.5299
| -0.49248 | -1.3292 | 0.27662 | 2.0858
| -0.38008 | 2.0181 | -0.82641 | 4.9002
| -0.48716 | -0.80596 | -0.19173 | 0.92366
| 0.33377 | -0.65857 | 0.5473 | 0.84466
| -0.71957 | -2.2058 | 0.68037 | 5.8462
| 0.10721 | -1.9668 | 0.54726 | 4.1794
| 0.25368 | 1.673 | -0.80085 | 3.5048
| -0.46689 | 0.88205 | 0.92417 | 1.8501
| 0.81595 | 1.3485 | 1.0818 | 3.6543
| -0.8134 | -0.53634 | -1.3375 | 2.7382
| 0.19979 | -0.29713 | -1.7643 | 3.2411
| 1.3002 | -0.94495 | -0.27421 | 2.6587
| -0.71183 | 0.62371 | 0.38383 | 1.043
| 1.8833 | 0.48786 | 0.78389 | 4.3995
| 1.3714 | 1.3079 | 0.36935 | 3.7277
| 0.23953 | 0.62216 | 1.4811 | 2.6382
| 0.78521 | -0.57444 | -1.8971 | 4.5455
| 0.17089 | -0.16881 | 0.034543 | 0.058893
| 0.079757 | 0.29641 | -1.1124 | 1.3317
| -0.32692 | 0.30473 | -0.20213 | 0.24059
| -0.39659 | -0.14763 | 0.29934 | 0.26868
| 0.22192 | -1.062 | 1.7641 | 4.2891
| 0.81764 | 0.47404 | 1.3092 | 2.6074
| -0.056386 | -1.6397 | -0.28935 | 2.7754
| 1.6113 | -0.13058 | 1.0832 | 3.7867
| 0.42364 | 0.85032 | -0.76936 | 1.4944
| 0.7876 | 0.96094 | 0.14938 | 1.566
| -0.31622 | 0.48246 | 0.18111 | 0.36556
| -0.29666 | -0.045025 | -0.88189 | 0.86776
| 1.2843 | -1.391 | 0.27303 | 3.6589
| 0.16079 | 0.49071 | -0.030737 | 0.26759
| 0.78739 | 0.98994 | -1.9071 | 5.237
| -0.52749 | 0.57065 | -0.55974 | 0.9172
| 0.59205 | -0.8184 | 0.17385 | 1.0505
| 1.1725 | 0.131 | 1.1383 | 2.6877
| -0.56515 | 0.79662 | 1.3693 | 2.8289
| -0.019649 | -0.82347 | -1.3307 | 2.4493
| -1.3522 | 1.4556 | 1.5409 | 6.3218
| -1.5757 | -0.72315 | 0.25581 | 3.0711
| -1.591 | -2.0526 | -0.59355 | 7.0967
| 0.77767 | 0.76246 | 0.024507 | 1.1867
| 1.4085 | -2.0705 | -0.65487 | 6.6997
| 0.88824 | 0.14154 | -0.57291 | 1.1372
| 0.87285 | 0.96137 | -3.3052 | 12.6105
| 0.61151 | 1.5926 | 0.52216 | 3.1831
| 1.0884 | 0.77009 | -0.71749 | 2.2925
| -0.45444 | -1.3412 | 0.034704 | 2.0065
| 0.37426 | 0.0046973 | 1.9323 | 3.8737
| -0.46328 | 1.3385 | -0.18476 | 2.0403
| 1.3282 | 0.4318 | -1.0096 | 2.97
| 0.91128 | 0.52619 | -0.98032 | 2.0684
| -0.34088 | 2.1636 | 2.4448 | 10.7743
| 0.08929 | -0.85129 | 1.9413 | 4.5014
| -0.28798 | -0.63947 | -0.38995 | 0.64391
| -1.2818 | -0.83139 | -1.1942 | 3.7602
| 0.22476 | -1.1173 | -0.61663 | 1.6792
| -0.3818 | -0.23973 | -0.58088 | 0.54067
| 0.61421 | -0.7451 | 1.0315 | 1.9965
| -1.9428 | -0.61006 | 0.48851 | 4.3853
| -0.45298 | -1.3435 | 1.8885 | 5.5766
| -0.57003 | -0.3406 | -0.78769 | 1.0614
| 0.90231 | -0.0095319 | -0.44218 | 1.0098
| -1.0777 | -1.9155 | -0.076344 | 4.8365
| 1.5019 | -1.3884 | 0.54826 | 4.4839
| 0.042975 | -1.77 | -0.077176 | 3.1406
| -0.08232 | 1.0939 | 0.65258 | 1.6293
| 0.162 | -0.39292 | -0.90381 | 0.99751
| -0.38045 | -0.73301 | -0.42689 | 0.86428
| 0.30566 | 0.34947 | -0.22136 | 0.26456
| 1.2342 | -0.28146 | -0.51656 | 1.8693
| 1.1066 | 0.029684 | 1.3314 | 2.9979
| 1.0001 | 0.19155 | 0.70486 | 1.5338
| -0.81816 | -1.2642 | 0.58331 | 2.6078
| 0.84283 | -1.3692 | 0.40608 | 2.75
| 0.17363 | 1.0948 | -1.9028 | 4.8492
| 0.43054 | -1.5189 | -0.061439 | 2.4961
| 0.086119 | 0.50983 | 1.3753 | 2.1589
| -1.7328 | -0.70849 | -0.54098 | 3.7972
| -0.63204 | 0.47652 | -0.8911 | 1.4206
| 0.26215 | 1.2378 | -0.308 | 1.6958
| 0.35542 | -0.2562 | -0.95895 | 1.1115
| 1.4576 | -0.53059 | 0.77662 | 3.0093
| -0.52247 | 2.7418 | 1.2356 | 9.3172
| -0.96717 | 0.28438 | 0.10026 | 1.0263
| 0.28245 | -0.68147 | 1.8756 | 4.062
| -0.40992 | 0.45513 | -0.39497 | 0.53117
| 0.98037 | 1.0433 | -0.16554 | 2.0769
| 0.29124 | -0.21071 | 0.18027 | 0.16171
| -0.20846 | 0.61736 | 0.01161 | 0.42472
| -1.519 | -0.69663 | -0.99841 | 3.7894
| 0.48866 | 1.4909 | 0.10615 | 2.4728
| -0.68789 | 0.025947 | -0.41963 | 0.64995
| -0.78977 | 1.4091 | 1.7129 | 5.5433
| 0.16603 | 0.87037 | -1.379 | 2.6868
| 0.72722 | -0.52972 | 0.42717 | 0.99194
| 0.67018 | 0.20904 | -0.10988 | 0.50491
| 1.8758 | -0.33885 | -1.5047 | 5.8975
| 0.29525 | -2.7049 | -0.38278 | 7.5499
| -1.2125 | -0.72477 | -1.5668 | 4.4503
| 0.834 | 2.2119 | -1.9699 | 9.4687
| 0.60878 | 0.76161 | 0.0065062 | 0.9507
| -0.26979 | -1.3106 | -1.548 | 4.1867
| -0.053845 | 1.0991 | 0.23366 | 1.2654
| -1.4792 | 0.08272 | -0.2034 | 2.2363
| 0.77316 | 0.56937 | -1.2607 | 2.5114
| -0.53523 | 1.0231 | -0.024339 | 1.3338
| -0.42116 | 0.81566 | -1.2417 | 2.3845
| -0.79974 | -1.2937 | 0.30118 | 2.4039
| 0.59027 | 0.92074 | 2.4056 | 6.9831
| -1.1632 | -0.2455 | 0.76719 | 2.0019
| -1.5762 | 0.34503 | 0.058361 | 2.6068
| 0.35372 | -1.3118 | -0.80904 | 2.5005
| -0.18015 | 0.27918 | -1.2995 | 1.799
| 0.77528 | -1.9331 | 0.68543 | 4.8077
| -0.099719 | 0.52215 | -0.70173 | 0.77501
| 0.13032 | 0.55505 | -0.83197 | 1.0172
| -1.3808 | 1.7163 | 1.2349 | 6.3774
| -0.15453 | -0.95462 | -0.49014 | 1.1754
| 0.33509 | 0.91725 | 0.27628 | 1.03
| -0.40808 | -0.6947 | 0.59217 | 0.99981
| 1.3441 | 1.961 | -0.15477 | 5.6759
| -2.1985 | 0.48306 | 1.1 | 6.2769
| -0.58207 | 1.0549 | 0.2111 | 1.4963
| -0.34566 | -1.0363 | 1.5525 | 3.6037
| 0.69826 | 1.4858 | 1.7805 | 5.8652
| 0.5601 | -2.5293 | 0.70632 | 7.2101
| 1.7249 | 0.36847 | -0.38168 | 3.2566
| -0.87933 | 0.056905 | 0.60135 | 1.1381
| 0.97522 | -0.1572 | 0.92542 | 1.8322
| 0.86807 | -1.6771 | -0.15467 | 3.5902
| 0.61467 | 0.36457 | 1.4741 | 2.6837
| 0.62304 | -0.53729 | -2.7918 | 8.471
| -0.70241 | -0.3592 | -0.73919 | 1.1688
| 0.62835 | 0.7818 | 0.77718 | 1.6101
| -1.6197 | 1.3689 | -0.94127 | 5.3834
| -0.40832 | 0.19312 | 0.92835 | 1.0659
| 1.5489 | 0.38219 | -0.55852 | 2.8571
| -0.55356 | -0.46632 | -0.67778 | 0.98327
| 0.15347 | -0.26642 | 0.55865 | 0.40663
| -0.70392 | -1.8517 | 1.6167 | 6.5382
| 1.6044 | -0.15481 | -1.8465 | 6.0077
| -0.079304 | -0.063372 | 0.5462 | 0.30864
| 0.64729 | 1.5484 | -0.2137 | 2.8622
| 0.33823 | -1.2909 | 0.60017 | 2.1412
| -0.20154 | -0.88403 | 1.8419 | 4.2149
| 0.50028 | 0.2357 | 0.59709 | 0.66235
| -0.80456 | 1.3447 | -0.8678 | 3.2087
| 0.42322 | -0.72163 | -0.97943 | 1.6591
| 0.36672 | -1.6318 | -0.24296 | 2.8563
| 0.60912 | -0.049886 | -0.24729 | 0.43467
| 0.51475 | 1.3601 | 0.86832 | 2.8689
| 0.90712 | 0.86322 | -2.5155 | 7.8956
| -0.9956 | -0.82691 | -0.69107 | 2.1526
| -0.57025 | 0.080867 | 2.07 | 4.6166
| 0.76888 | -0.38924 | -0.72506 | 1.2684
| 0.30699 | 0.11616 | 0.65723 | 0.53969
| -1.1328 | -2.538 | -0.55645 | 8.0341
| 1.6396 | 0.73845 | 0.77346 | 3.8317
| -0.96037 | -0.80981 | -1.0704 | 2.7238
| 0.26644 | 0.71598 | 1.3963 | 2.5333
| 0.61356 | -0.69428 | 1.1412 | 2.1607
| 1.3439 | 0.16386 | -1.8179 | 5.1378
| 0.94159 | 0.0060958 | -0.98745 | 1.8617
| 1.044 | -1.0695 | 0.64927 | 2.6552
| -0.15456 | 0.53459 | -0.82565 | 0.99138
| -0.18209 | 2.4425 | -1.3032 | 7.697
| -1.423 | 0.25492 | 2.2159 | 7.0003
| -0.34948 | -0.21977 | -1.135 | 1.4586
| -1.0777 | 0.97845 | -1.8926 | 5.7007
| 0.63462 | -0.70349 | -0.11377 | 0.91059
| -1.4301 | -0.43056 | 0.7215 | 2.7511
| 0.41407 | -0.26435 | 0.29653 | 0.32927
| -0.60688 | -0.33287 | 0.5112 | 0.74043
| 1.2734 | 0.7889 | -0.052678 | 2.2466
| -0.27339 | 0.033072 | -0.38491 | 0.22399
| 0.074388 | 0.27519 | 0.35834 | 0.20967
| 0.68539 | 1.6073 | 0.22079 | 3.1021
| -0.81182 | -1.9614 | -0.712 | 5.0132
| -0.32757 | 0.093173 | -0.70824 | 0.61759
| -0.49229 | 0.45844 | -1.2592 | 2.0381
| 0.42287 | 0.37372 | -0.021105 | 0.31893
| -0.60556 | 1.9041 | 2.1957 | 8.8135
| 0.22918 | -0.28852 | -0.4291 | 0.31989
| -0.11272 | -0.31133 | 1.8193 | 3.4194
| 0.27338 | 0.27954 | 0.76749 | 0.74192
| 0.5081 | 0.54669 | -1.0586 | 1.6777
| -1.0833 | 1.115 | -0.51896 | 2.6861
| -2.0119 | -0.038784 | -1.5424 | 6.4281
| -0.078864 | -0.45537 | -0.41797 | 0.38828
| 0.96899 | -0.30656 | -0.13731 | 1.0518
| -0.0046367 | -1.4577 | -1.2475 | 3.6813
| 1.6026 | 2.7825 | 0.22032 | 10.3589
| -0.031576 | -0.76975 | -0.19072 | 0.62989
| 1.3271 | -0.015531 | 1.4638 | 3.9044
| 0.15607 | 0.45628 | 2.3692 | 5.8457
| 1.2394 | 0.58513 | 0.12012 | 1.8929
| -1.1203 | -0.65019 | 0.81786 | 2.3466
| -0.039803 | -0.12645 | -0.75831 | 0.5926
| -1.737 | -0.821 | -0.27401 | 3.7665
| 0.63737 | 2.1683 | 1.1179 | 6.3574
| 0.22638 | 1.6793 | -1.0076 | 3.8865
| 0.12535 | 0.87681 | 0.43577 | 0.9744
| 0.38378 | 1.6221 | 0.53659 | 3.0664
| -0.36243 | -1.9407 | 0.077776 | 3.9036
| 0.74059 | -0.098038 | 0.74377 | 1.1113
| -0.38141 | -1.8422 | 2.6107 | 10.3552
| 0.43179 | 0.28902 | 0.12921 | 0.28667
| -0.082875 | -0.1294 | 1.1342 | 1.3099
| 0.77373 | -0.49604 | -0.56904 | 1.1685
| -0.24625 | -0.53595 | 0.91684 | 1.1885
| -1.374 | -0.2506 | -1.4837 | 4.1519
| -2.4472 | 0.96469 | -0.60184 | 7.2815
| -0.18415 | -0.19366 | 0.14633 | 0.092828
| 1.9858 | -0.18969 | -1.3536 | 5.8117
| -0.03038 | -0.88473 | -0.047338 | 0.78592
| -0.24042 | 1.1161 | 0.51549 | 1.5693
| -1.1965 | 0.77762 | -0.25344 | 2.1006
| -0.34076 | 1.4372 | 0.90608 | 3.0027
| 1.1095 | -0.55701 | 1.2551 | 3.1164
| 1.4075 | 1.5708 | -0.25424 | 4.5132
| 0.73452 | 1.0456 | 0.13606 | 1.6513
| -0.70313 | 0.13245 | 0.59034 | 0.86043
| -0.41916 | 1.093 | -0.20576 | 1.4126
| -0.12268 | 1.3168 | -0.82545 | 2.4304
| 0.25637 | 0.38299 | -0.20826 | 0.25578
| -0.95752 | 1.6619 | 1.0852 | 4.8564
| 0.60011 | 0.26022 | -1.1394 | 1.7261
| 0.66661 | 1.8822 | -0.96865 | 4.9253
| 0.30361 | 0.64852 | 0.16449 | 0.53982
| -0.71069 | -0.11655 | -0.13613 | 0.5372
| 0.9425 | -1.0193 | 0.72619 | 2.4546
| 1.9393 | -0.036974 | 0.095715 | 3.7715
| 0.28592 | -0.29214 | 0.90747 | 0.99058
| 1.0135 | -1.1045 | 0.32383 | 2.352
| 0.25992 | -0.022941 | 0.87726 | 0.83766
| -0.45563 | -0.5976 | 0.05221 | 0.56745
| -0.62869 | -1.3588 | -1.1964 | 3.6729
| 1.1062 | -0.71412 | -0.1632 | 1.7603
| 1.2695 | 0.0054045 | -0.77158 | 2.207
| -0.77408 | 0.45858 | -0.055201 | 0.81254
| 0.81055 | -0.88301 | -2.1884 | 6.2259
| 0.20228 | 0.40751 | 0.090202 | 0.21511
| -1.1326 | 0.64849 | 0.89661 | 2.5073
| 1.5174 | -0.18528 | -0.1586 | 2.362
| 1.6314 | -0.46052 | 0.11165 | 2.8861
| -0.38382 | 0.32013 | 0.22831 | 0.30193
| 0.98224 | -0.81688 | -1.5152 | 3.9278
| -1.0324 | 0.94932 | 0.15581 | 1.9914
| 0.70566 | 1.0903 | -0.90252 | 2.5013
| 1.569 | 1.1175 | 1.1172 | 4.9584
| -1.4077 | 0.58029 | -1.155 | 3.6525
| 0.025494 | 0.50809 | 0.97058 | 1.2008
| -1.366 | 0.39005 | -2.0705 | 6.3049
| 2.1086 | 1.208 | -0.60554 | 6.2723
| -0.12385 | 0.096758 | 0.80488 | 0.67254
| 1.4616 | 0.59992 | 0.26987 | 2.5689
| -0.96434 | 1.5648 | 0.084741 | 3.3857
| 0.03225 | -0.13224 | 1.5824 | 2.5225
| -0.50041 | 0.66533 | 1.4704 | 2.8553
| 0.66584 | 0.79974 | 0.7209 | 1.6026
| -0.087516 | 0.60239 | -0.065218 | 0.37478
| -0.33328 | -0.78179 | -0.15489 | 0.74626
| -2.2746 | -0.16261 | 1.4782 | 7.3852
| 0.10101 | -1.2375 | -0.2268 | 1.593
| -0.28535 | -2.1084 | 0.55136 | 4.8308
| 0.77279 | -0.20222 | -1.1547 | 1.9715
| -1.445 | -0.37778 | -0.10882 | 2.2424
| 0.11532 | -1.1142 | 0.31569 | 1.3543
| 2.2634 | -1.5462 | 0.72865 | 8.0447
| 0.3982 | 0.47386 | 1.0138 | 1.4109
| -1.0591 | 1.222 | -1.9303 | 6.3412
| -0.76564 | -1.2117 | -2.1655 | 6.7436
| -0.025823 | 0.58247 | 0.32882 | 0.44806
| -1.0614 | -1.1714 | -0.67967 | 2.9606
| -1.2534 | -1.2924 | 0.71211 | 3.7485
| 1.1556 | -0.97866 | 0.47176 | 2.5157
| 2.1516 | 0.35146 | -0.60548 | 5.1194
| -1.173 | -0.29424 | -0.98821 | 2.4391
| 1.7655 | -1.1017 | 0.1821 | 4.3637
| -0.65827 | 0.6227 | 0.91326 | 1.6551
| -1.1775 | -0.10013 | -1.4187 | 3.4093
| 1.2465 | 0.70801 | -0.17312 | 2.085
| -0.12798 | -0.52801 | -0.84277 | 1.0054
| 0.41137 | 0.97779 | 1.5954 | 3.6705
| 0.92647 | 0.65499 | -1.7768 | 4.4445
| 1.4367 | 1.7208 | 0.26148 | 5.0938
| 1.1231 | -0.31383 | -0.22892 | 1.4122
| 0.70653 | -1.4492 | -1.7665 | 5.7199
| 0.63227 | -1.543 | -0.53354 | 3.0652
| -1.4182 | 1.315 | -0.27501 | 3.8163
| 0.53259 | -0.15259 | -1.4208 | 2.3255
| 1.4775 | 2.1845 | 0.36908 | 7.0913
| 0.14272 | 2.1119 | -0.99539 | 5.4711
| 1.1556 | 1.3983 | 0.43181 | 3.4773
| -1.8549 | -0.45883 | -0.27534 | 3.7271
| 0.22022 | -1.3733 | -1.1454 | 3.2464
| 0.3377 | -0.18563 | -1.1737 | 1.5261
| 0.89578 | -0.088268 | -0.45403 | 1.0164
| 0.7714 | -0.1429 | -1.0538 | 1.726
| -0.46767 | -0.93593 | 0.15648 | 1.1192
| -0.37643 | -1.3136 | 0.24253 | 1.926
| -0.2995 | -0.28122 | 0.18769 | 0.20402
| 0.29312 | -0.051416 | -0.25092 | 0.15152
| -0.36599 | -0.35881 | -2.2226 | 5.2024
| 0.6517 | 0.9254 | 0.94486 | 2.1738
| -0.04436 | 0.88863 | -0.41311 | 0.96229
| -0.62137 | 0.75889 | -1.6226 | 3.595
| -1.4675 | -0.24799 | 0.60244 | 2.578
| -0.010935 | 0.74541 | -1.782 | 3.7314
| 0.4037 | 1.1589 | 0.38258 | 1.6523
| 0.42012 | 1.1364 | -0.61308 | 1.8437
| -1.8669 | -1.294 | 1.5548 | 7.5771
| 0.084629 | -1.8401 | 0.83031 | 4.0826
| -0.88244 | 2.9269 | 1.3305 | 11.1157
| 0.78929 | 0.027633 | -1.2787 | 2.2588
| 1.5327 | 1.0876 | 1.571 | 6
| 1.8115 | -2.1867 | -0.24972 | 8.1254
| 1.1126 | 0.75971 | -0.25411 | 1.8796
| -1.6003 | 0.38391 | 0.29199 | 2.7935
| -0.070254 | 0.26447 | -1.2228 | 1.5702
| -1.816 | 0.46855 | -0.32985 | 3.6261
| -0.084624 | 0.494 | -1.2884 | 1.9112
| -2.4873 | 0.17567 | 1.7947 | 9.4384
| 0.31665 | -0.2956 | -0.1261 | 0.20355
| -0.041075 | 0.56657 | -0.047098 | 0.32491
| 0.49226 | -1.6438 | 0.42887 | 3.1282
| -0.32424 | 0.50036 | -1.849 | 3.7744
| -0.05424 | 0.93077 | 1.1043 | 2.0888
| -0.54305 | 1.161 | 0.020916 | 1.6433
| -0.65775 | 0.30126 | -1.3095 | 2.2382
| -1.6546 | -0.65545 | 2.4454 | 9.1474
| -0.27178 | 0.40864 | 0.37445 | 0.38107
| -1.6178 | -1.2191 | -1.5183 | 6.4087
| 0.16135 | 0.02377 | -0.50958 | 0.28627
| -0.5301 | -0.22344 | 0.78662 | 0.9497
| -1.4937 | 0.80981 | -1.4523 | 4.996
| 0.3032 | 0.33249 | 1.0882 | 1.3866
| -0.26856 | -0.55468 | -0.36187 | 0.51075
| -1.5409 | -0.78339 | -1.0541 | 4.0994
| -0.69232 | 0.44241 | -0.44821 | 0.87593
| 0.39017 | -0.39696 | -0.29096 | 0.39447
| 1.6661 | 1.2268 | 0.20678 | 4.3236
| -1.0014 | -1.2043 | 2.0213 | 6.5386
| 0.81781 | -1.3828 | -0.38665 | 2.7306
| 1.0803 | -0.56345 | 0.55133 | 1.7884
| 1.4903 | 0.80914 | -0.81579 | 3.5413
| -0.53444 | -0.2468 | 0.60102 | 0.70775
| -0.23077 | -0.36263 | -0.09546 | 0.19386
| -1.0289 | 0.017435 | -1.4509 | 3.164
| 0.76122 | -0.45029 | 0.011531 | 0.78236
| -0.94558 | 1.6635 | -0.86423 | 4.4083
| 0.4553 | 0.54266 | 0.98682 | 1.4756
| 1.664 | -1.7779 | 0.31745 | 6.0307
| 0.77082 | -1.5923 | -0.26554 | 3.2
| -0.3037 | 0.40203 | 0.23585 | 0.30949
| -0.34403 | -1.4151 | 1.5831 | 4.6272
| -0.26374 | 0.89473 | -1.4706 | 3.0327
| 1.0153 | 0.035739 | -0.70174 | 1.5245
| -0.93639 | 1.1763 | -1.6168 | 4.8746
| 0.16776 | -0.29296 | 1.0678 | 1.2541
| -0.06934 | -0.77114 | 1.3465 | 2.4125
| 0.53605 | 0.12676 | 0.042951 | 0.30526
| 0.82264 | 2.1876 | 0.082174 | 5.4691
| -0.54288 | -0.39501 | -0.18255 | 0.48407
| 0.14687 | 0.31737 | 0.15472 | 0.14623
| 0.32476 | -0.54338 | -0.36763 | 0.53589
| -0.61895 | 0.081816 | 0.49104 | 0.6309
| 0.10676 | 0.42308 | -2.5221 | 6.5514
| -0.1888 | 1.021 | -0.25304 | 1.1421
| 0.41546 | -0.35989 | -0.82567 | 0.98385
| 0.26101 | 0.60492 | 0.92234 | 1.2848
| 2.4096 | 1.1757 | 0.9938 | 8.1761
| -0.1127 | 0.060875 | 0.5089 | 0.27538
| -1.3013 | -0.41216 | 0.16448 | 1.8903
| 0.092772 | -0.11573 | -0.15149 | 0.044948
| -1.8631 | 1.7238 | 2.072 | 10.7361
| -0.31745 | -0.055639 | -0.83386 | 0.7992
| -0.59491 | 0.054968 | 1.1158 | 1.6019
| -1.8018 | -0.47566 | -0.85898 | 4.2106
| -0.56867 | -0.14875 | 0.030424 | 0.34644
| -1.5902 | -0.24607 | 0.091991 | 2.5977
| -0.0556 | -0.59048 | 0.18567 | 0.38623
| -2.2067 | 1.5063 | -0.92006 | 7.985
| -0.73903 | -1.3435 | -0.16194 | 2.3774
| -1.7151 | 0.58256 | 0.70372 | 3.7761
| 1.211 | 1.3053 | -0.81088 | 3.8279
| 0.14257 | -0.71137 | -1.0437 | 1.6156
| 0.56379 | 1.7433 | 0.62425 | 3.7466
| -1.1838 | 1.224 | 0.19338 | 2.937
| 0.89233 | 0.83853 | 0.032838 | 1.5005
| -1.2458 | -0.86813 | 0.62762 | 2.6996
| 0.82458 | -1.5356 | 1.8458 | 6.4451
| 0.026827 | -0.065903 | -0.89332 | 0.80309
| 1.0393 | -0.12018 | -1.5955 | 3.6405
| 2.2982 | 0.15959 | -0.58988 | 5.655
| -0.93658 | -0.2223 | 0.92371 | 1.7798
| -0.22706 | -0.1901 | -0.25937 | 0.15496
| 0.93819 | -1.7564 | 1.0595 | 5.0878
| 0.5592 | 1.6152 | 0.63613 | 3.3263
| 1.9342 | -0.097542 | -1.6627 | 6.5151
| -0.72708 | -1.723 | 0.31961 | 3.5994
| 0.5091 | 0.049964 | 0.49007 | 0.50184
| -0.263 | -0.49194 | 0.42548 | 0.49221
| 2.0668 | -0.70725 | -0.78283 | 5.3848
| -1.3265 | 0.43574 | 0.27312 | 2.024
| 0.076277 | -1.5877 | 0.76696 | 3.1148
| 1.3792 | 0.12486 | 0.35947 | 2.047
| 1.3983 | 0.10335 | -1.0581 | 3.0855
| 1.0275 | 0.82214 | 0.75136 | 2.2962
| -0.45749 | 0.52843 | 1.5124 | 2.7758
| -1.3272 | -1.8428 | -1.0905 | 6.3464
| -0.4894 | 0.90676 | 0.19198 | 1.0986
| -1.2099 | 1.6211 | -0.6457 | 4.5086
| 1.03 | -0.32697 | -0.29371 | 1.254
| 0.52158 | 1.4253 | -1.1594 | 3.6479
| -0.043823 | -1.3745 | -1.0654 | 3.0263
| 1.0131 | -0.36343 | -0.50063 | 1.4091
| -0.20682 | 0.18164 | 1.4928 | 2.3041
| 0.30248 | 0.5453 | 0.3857 | 0.53761
| -0.98313 | 1.1389 | 1.1672 | 3.6261
| -0.25629 | -0.29038 | 0.3809 | 0.29509
| 0.082215 | -0.20865 | -0.95404 | 0.96049
| -2.7553 | 0.47212 | -0.69118 | 8.292
| 0.45106 | 2.1804 | -0.7771 | 5.5615
| 1.4395 | 2.109 | 0.87846 | 7.2916
| 0.42004 | -1.5051 | -0.37577 | 2.583
| 0.20785 | 0.4111 | 1.5637 | 2.6575
| 1.0918 | 0.25468 | 0.1118 | 1.2693
| 1.2758 | -0.36754 | 1.5481 | 4.1594
| 0.6325 | -0.79847 | 0.5101 | 1.2978
| 1.6047 | 0.26209 | -0.17435 | 2.6743
| 1.7077 | -0.4335 | -0.4153 | 3.2766
| -0.85163 | -2.0798 | 0.8692 | 5.8065
| 0.74078 | -1.3366 | 0.91762 | 3.1774
| -0.32583 | -0.11885 | -0.2775 | 0.19729
| 0.48047 | 0.48523 | -0.49657 | 0.71288
| 1.1633 | -0.18901 | 0.073304 | 1.3944
| -1.0787 | -0.72426 | 2.6771 | 8.8551
| 0.6537 | 0.47774 | 1.6121 | 3.2543
| -0.031022 | 0.89563 | -0.66188 | 1.2412
| 0.26087 | 0.077778 | 0.52279 | 0.34741
| 0.057581 | -1.9199 | 1.1586 | 5.0317
| -1.2038 | -0.31465 | 0.21556 | 1.5947
| -1.2579 | 1.4414 | 0.31129 | 3.7569
| 2.2255 | 0.34166 | 1.028 | 6.1265
| 0.87161 | 0.44923 | -0.10799 | 0.97317
| -0.0018791 | -0.21007 | 0.20527 | 0.086267
| 0.1518 | 0.67376 | -0.67048 | 0.92654
| 2.6569 | 0.20868 | 1.9679 | 10.9749
| 0.28425 | -1.4338 | 0.36413 | 2.2693
| -1.7867 | 1.3849 | 0.020476 | 5.1108
| -0.23353 | -0.10582 | -0.64848 | 0.48626
| 1.3706 | 0.37103 | 0.50409 | 2.2704
| -0.81884 | -0.57707 | 0.49457 | 1.2481
| -1.3038 | -0.37829 | 3.501 | 14.1001
| -0.9755 | 1.0289 | -1.2182 | 3.4942
| 0.38568 | 1.1159 | -0.34684 | 1.5143
| 0.2677 | 0.77593 | 1.3656 | 2.5385
| -0.94835 | -0.55514 | 0.23995 | 1.2651
| 1.5453 | -0.65186 | 0.95706 | 3.729
| -1.0176 | 0.030749 | -0.27334 | 1.1112
| 0.81647 | 0.98328 | -2.3049 | 6.9459
| -0.94693 | 0.98588 | -0.17369 | 1.8988
| -0.0034753 | -1.1297 | 0.31936 | 1.3782
| -1.36 | -0.88964 | 0.14022 | 2.6608
| -0.18616 | 0.27228 | 0.84958 | 0.83058
| -0.37086 | -0.90287 | -1.3206 | 2.6966
| -0.36616 | 0.7792 | -0.50922 | 1.0005
| 1.9461 | 0.17316 | -0.74488 | 4.3721
| -0.46931 | 0.034344 | -1.7606 | 3.3211
| 0.017835 | 0.9198 | 0.70162 | 1.3386
| -0.074652 | -1.2481 | -1.2476 | 3.1197
| 0.35258 | -0.80286 | 0.47373 | 0.99332
| 2.1853 | -1.472 | 0.096425 | 6.9515
| 0.032004 | 0.6684 | 0.78446 | 1.0632
| -1.7649 | -0.5635 | 0.4987 | 3.6812
| -0.60212 | 0.12162 | 0.3036 | 0.46951
| 0.023397 | -0.96737 | 1.5289 | 3.2739
| 1.8321 | -1.0684 | 0.19075 | 4.5347
| -0.19779 | -0.26405 | 0.4717 | 0.33134
| 1.8652 | 0.28476 | -0.027029 | 3.5606
| -0.12774 | -0.27446 | -1.4144 | 2.0921
| 0.072109 | 0.34809 | -2.9707 | 8.9512
| -0.4037 | -0.79984 | 0.14773 | 0.82454
| -1.1463 | 0.32507 | -3.9769 | 17.2355
| 1.2405 | 0.54596 | -0.061239 | 1.8407
| -0.4299 | 2.3429 | -1.9153 | 9.3424
| 2.0681 | -0.48678 | -2.7878 | 12.2855
| 0.38074 | 0.87827 | -0.85838 | 1.6532
| 0.66827 | 0.53776 | 0.013627 | 0.73595
| -0.58766 | -0.020317 | -0.6084 | 0.71591
| -0.53773 | 1.099 | -0.5252 | 1.7727
| 1.2205 | -0.70974 | -0.96159 | 2.9179
| -1.184 | 0.2844 | 0.53083 | 1.7646
| -1.2699 | -0.072332 | -0.54978 | 1.9202
| -0.2775 | 0.57812 | -0.49701 | 0.65824
| -1.5721 | -0.42102 | -2.2113 | 7.5387
| 0.016971 | -0.37042 | -1.5177 | 2.4409
| 0.36264 | -1.7431 | 0.32128 | 3.273
| -0.54677 | 0.75537 | 1.7717 | 4.0085
| 0.59223 | -1.6505 | -0.13157 | 3.0922
| 0.49646 | 1.41 | -0.67647 | 2.6921
| -0.10239 | -0.14093 | 1.7746 | 3.1795
| -0.052212 | -0.075528 | 0.13594 | 0.02691
| 1.1638 | -0.32991 | -0.71498 | 1.9744
| -2.3545 | 0.22457 | 1.7437 | 8.6346
| -1.5162 | 0.090756 | 1.0386 | 3.3859
| -0.30699 | 0.69141 | -0.96691 | 1.5072
| -0.14941 | 0.62754 | -0.65324 | 0.84284
| -1.9957 | -0.2273 | 0.50272 | 4.2874
| -0.99828 | 1.1764 | -0.58322 | 2.7206
| 0.27007 | 1.043 | 1.1823 | 2.5584
| 0.68604 | -2.2585 | 1.562 | 8.0112
| -0.095727 | -0.37509 | -0.71805 | 0.66545
| -1.1223 | -0.33011 | -0.71018 | 1.8728
| 1.6108 | -0.50556 | -0.62215 | 3.2374
| 1.3989 | -1.8359 | 0.69652 | 5.8124
| 0.54655 | -0.25542 | 1.2215 | 1.856
| 0.64738 | -0.66564 | 0.77937 | 1.4696
| -1.7031 | 0.16324 | -0.051851 | 2.93
| 3.3583 | -0.10354 | -1.4374 | 13.3553
| 0.61923 | -0.48409 | -1.4056 | 2.5936
| 0.36018 | -1.3423 | -0.40577 | 2.0961
| 1.0854 | -0.25084 | 0.4948 | 1.4858
| -0.61645 | -0.15938 | -0.80342 | 1.0509
| 1.0487 | 0.21411 | -0.28282 | 1.2255
| -0.83413 | 1.7473 | -1.3444 | 5.5562
| 0.37395 | 0.44422 | 0.27519 | 0.41289
| 0.94034 | 0.83887 | 0.12962 | 1.6048
| -0.14443 | -0.46296 | -0.13807 | 0.25425
| 0.09682 | 0.72463 | 0.36327 | 0.66643
| 0.17911 | 1.9732 | -0.18698 | 3.9607
| 1.9259 | -0.20194 | 0.77576 | 4.3516
| 0.042413 | -0.079772 | -0.45537 | 0.21552
| -0.39779 | -0.70022 | -0.94149 | 1.5349
| -0.74438 | 0.59566 | -0.31212 | 1.0063
| -0.17839 | 1.202 | 1.3461 | 3.2887
| 0.74225 | -0.30192 | 1.0588 | 1.7632
| -0.82786 | 0.6217 | -0.65529 | 1.5013
| -0.16931 | 0.045156 | 2.5802 | 6.6883
| 1.63 | -0.69107 | 0.12251 | 3.1496
| 0.40972 | -0.88122 | -0.15473 | 0.96836
| 0.34044 | 1.6284 | -0.76609 | 3.3544
| -0.32435 | 0.22621 | 0.5474 | 0.45602
| -0.17607 | 1.3792 | -1.4658 | 4.0818
| -0.54975 | -0.43598 | 0.72311 | 1.0152
| -0.29709 | 0.24938 | 0.70917 | 0.65337
| -1.3412 | -1.0866 | -0.18968 | 3.0156
| -0.32698 | -0.60832 | -0.88377 | 1.258
| -1.237 | -1.2346 | -0.020471 | 3.0549
| 1.605 | 0.59275 | -2.1297 | 7.4631
| -0.36056 | -0.62973 | 0.12284 | 0.54165
| -0.29541 | 0.68872 | -0.24132 | 0.61984
| 0.71115 | 0.65097 | 0.39754 | 1.0875
| 0.15446 | 0.64203 | -0.36293 | 0.56777
| -0.65327 | -1.9594 | -0.015212 | 4.2662
| -0.12122 | 0.95382 | 0.18118 | 0.9573
| 1.1049 | -0.46255 | 0.57849 | 1.7695
| -1.9431 | 2.1055 | 0.40305 | 8.3714
| 0.29298 | -0.49216 | 0.17085 | 0.35725
| -0.44154 | -1.1467 | 0.12862 | 1.5265
| 1.1221 | 1.397 | 1.6733 | 6.0108
| 0.083648 | 2.022 | -1.0712 | 5.243
| -0.66875 | -1.8737 | -0.65802 | 4.391
| 1.9372 | 0.15234 | 0.13829 | 3.7951
| -0.44064 | -0.14131 | 0.14476 | 0.23509
| 0.42106 | -1.7682 | -0.61456 | 3.6817
| -1.0528 | -0.26221 | -0.74361 | 1.73
| -0.65238 | -0.27165 | 1.9904 | 4.4612
| -0.13637 | -0.40541 | -1.7412 | 3.2146
| -0.053546 | 0.8426 | 1.5563 | 3.1349
| 0.74323 | -0.057741 | -0.62983 | 0.95241
| 1.7324 | 0.5445 | 0.36716 | 3.4326
| 0.99262 | 1.5794 | 0.2882 | 3.5628
| -0.74403 | 0.67323 | 0.18663 | 1.0417
| -0.48831 | 1.1664 | -0.29313 | 1.685
| -0.1228 | 1.7974 | -0.044822 | 3.2477
| -0.23605 | -0.28739 | 0.13791 | 0.15733
| -0.40964 | -0.13573 | 0.31095 | 0.28292
| 0.33011 | 0.35356 | -0.60638 | 0.60168
| -0.46019 | 0.46992 | 1.7779 | 3.5936
| 0.69108 | -0.61391 | 0.12734 | 0.87069
| -0.63888 | 0.69744 | 0.19228 | 0.93156
| -0.34301 | 1.3368 | 0.80562 | 2.5538
| 0.3187 | -0.83458 | 1.7888 | 3.998
| 0.25571 | 1.288 | 0.26361 | 1.7938
| 0.25467 | -1.3124 | -0.26167 | 1.8557
| -0.50058 | 0.20105 | -0.22493 | 0.3416
| -0.72825 | 1.6207 | -0.064428 | 3.161
| -0.0098216 | 1.1646 | 1.1859 | 2.7627
| 1.0756 | 0.2113 | -1.7785 | 4.3648
| 0.19631 | 0.9556 | -0.73321 | 1.4893
| -0.63913 | -0.0010888 | 0.068661 | 0.4132
| 0.34423 | -0.6193 | 0.37521 | 0.64281
| 0.64492 | 0.91386 | -0.80553 | 1.8999
| 1.1017 | 1.9054 | -0.17485 | 4.8748
| -0.0010677 | 0.79309 | -0.45202 | 0.83331
| 0.026252 | 0.52148 | 0.36809 | 0.40812
| -0.77068 | 1.5612 | 0.43783 | 3.2231
| -1.8951 | 0.88917 | -0.54971 | 4.6842
| 0.44287 | -0.20133 | 1.4248 | 2.2667
| 1.0784 | 0.18729 | -1.3816 | 3.1068
| -0.28678 | 0.92866 | 0.39024 | 1.0969
| 0.71887 | 0.14026 | 0.40006 | 0.6965
| -0.60471 | -1.5054 | 0.66115 | 3.069
| 0.93589 | -1.6112 | 0.24624 | 3.5326
| 0.41333 | 0.25944 | -1.5354 | 2.5956
| -0.43414 | -1.3739 | 0.76907 | 2.6676
| 1.3712 | -1.4708 | 1.3816 | 5.9523
| -0.10145 | 0.71767 | -0.51959 | 0.79532
| 0.942 | -0.84757 | 0.79326 | 2.235
| 0.25244 | -1.0759 | -1.1394 | 2.5196
| 1.408 | 0.89519 | -1.7502 | 5.8468
| 0.81308 | -0.68053 | 1.5024 | 3.3813
| -1.2145 | -0.39899 | 0.25218 | 1.6979
| 0.97729 | -1.4036 | 0.50153 | 3.1766
| 0.23268 | 0.014992 | -0.83465 | 0.75101
| -0.49001 | 1.0487 | -1.3334 | 3.1176
| -0.93861 | 0.31348 | 1.4243 | 3.0079
| 0.22477 | -0.73381 | 0.1334 | 0.60679
| 0.35539 | 0.39283 | -0.60225 | 0.64332
| -0.010849 | -1.5385 | 0.73573 | 2.9084
| -0.57491 | -0.88632 | 1.8891 | 4.6849
| 0.57496 | -0.40993 | -0.51665 | 0.76555
| 0.045192 | 0.39201 | 0.80498 | 0.8037
| -0.73404 | -0.86542 | -0.35518 | 1.4139
| 0.72193 | 1.726 | -0.84151 | 4.2085
| 0.99659 | 1.1508 | 0.83588 | 3.0162
| -0.45512 | 0.32945 | -0.79661 | 0.95026
| 1.5536 | 0.17274 | 0.22466 | 2.494
| 1.2412 | -0.0014587 | -1.2585 | 3.1243
| 1.1273 | 1.1336 | 2.0297 | 6.6755
| -1.2626 | 0.3364 | -1.0592 | 2.8292
| -0.37641 | 0.80284 | 1.5192 | 3.0941
| 0.16303 | 1.0948 | -0.053482 | 1.2281
| 1.4422 | 0.67034 | 1.1108 | 3.7632
| 0.26346 | 1.2261 | -0.251 | 1.6356
| 1.1514 | 0.1083 | 0.60904 | 1.7083
| 0.045046 | 0.77899 | -0.336 | 0.72175
| -0.27338 | 2.2716 | 1.869 | 8.7278
| 0.079733 | 0.26611 | -0.063273 | 0.081175
| 0.14886 | 0.061319 | 1.1304 | 1.3036
| 1.9373 | -0.18384 | 1.2282 | 5.2955
| -0.48994 | 0.021141 | 0.88593 | 1.0254
| -0.54492 | 0.58747 | 0.0922 | 0.65056
| -0.49784 | 0.53246 | -0.53764 | 0.82042
| 0.15985 | -0.10758 | -1.0103 | 1.0578
| -0.10455 | 0.75777 | 0.83944 | 1.2898
| 0.67581 | -0.26922 | 0.69663 | 1.0145
| -0.92637 | 0.51842 | 0.49783 | 1.3748
| -1.3121 | 0.25191 | 1.4547 | 3.9013
| -0.49067 | 1.1305 | -0.9475 | 2.4166
| 0.65747 | 1.3068 | 0.24372 | 2.1994
| 1.0402 | -0.60483 | 0.88163 | 2.2251
| 0.15328 | 0.049297 | -0.17259 | 0.055712
| 0.31315 | 1.0685 | 1.0768 | 2.3994
| 0.72889 | 0.43563 | 0.31144 | 0.81805
| 0.27478 | 0.37563 | 0.87169 | 0.97644
| 0.18778 | -1.095 | 0.13826 | 1.2533
| -1.1494 | 0.79998 | -0.47621 | 2.1879
| -0.19158 | -0.25686 | 0.65965 | 0.53783
| 1.0386 | 1.1261 | 1.4267 | 4.3823
| -0.37283 | 0.74973 | 0.53159 | 0.98368
| 1.6155 | 1.4209 | 0.52568 | 4.9053
| -0.56268 | -0.074754 | 0.58372 | 0.66293
| 0.54624 | 0.068517 | -0.79203 | 0.93039
| -0.24987 | -0.98021 | 0.30286 | 1.115
| -2.2038 | -0.71414 | -0.56834 | 5.6896
| 0.55178 | 1.2428 | 0.17963 | 1.8814
| 0.060783 | -0.74837 | 0.15412 | 0.5875
| 0.43647 | -0.64672 | 0.63724 | 1.0148
| 0.0070294 | 1.1484 | 0.32168 | 1.4224
| -0.69721 | -2.066 | 0.54902 | 5.0559
| 0.059624 | -0.56587 | -1.1183 | 1.5744
| -1.1923 | -0.33788 | -0.1004 | 1.5459
| 1.2 | 1.1816 | -0.55737 | 3.1468
| 0.42333 | 0.41147 | 0.099514 | 0.35842
| -1.004 | 0.31396 | -2.0144 | 5.1643
| 0.44987 | 0.56832 | 0.33841 | 0.63989
| 0.090393 | -0.47876 | -0.80987 | 0.89327
| -0.17006 | -1.0359 | 0.91857 | 1.9458
| 0.03133 | -2.4199 | 1.7684 | 8.9838
| -0.25099 | 1.8784 | 0.56628 | 3.912
| 0.61344 | -1.6976 | -0.92282 | 4.1096
| 0.4766 | -2.5133 | 2.1994 | 11.3812
| -0.53878 | 0.033664 | 1.571 | 2.7596
| -1.0487 | 0.38158 | 0.32715 | 1.3524
| -1.4441 | 0.45417 | 0.40179 | 2.453
| -0.1961 | 0.69603 | 2.1605 | 5.1905
| 0.48524 | 2.702 | 1.2451 | 9.0863
| 0.35693 | 0.72908 | 0.85673 | 1.3929
| -0.17516 | -0.31882 | -1.1775 | 1.5189
| 0.15462 | -0.063323 | 0.21745 | 0.075203
| 2.9464 | -0.36994 | -0.3475 | 8.9391
| 0.16679 | -0.22939 | -0.39151 | 0.23372
| 0.83726 | 0.87461 | -0.40355 | 1.6288
| 1.1734 | 0.29875 | 0.0058276 | 1.4663
| 0.010683 | 1.051 | 0.68995 | 1.5807
| 2.0738 | -1.6334 | -0.6104 | 7.3413
| -0.36287 | 0.5602 | -0.93591 | 1.3214
| 1.5984 | -1.0889 | 2.2649 | 8.8704
| -0.2285 | 1.0056 | -0.84774 | 1.782
| -0.34932 | 0.84297 | -0.26814 | 0.90451
| -1.5226 | 1.5305 | 1.5323 | 7.0087
| -0.34869 | 0.62941 | 1.9416 | 4.2874
| 0.78424 | 1.3083 | -0.36175 | 2.4576
| 0.52947 | -0.17293 | 1.816 | 3.6082
| -0.35268 | -1.7957 | 2.0519 | 7.5594
| 1.3626 | -2.3628 | 1.86 | 10.899
| -0.0075656 | 1.6127 | -1.0207 | 3.6427
| 0.15291 | 0.41048 | -1.3598 | 2.041
| -1.7378 | 0.85769 | -1.0395 | 4.8363
| -0.018305 | -0.16493 | 1.9074 | 3.6658
| 0.31345 | 0.67573 | 0.011536 | 0.55499
| 0.66963 | -0.87607 | -0.19856 | 1.2553
| -1.0325 | 0.23318 | 0.4332 | 1.308
| -0.0021052 | 1.0467 | 0.7212 | 1.6156
| -0.7354 | -0.37614 | -0.64212 | 1.0946
| -1.7358 | 0.66942 | -1.3494 | 5.2819
| 0.36859 | -1.191 | 1.2337 | 3.0765
| 1.1019 | 0.31231 | 1.3213 | 3.0575
| -0.30063 | 0.71197 | -1.0771 | 1.7575
| -0.64387 | -0.52873 | 0.85178 | 1.4196
| -0.17833 | -0.12334 | -0.29141 | 0.13193
| 0.57579 | -0.48994 | -1.0995 | 1.7805
| -0.51552 | 0.38655 | -1.8404 | 3.8022
| 0.4632 | 1.5312 | 0.38101 | 2.7043
| 1.6819 | 0.24579 | 0.5957 | 3.2442
| -0.85096 | -1.4319 | -0.21445 | 2.8205
| -1.0225 | -0.2245 | -0.74356 | 1.6489
| -0.12492 | 0.40282 | 1.0643 | 1.3105
| 1.0144 | -1.0362 | 0.65256 | 2.5286
| -0.66667 | 0.20619 | 0.66698 | 0.93183
| 0.36731 | -1.8306 | 0.21384 | 3.5316
| 1.2897 | 0.52169 | -1.2384 | 3.4693
| -0.47424 | -1.5959 | -0.97016 | 3.7131
| 1.7732 | -1.2847 | 0.15207 | 4.8177
| 0.77361 | -1.3214 | 1.3632 | 4.2029
| -2.1652 | -1.3995 | 0.99017 | 7.6272
| 2.4816 | 0.60881 | 0.66703 | 6.9737
| 0.77122 | 1.0325 | 0.069268 | 1.6656
| -0.28676 | -0.15559 | -0.59686 | 0.46268
| 2.149 | 0.46177 | -0.56212 | 5.1474
| 0.9722 | -0.60553 | -0.13345 | 1.3296
| 0.5563 | 0.22947 | -0.48073 | 0.59323
| -1.0794 | 0.63494 | 1.0305 | 2.6302
| 0.50721 | 0.61649 | 0.95701 | 1.5532
| -0.46877 | -0.51729 | 0.090382 | 0.4955
| -0.090149 | 0.23544 | -0.46048 | 0.2756
| 0.91367 | 0.23576 | 2.3999 | 6.6497
| -0.36529 | 0.95018 | -0.68147 | 1.5007
| 0.5216 | -0.68872 | 0.72347 | 1.2698
| 1.3807 | 0.89284 | -1.2703 | 4.3173
| -0.68247 | 0.74077 | 0.86155 | 1.7568
| 1.1712 | -0.44631 | 0.16364 | 1.5977
| 1.5534 | 1.0066 | -1.1435 | 4.7338
| 1.0382 | 1.1149 | 0.69649 | 2.8059
| 0.035833 | 0.13708 | -1.4206 | 2.0381
| -0.020747 | 0.32884 | -0.76547 | 0.69451
| -0.017564 | -0.71138 | 1.0303 | 1.5679
| 0.40786 | 0.18216 | 0.10502 | 0.21056
| -2.1736 | 0.48491 | 0.68088 | 5.4233
| 1.7824 | 2.2394 | 0.66188 | 8.6299
| 0.53437 | -2.1879 | -1.2723 | 6.6912
| 0.41202 | 0.60142 | -0.49576 | 0.77725
| 0.8362 | -1.6085 | 0.044504 | 3.2886
| 0.24428 | 0.25684 | -0.56011 | 0.43936
| 0.62286 | 3.0764 | -1.335 | 11.6341
| -0.65209 | -0.24365 | 1.7392 | 3.5095
| 0.5878 | -1.4172 | 0.70065 | 2.845
| 1.7057 | 0.35837 | 0.95275 | 3.9457
| 0.23528 | -1.5096 | 1.441 | 4.4109
| -0.82974 | -0.62206 | 1.1949 | 2.5032
| -0.83014 | 0.31168 | -0.41938 | 0.96216
| 0.23098 | -1.5535 | -1.2823 | 4.1111
| -0.57299 | -0.60823 | 0.40264 | 0.86039
| -0.40287 | 1.0123 | -1.9412 | 4.9552
| 0.60587 | 0.32864 | 1.5247 | 2.7999
| -1.3613 | -0.62742 | -0.19608 | 2.2853
| -0.57368 | 1.8152 | 0.17232 | 3.6536
| 0.10937 | 0.85926 | 0.018627 | 0.75064
| 0.85625 | 1.0375 | -0.64425 | 2.2246
| 0.21734 | 0.25182 | -0.28274 | 0.19059
| 1.1071 | -0.071014 | -1.9576 | 5.0628
| 0.61715 | 0.28156 | -1.0582 | 1.58
| 0.6344 | 1.1594 | 1.0777 | 2.908
| 0.052749 | 1.8374 | -0.54925 | 3.6806
| 1.9131 | 0.5407 | 1.7471 | 7.0047
| -0.45571 | 0.45221 | -0.72868 | 0.94314
| -0.71707 | 0.80553 | -0.067111 | 1.1676
| -0.82027 | -0.32068 | 1.1193 | 2.0285
| -0.43123 | 1.0131 | -0.65644 | 1.6433
| 0.56443 | -0.36052 | -0.57645 | 0.78086
| 0.42094 | -2.437 | -0.19873 | 6.1558
| -1.4742 | -1.2794 | -2.3138 | 9.1634
| 1.6299 | 1.3385 | 0.00054003 | 4.4479
| -0.23117 | 0.32749 | 0.042527 | 0.16249
| -0.080123 | -0.066321 | -0.15097 | 0.033609
| -1.0573 | 1.4406 | 0.22417 | 3.2434
| -0.44563 | 0.92706 | -1.3862 | 2.9796
| 0.033816 | 0.8983 | 0.53495 | 1.0943
| 1.6204 | -0.18047 | -0.32713 | 2.7653
| -0.6998 | 1.3198 | -0.023371 | 2.2322
| 0.70433 | -0.35613 | 0.59708 | 0.97942
| -0.85443 | 1.2799 | 0.55487 | 2.676
| -0.4146 | -0.10955 | 0.17851 | 0.21576
| -0.41544 | -0.82347 | -0.10057 | 0.8608
| -1.5082 | -0.47393 | -1.8762 | 6.0194
| -0.096729 | -1.6494 | 0.46714 | 2.948
| 0.75024 | -1.0709 | -0.72789 | 2.2394
| 0.92288 | -0.38494 | 0.33309 | 1.1108
| -0.059851 | -0.46256 | -0.72849 | 0.74824
| -0.12043 | 0.25186 | 0.81543 | 0.74286
| -0.31165 | -0.19397 | 0.48819 | 0.37308
| -0.25556 | 0.85432 | -1.1456 | 2.1076
| 1.1321 | -1.2574 | -1.2064 | 4.3183
| 0.032826 | -0.48098 | -1.1243 | 1.4965
| 0.038815 | -0.77216 | 0.12798 | 0.61412
| -0.12721 | -0.048959 | 0.85091 | 0.74263
| 2.3683 | 0.87514 | 0.39648 | 6.532
| -0.29853 | -0.75504 | 0.3859 | 0.80812
| -0.4788 | 0.24184 | -0.53565 | 0.57465
| 2.0717 | 0.57125 | -0.27742 | 4.6952
| -0.38919 | 0.0064555 | -0.99929 | 1.1501
| 1.242 | 0.39949 | -0.54191 | 1.9959
| -0.42139 | -1.9439 | -1.3788 | 5.8573
| -0.14649 | 2.4008 | 0.94996 | 6.6876
| 2.097 | -0.64519 | -0.94341 | 5.7038
| -0.5709 | 0.29371 | 0.047695 | 0.41447
| 0.55834 | 1.2272 | -0.12588 | 1.8337
| -1.6501 | 0.78849 | 0.29201 | 3.4299
| -0.72909 | 0.82884 | 0.54766 | 1.5185
| -0.74261 | -1.0485 | 0.11984 | 1.6652
| -1.9028 | 1.0865 | 0.80838 | 5.4548
| -1.2708 | -1.5703 | 0.4098 | 4.2487
| 0.40927 | 0.009092 | 0.032633 | 0.16865
| 2.2374 | -0.80632 | -0.6157 | 6.035
| -2.7474 | 0.5579 | -1.1824 | 9.2577
| 1.3596 | -0.22445 | -1.5946 | 4.4418
| 0.89593 | 1.5363 | 2.0704 | 7.4494
| 0.66945 | 0.13581 | -1.6441 | 3.1697
| -0.52484 | -2.2529 | 2.3915 | 11.0704
| 0.30999 | -0.20518 | -0.22494 | 0.18879
| 2.4994 | 1.0114 | -1.6059 | 9.849
| 1.2731 | 0.53153 | 0.80377 | 2.5493
| 1.5793 | 1.144 | 0.63898 | 4.2113
| 1.2163 | 0.047187 | -0.45833 | 1.6917
| -0.0048706 | 2.6887 | -0.4615 | 7.4419
| -0.77419 | 0.02555 | 0.56583 | 0.92019
| 1.7273 | -3.4972 | 0.58969 | 15.5617
| 0.40724 | -1.2066 | -1.1243 | 2.8859
| 1.1766 | -0.4128 | 0.91206 | 2.3866
| 1.9358 | 1.0005 | 2.0979 | 9.1495
| 0.64833 | 1.2253 | 1.6626 | 4.6861
| 0.88287 | 1.7408 | -0.8861 | 4.5951
| -0.41405 | -1.8718 | 1.0184 | 4.7122
| 0.083181 | -0.94929 | 0.93754 | 1.7871
| -1.2612 | -0.53869 | -0.49749 | 2.1282
| -0.82222 | -0.20951 | 0.39057 | 0.87249
| 0.41826 | 0.57532 | -1.845 | 3.9101
| 0.78639 | 0.085066 | 0.54286 | 0.92035
| -0.46003 | -0.74375 | -2.4344 | 6.6913
| -0.41513 | -0.69835 | 0.61464 | 1.0378
| -1.2433 | -0.74726 | 0.96199 | 3.0295
| -0.25421 | 1.4893 | 0.52481 | 2.5581
| 1.1772 | -0.47245 | 0.27735 | 1.6861
| -2.3579 | 1.5019 | -0.067567 | 7.8199
| -0.78617 | 0.43537 | 0.92783 | 1.6685
| -1.8723 | -0.28255 | -0.11429 | 3.5984
| 1.0914 | -0.69082 | -1.0013 | 2.6711
| 1.1783 | 0.26319 | 0.64018 | 1.8676
| -0.53936 | -0.099603 | 0.026471 | 0.30153
| 0.43908 | 1.6906 | 0.99285 | 4.0368
| 0.35638 | 2.1473 | 0.58128 | 5.0759
| 0.42675 | 1.7137 | -0.69831 | 3.6067
| -0.86775 | 1.385 | -2.2344 | 7.6637
| -1.0013 | 1.2192 | -0.64907 | 2.9103
| 0.15666 | -0.47702 | 1.3783 | 2.1517
| -0.8067 | 0.96986 | -0.40612 | 1.7563
| -1.4737 | 1.6734 | -2.2738 | 10.1426
| 2.0289 | -0.17686 | 0.65228 | 4.5732
| 0.1179 | 0.057843 | -0.20473 | 0.05916
| 1.2498 | -0.70036 | 1.7826 | 5.2302
| -0.387 | -1.1158 | -0.12722 | 1.4109
| -0.71509 | -0.25742 | -0.13794 | 0.59664
| -0.024569 | -0.7416 | 0.081974 | 0.5573
| -0.2235 | -0.70575 | 0.32528 | 0.65385
| -0.0043645 | -0.13679 | -0.86213 | 0.762
| 1.5715 | 0.63709 | 0.55752 | 3.1862
| -1.5031 | -0.27889 | -0.45801 | 2.5468
| -0.80286 | 2.2422 | -1.6372 | 8.3523
| 0.13508 | -2.2336 | 1.0487 | 6.1068
| -0.46143 | -0.65088 | 0.9505 | 1.54
| -2.1392 | 1.7723 | 0.34353 | 7.8353
| -0.75304 | 0.71206 | 0.84321 | 1.7851
| 1.1566 | -0.781 | 2.8367 | 9.9948
| 0.51887 | 0.88924 | -0.75866 | 1.6355
| -1.1694 | 0.15793 | -1.829 | 4.7379
| -1.6815 | 0.19845 | 1.1486 | 4.1861
| 0.53526 | -0.14645 | 0.87068 | 1.066
| -1.0492 | 0.66776 | -0.20374 | 1.5882
| 1.793 | -0.93966 | 2.3932 | 9.8251
| -0.073666 | 0.24084 | 0.72233 | 0.58519
| -0.37899 | -1.232 | -1.6709 | 4.4535
| -0.04096 | -0.58018 | 1.8544 | 3.7772
| -1.3592 | -0.0049498 | 1.0269 | 2.9021
| 1.5397 | -0.50198 | 1.933 | 6.3591
| 0.98316 | -2.2585 | 0.10085 | 6.0777
| -0.86467 | -1.4996 | 1.6627 | 5.7612
| -1.7423 | 0.57425 | 0.048185 | 3.3678
| 0.86419 | -1.9275 | 0.06476 | 4.4662
| -0.22868 | 0.23131 | 0.55974 | 0.41911
| -0.88355 | 1.0077 | 0.36022 | 1.9259
| 0.019797 | 0.89011 | 1.2091 | 2.2547
| -0.2045 | -0.16233 | 0.61605 | 0.44769
| 0.044166 | -0.28235 | -0.21239 | 0.12679
| -0.2126 | 1.9324 | 0.010421 | 3.7796
| 0.38158 | 0.51463 | 1.3128 | 2.1339
| -1.8757 | -0.12755 | 1.0647 | 4.6682
| 0.12504 | 0.37199 | -0.4859 | 0.39011
| 0.80741 | -0.5296 | -1.5715 | 3.402
| 1.0969 | 1.7037 | 0.2309 | 4.159
| 0.05667 | 0.46866 | 0.40845 | 0.38968
| 0.88942 | -1.898 | 0.32983 | 4.5024
| -0.21872 | 1.7344 | 0.36746 | 3.1911
| 1.3228 | 0.28873 | -0.17515 | 1.8639
| 1.1961 | -1.6022 | -1.007 | 5.0118
| 2.3718 | 1.1837 | -1.7029 | 9.9266
| 0.24377 | -0.51152 | -1.023 | 1.3677
| 1.3445 | 0.10767 | 0.64405 | 2.2341
| -0.69181 | -0.58174 | -1.5843 | 3.327
| 0.32937 | 1.9381 | 0.06424 | 3.8688
| 1.2866 | 0.0042926 | 0.48026 | 1.8861
| -0.2862 | 0.7568 | -1.1754 | 2.0363
| -1.0839 | -1.3999 | -0.024232 | 3.1351
| -0.019531 | -0.55142 | 0.90002 | 1.1145
| -1.0455 | -0.82402 | -0.89206 | 2.5677
| 0.01432 | -2.514 | 0.025445 | 6.321
| -0.5053 | 0.29472 | -0.29805 | 0.43102
| -0.94696 | 0.11142 | 0.25384 | 0.97358
| -0.27041 | 0.27957 | 0.92028 | 0.9982
| -0.21022 | -0.64588 | 0.71151 | 0.9676
| -0.36599 | 0.95286 | 1.7912 | 4.2504
| 0.36132 | 0.53624 | 0.24967 | 0.48044
| -0.21636 | -0.19145 | -0.95783 | 1.0009
| -0.6062 | 0.16816 | 1.2893 | 2.0581
| -0.77879 | -0.40541 | 0.59784 | 1.1283
| -0.92658 | -0.030959 | -0.4089 | 1.0267
| 1.5962 | 0.37809 | -0.23875 | 2.7477
| 0.59727 | -1.1316 | 0.92747 | 2.4974
| -0.74255 | -0.42609 | 1.7518 | 3.8016
| -0.50841 | -0.30437 | 0.0013107 | 0.35112
| 3.6958 | -0.23276 | -1.0192 | 14.7518
| -0.18991 | -0.96137 | -0.85486 | 1.6911
| -0.29115 | 0.5392 | 0.42458 | 0.55578
| 0.6301 | -2.301 | -0.33701 | 5.8053
| 1.5952 | -0.57589 | -0.89062 | 3.6695
| 0.22475 | 2.4366 | -1.4905 | 8.2092
| -0.13746 | -0.5658 | -0.60367 | 0.70344
| 0.92261 | 0.053572 | 0.57223 | 1.1815
| 1.5021 | -0.38189 | -0.32399 | 2.5071
| -1.0448 | 1.7698 | -0.056467 | 4.227
| 0.69705 | -0.57276 | -1.1003 | 2.0245
| 0.87483 | -0.76242 | -0.072623 | 1.3519
| 0.043202 | 0.20102 | 1.3178 | 1.7788
| -0.82701 | -2.8715 | -0.38094 | 9.0748
| -0.88323 | 0.84336 | 0.021644 | 1.4918
| 1.0881 | -0.81255 | 0.0033111 | 1.8442
| 0.52713 | -0.1072 | 0.040742 | 0.29101
| 1.158 | 0.36953 | 0.74202 | 2.028
| -0.70726 | -1.0378 | -2.0767 | 5.8899
| 1.8017 | -0.62764 | -0.023644 | 3.6405
| 0.28469 | -0.76658 | -0.87835 | 1.4402
| 2.3818 | -0.89475 | -0.2255 | 6.5246
| -0.97054 | 0.34017 | -0.0023714 | 1.0577
| 0.14881 | 0.065474 | 0.62487 | 0.41689
| -0.68067 | 1.6789 | -0.54495 | 3.5789
| -2.7411 | -0.47413 | -0.46343 | 7.9533
| 1.2595 | 1.4416 | 1.2629 | 5.2596
| 1.3471 | -2.5946 | 0.54637 | 8.845
| 0.38042 | -0.084407 | -0.28016 | 0.23033
| 0.675 | -0.31093 | -0.090971 | 0.56058
| -0.85588 | 0.5257 | -0.17053 | 1.038
| 0.42391 | 0.85257 | -0.78405 | 1.5213
| -1.8313 | 0.32141 | 0.75885 | 4.033
| 1.6396 | 0.054566 | 0.52453 | 2.9665
| 1.6784 | -0.8529 | 0.41404 | 3.7158
| -0.26552 | -0.35527 | 1.9831 | 4.1293
| 0.47296 | -0.61249 | 0.133 | 0.61652
| -0.4606 | 0.078219 | 1.8383 | 3.5976
| -0.94656 | -0.28428 | -0.26296 | 1.0459
| 1.1855 | -1.1775 | -0.73841 | 3.3372
| 0.051641 | -0.077411 | -0.095635 | 0.017805
| -0.78961 | -0.4912 | -0.62246 | 1.2522
| -0.052075 | 0.68947 | -0.69993 | 0.96799
| 0.1408 | 0.84398 | -1.9101 | 4.3807
| 0.44507 | 0.97262 | -0.67002 | 1.593
| 0.68453 | -0.15235 | -0.78311 | 1.1051
| -0.056472 | -0.70563 | 1.2223 | 1.9952
| -0.79193 | 0.025216 | 1.0398 | 1.7091
| -0.63915 | -2.6505 | -1.5582 | 9.8615
| 0.25708 | -0.82152 | 1.9772 | 4.6505
| -0.25707 | -0.29135 | -0.75634 | 0.72302
| -0.31646 | -1.0494 | -2.6832 | 8.4011
| -1.6378 | -0.60364 | 0.99681 | 4.0405
| -1.159 | -0.23591 | 0.32514 | 1.5047
| -0.31086 | -0.50252 | 0.076548 | 0.35502
| 0.80397 | -0.23059 | 0.38978 | 0.85147
| 1.4011 | -0.83362 | 0.1694 | 2.6867
| -0.80929 | 0.0356 | -0.74038 | 1.2044
| 0.42071 | 0.38195 | 1.3644 | 2.1846
| 1.6404 | 0.14176 | -0.33672 | 2.8244
| -0.73116 | -0.10855 | -0.072683 | 0.55166
| 0.34369 | 0.80567 | 1.0509 | 1.8717
| 0.19021 | 1.7086 | 1.4649 | 5.1014
| 1.0115 | 0.86085 | -1.1782 | 3.1524
| 0.85642 | -0.82679 | 1.6234 | 4.0523
| -0.6207 | -1.1578 | 1.7072 | 4.6403
| -0.74612 | -0.83306 | 0.74664 | 1.8082
| 0.51564 | 1.4856 | -0.12488 | 2.4885
| -0.24146 | 3.0437 | -0.11131 | 9.3346
| -1.025 | 0.97072 | -1.3585 | 3.8383
| -0.34369 | 0.73795 | -2.4526 | 6.6777
| 0.35913 | 0.28759 | 1.1556 | 1.5471
| 1.1219 | 0.92152 | -1.2225 | 3.6026
| 0.077784 | -1.1552 | 1.2181 | 2.8245
| -0.053512 | -0.7194 | 1.4006 | 2.4822
| 0.79623 | -1.62 | -0.0018402 | 3.2583
| 1.0277 | 0.37737 | -0.16595 | 1.2261
| 1.8359 | 0.31709 | -0.41084 | 3.6397
| -0.866 | 0.69152 | 0.18515 | 1.2624
| -1.461 | -0.23346 | -0.02005 | 2.1895
| -0.77012 | 0.1096 | 0.17422 | 0.63545
| -0.24483 | -0.39743 | 0.3442 | 0.33637
| 0.17012 | -1.1071 | 0.55967 | 1.5679
| -0.11786 | 1.4006 | 1.6611 | 4.7348
| 0.55559 | -0.061545 | -0.39312 | 0.46702
| -0.49022 | -1.1252 | -0.64093 | 1.9171
| -0.26986 | -0.29006 | 1.2154 | 1.6343
| 0.48182 | -0.028628 | -0.59902 | 0.5918
| -1.1613 | 1.8515 | -0.82765 | 5.4615
| 0.66352 | 0.45462 | 0.52591 | 0.92351
| -0.51451 | 0.27108 | 0.43922 | 0.53112
| -0.28778 | 1.5552 | 1.7504 | 5.5653
| 0.35457 | 0.29486 | 0.52103 | 0.48414
| -0.85107 | -0.012654 | 0.34287 | 0.84204
| 0.61482 | -0.072778 | 0.097581 | 0.39283
| -0.60777 | -0.80526 | -1.1917 | 2.4379
| 0.64817 | -1.1063 | 0.58032 | 1.9807
| 0.50952 | 0.96971 | 0.87248 | 1.9612
| -1.2948 | -0.69698 | -0.019703 | 2.1627
| -1.5279 | 2.4735 | -0.094038 | 8.4614
| 0.46191 | -1.6696 | -0.29335 | 3.0871
| 1.0518 | 0.096571 | -1.2024 | 2.5614
| 0.25321 | -2.0133 | 1.5531 | 6.5297
| 1.0108 | 1.0899 | -0.31884 | 2.3114
| 1.0829 | -0.41405 | -0.39281 | 1.4983
| 0.34422 | 0.90675 | -0.60186 | 1.3029
| 1.9802 | 0.6454 | 0.54355 | 4.6333
| 0.76364 | 1.5725 | 0.1066 | 3.0672
| 0.25351 | -0.68654 | 0.56491 | 0.85473
| 0.45787 | 0.89934 | 0.31275 | 1.1163
| 1.2527 | 0.87941 | 0.42065 | 2.5196
| -0.83055 | 1.2053 | 1.1049 | 3.3633
| 0.085275 | -0.24764 | -0.10322 | 0.079253
| 1.0576 | 0.45608 | -0.3816 | 1.472
| 0.47285 | 0.81182 | -0.94943 | 1.7841
| 0.11707 | 0.78638 | -0.53023 | 0.91324
| -1.3379 | -0.37944 | 1.3843 | 3.8503
| 0.2721 | -0.82626 | 1.4712 | 2.9213
| 0.67313 | 1.0106 | 1.9091 | 5.1192
| 0.1627 | 0.5737 | 0.71042 | 0.86029
| -1.1767 | -0.51791 | 1.098 | 2.8584
| 0.6315 | -1.5057 | 0.28841 | 2.7491
| -2.1018 | 1.1044 | 1.318 | 7.3745
| -0.33971 | 0.55048 | -0.58369 | 0.75912
| -0.13593 | 1.5359 | 1.0547 | 3.4899
| -0.86877 | -0.28462 | -0.93238 | 1.7051
| -0.55938 | -0.34463 | -1.2009 | 1.8738
| 0.69074 | -0.32374 | -1.6439 | 3.2843
| 0.95159 | 1.0148 | 0.29767 | 2.0239
| -0.38447 | 2.3289 | -1.5482 | 7.9686
| 1.6452 | 2.1348 | -0.10073 | 7.274
| -1.0123 | -0.033264 | -0.2917 | 1.1109
| -1.0482 | -0.20539 | 3.1491 | 11.0575
| 0.29872 | -0.33872 | 2.0297 | 4.3237
| 0.56394 | 1.0992 | 1.4338 | 3.5821
| -0.1659 | -0.89656 | 0.33347 | 0.94255
| 0.30557 | 0.021802 | 0.55118 | 0.39765
| 0.16888 | -0.53587 | 0.81667 | 0.98262
| -0.2533 | -0.64981 | 1.3895 | 2.4171
| 0.068202 | -1.9199 | -1.0353 | 4.7625
| 0.057823 | 1.1904 | 0.13817 | 1.4394
| -0.57776 | -1.0914 | 0.60357 | 1.8892
| -0.021699 | -2.0951 | -0.13961 | 4.4096
| 1.5021 | 1.4348 | 0.99997 | 5.3149
| -0.19719 | -1.1723 | 1.0539 | 2.5239
| -1.4869 | -0.6827 | -0.62152 | 3.0633
| 0.16053 | 1.1355 | -0.65556 | 1.745
| 1.4812 | 0.065969 | -0.38689 | 2.3479
| -0.054075 | 0.91619 | 1.2885 | 2.5026
| -0.94837 | -1.5629 | 0.36632 | 3.4761
| 1.4808 | -0.51901 | 0.6367 | 2.8676
| -0.25332 | 1.079 | 0.86232 | 1.972
| -1.6878 | 0.081432 | 0.65106 | 3.2791
| -2.7639 | -0.1011 | 1.1488 | 8.9693
| 0.83862 | -0.20268 | 2.7908 | 8.5327
| 0.25389 | -0.6785 | -0.14152 | 0.54486
| 0.85674 | 1.0066 | 0.040894 | 1.7489
| -0.33211 | -0.11674 | -0.61243 | 0.499
| 0.66349 | -1.9778 | 0.36584 | 4.4857
| -0.30753 | -0.36838 | 0.40465 | 0.39402
| -0.035446 | -0.7798 | -1.129 | 1.8839
| 1.0602 | -0.12193 | 0.37324 | 1.2783
| 1.4742 | 1.3373 | 0.089741 | 3.9697
| -1.0695 | 1.7295 | 0.028751 | 4.1357
| -0.93433 | 0.091917 | -0.4412 | 1.0761
| 0.27831 | 1.7653 | -0.028208 | 3.1945
| 1.0936 | 0.2344 | 1.8477 | 4.6647
| 1.3407 | 0.83537 | -0.14056 | 2.515
| 1.1818 | 1.8075 | 0.07991 | 4.6699
| 0.76434 | 0.13728 | -0.50529 | 0.85838
| -0.23746 | 0.60795 | 0.07885 | 0.43221
| -1.6197 | 1.7212 | -0.26328 | 5.6552
| -1.2471 | 1.6777 | 0.69541 | 4.8534
| -1.0965 | 1.218 | -0.653 | 3.1121
| -0.74691 | 0.18161 | 0.93172 | 1.459
| -0.086617 | -0.10574 | -0.77846 | 0.62468
| 1.6716 | 1.4661 | 0.49486 | 5.1884
| 0.16216 | -0.55699 | -0.84289 | 1.047
| -0.5759 | -0.50382 | 1.4459 | 2.6761
| 0.95294 | -0.14852 | -0.45152 | 1.134
| 1.1581 | 1.2933 | 1.5042 | 5.2765
| -0.66487 | -1.1802 | 1.8956 | 5.4284
| 0.50852 | 0.76455 | -0.057552 | 0.84644
| 0.51947 | 0.48354 | -1.6763 | 3.3135
| -0.83323 | 0.30444 | -0.028812 | 0.78779
| 0.89079 | -0.39199 | 0.9037 | 1.7638
| 0.78404 | 0.062288 | 0.24721 | 0.67971
| -0.6658 | 0.047902 | -0.49923 | 0.69482
| 0.68484 | -0.33455 | -1.3282 | 2.3451
| -1.1449 | -0.46079 | -0.91687 | 2.3637
| -1.231 | 0.17143 | -1.4718 | 3.7109
| -0.0041538 | 0.12905 | 0.98638 | 0.98962
| -0.04801 | -2.0456 | 0.3553 | 4.3131
| 1.1521 | 1.8172 | 1.2074 | 6.0876
| 0.5148 | 1.0673 | -0.040316 | 1.4059
| 1.1324 | -0.42571 | -0.96076 | 2.3866
| -1.4624 | -1.1817 | 1.8016 | 6.7806
| 0.096745 | 1.2424 | 0.58684 | 1.8973
| 0.54119 | 0.39526 | -0.00080991 | 0.44911
| -0.51303 | 0.53428 | -0.92755 | 1.409
| 0.35237 | 0.14956 | -0.81056 | 0.80355
| 0.025826 | 0.58611 | 0.14102 | 0.36408
| -0.45412 | 1.1643 | -0.40869 | 1.7288
| -1.0402 | 1.699 | 1.0593 | 5.0908
| 0.87442 | 0.79612 | 0.65657 | 1.8295
| 0.42876 | 0.34987 | -0.82616 | 0.98878
| -0.4113 | -0.76239 | 0.40838 | 0.91718
| 0.5064 | 0.59552 | 0.30037 | 0.7013
| 0.26629 | -0.28567 | 0.27069 | 0.22579
| -1.0866 | 0.20648 | 1.142 | 2.5276
| 0.25025 | 0.62641 | 0.18208 | 0.48816
| 0.063979 | 1.7085 | 0.93062 | 3.7891
| 0.84999 | 0.66285 | 1.1758 | 2.5444
| -1.2753 | 0.97394 | -2.0994 | 6.9825
| -0.3656 | 0.48105 | -2.4426 | 6.3315
| -0.048019 | -2.2039 | 0.72177 | 5.3804
| -0.26895 | 1.1792 | -1.8607 | 4.9251
| 0.035352 | 0.18461 | -0.075787 | 0.041074
| 1.9964 | 0.16449 | 0.414 | 4.1841
| 1.0926 | -0.82708 | -0.2697 | 1.9506
| -0.035524 | 0.85128 | -0.0038827 | 0.72595
| -2.0598 | -2.189 | 0.78215 | 9.6464
| -3.1152 | -0.70928 | 0.19888 | 10.2469
| -0.63822 | -0.2089 | 1.48 | 2.6414
| -1.373 | -0.93347 | -0.021175 | 2.757
| -0.81793 | 0.44288 | -0.18045 | 0.89771
| 0.98655 | -1.1659 | -1.0442 | 3.4231
| -0.79761 | 1.1236 | -0.28212 | 1.9783
| -0.10827 | -0.10573 | 1.0012 | 1.0253
| 1.9198 | -0.60454 | -0.015077 | 4.0514
| 0.24172 | 2.4094 | 0.011834 | 5.864
| 1.3436 | -1.0174 | 0.18238 | 2.8734
| 0.26173 | 0.51774 | -1.2547 | 1.9108
| -2.1065 | -0.97659 | -0.40983 | 5.559
| -0.72749 | -0.75623 | 0.28333 | 1.1814
| -0.47523 | 0.27572 | 2.0616 | 4.5521
| -0.39871 | 0.78591 | 1.6529 | 3.5088
| 1.0023 | -1.0726 | -1.6519 | 4.8841
| 1.0453 | -0.23575 | 1.2623 | 2.7415
| 1.5295 | 0.057996 | 0.77029 | 2.936
| 0.54154 | 1.4823 | -0.99513 | 3.4808
| -0.38122 | 0.6861 | -1.1508 | 1.9404
| 1.5734 | -0.29875 | 1.0286 | 3.6228
| 0.41169 | 0.28418 | 1.0065 | 1.2633
| -1.2953 | -0.8415 | 0.47618 | 2.6127
| -0.75681 | 0.78084 | 0.98372 | 2.1502
| 0.36663 | -0.47683 | -0.8564 | 1.0952
| -0.47107 | -0.3281 | -0.97749 | 1.285
| -1.3863 | -0.96815 | -0.58799 | 3.2049
| 0.71533 | -0.55258 | -0.22837 | 0.86919
| 0.72122 | -1.0896 | 1.566 | 4.1597
| -0.4778 | 0.17239 | -1.5642 | 2.7048
| -1.9803 | -0.43994 | 0.19138 | 4.1519
| -0.44433 | 0.085379 | 0.44143 | 0.39959
| 0.64761 | -0.32678 | -0.77655 | 1.1292
| -0.70888 | -0.11584 | 0.16517 | 0.54321
| -0.19318 | 0.0056566 | -0.1496 | 0.059731
| -1.2928 | -1.0731 | 0.31856 | 2.9246
| 0.36357 | 0.75158 | -1.7152 | 3.6389
| 1.623 | -0.43428 | -2.1284 | 7.3528
| -0.64735 | 0.45807 | -0.66601 | 1.0724
| 1.0864 | -0.45787 | -0.032087 | 1.3909
| 2.1753 | 1.1867 | 1.2276 | 7.6473
| -1.1443 | 0.84233 | 0.9557 | 2.9324
| -1.575 | -0.45839 | -0.79513 | 3.3231
| -0.90703 | 1.8175 | -0.8544 | 4.856
| 1.7815 | 0.70037 | -2.4003 | 9.426
| 0.68449 | -0.65573 | -0.32188 | 1.0021
| 0.20745 | -0.99033 | 0.87406 | 1.7878
| 0.38561 | -1.9312 | 0.49334 | 4.1218
| 0.53625 | -0.65162 | 0.64449 | 1.1275
| -0.3592 | 0.23705 | -1.2067 | 1.6412
| 0.30063 | 0.55074 | -1.1282 | 1.6665
| -1.2998 | -0.12655 | -0.39689 | 1.8631
| -0.68767 | 0.47019 | -0.79448 | 1.3252
| 0.94068 | -0.042274 | -1.2674 | 2.4928
| 0.91765 | -0.2767 | -1.2149 | 2.3945
| -0.0066282 | -0.14957 | 0.40718 | 0.18821
| 1.4268 | -0.069337 | 1.5576 | 4.4669
| -1.4068 | 2.2208 | -2.6086 | 13.716
| -1.4555 | -0.085787 | -0.76398 | 2.7094
| -0.74978 | 0.48698 | -0.19622 | 0.83783
| 1.153 | 1.3713 | -0.036151 | 3.211
| -0.7043 | -0.61176 | 0.32325 | 0.97478
| -1.144 | -0.45587 | 0.61066 | 1.8895
| -1.3659 | -1.5178 | 0.71551 | 4.6815
| -0.83452 | 0.80088 | 0.5075 | 1.5954
| 0.31791 | -0.38583 | -0.28518 | 0.33126
| 1.6039 | 0.064206 | 0.15852 | 2.6016
| 0.88567 | -0.0038216 | -0.6618 | 1.2224
| 0.67323 | -0.63599 | 0.33635 | 0.97085
| -0.30395 | -2.176 | -1.5252 | 7.1536
| 0.55902 | -1.0325 | 1.1014 | 2.5918
| -0.27683 | 1.2841 | -2.117 | 6.2075
| 0.4213 | -0.31538 | 1.1446 | 1.5871
| 0.71573 | -0.9018 | 0.20406 | 1.3672
| -2.0693 | 0.40618 | 0.020408 | 4.4474
| 0.25063 | 0.73525 | -0.56641 | 0.92423
| -0.98953 | -2.4519 | 0.54271 | 7.2855
| 0.49535 | 1.1201 | 0.7053 | 1.9975
| 0.32519 | 0.99627 | -0.03952 | 1.0999
| 0.1377 | -1.1153 | 0.5711 | 1.589
| -1.738 | 1.0358 | -0.78632 | 4.712
| -0.39962 | -0.56991 | -0.32712 | 0.59149
| -0.85088 | -0.19957 | 0.61232 | 1.1388
| 1.4405 | -0.59597 | -0.3099 | 2.5262
| 0.38387 | -1.5246 | 0.56114 | 2.7868
| 0.8585 | 0.82883 | -0.69128 | 1.9019
| -1.367 | -0.81531 | -0.97678 | 3.4876
| -0.28282 | -0.54486 | 0.69548 | 0.86055
| 1.1118 | 1.8943 | 0.38929 | 4.976
| 0.60682 | -0.88648 | -2.8305 | 9.1657
| -0.0022119 | 0.0072308 | 0.96374 | 0.92886
| 0.13087 | -0.39522 | -0.26109 | 0.24149
| 0.26206 | 0.78867 | -1.4739 | 2.8629
| -0.44738 | -0.55323 | -0.77366 | 1.1048
| 0.43566 | 0.2499 | -1.42 | 2.2687
| 0.69419 | -0.6315 | 0.25024 | 0.94331
| -0.0024614 | -1.6322 | 0.94064 | 3.549
| 0.64648 | -0.024039 | -0.29085 | 0.50311
| 0.094623 | 0.4187 | 0.21234 | 0.22935
| 0.43821 | -1.5455 | 0.207 | 2.6236
| 0.55793 | -0.62452 | 0.88836 | 1.4905
| 2.0682 | 3.3396 | -0.3116 | 15.5272
| 0.26779 | -0.37753 | -0.63763 | 0.62081
| -0.8202 | 1.1413 | 1.5839 | 4.4842
| -0.21161 | -1.0715 | -0.31921 | 1.2947
| -1.106 | -0.062745 | 0.14751 | 1.2489
| 0.048138 | 3.0967 | -0.89986 | 10.4015
| 0.20428 | -0.38672 | 0.071258 | 0.19636
| -0.5405 | -1.5459 | 0.58276 | 3.0214
| 2.0327 | -1.9427 | 0.16868 | 7.9346
| 0.022811 | -0.38293 | -0.67191 | 0.59862
| -0.61298 | -0.93363 | 1.366 | 3.1133
| -0.080945 | 0.45128 | 0.27258 | 0.2845
| 1.9359 | 0.26099 | 0.0099115 | 3.8159
| 0.55735 | 0.87199 | 0.80779 | 1.7235
| -0.95068 | -0.022265 | 0.62784 | 1.2985
| 0.53409 | 0.3179 | 1.1559 | 1.7224
| 1.0137 | -1.6504 | -0.090763 | 3.7595
| 1.099 | 2.0293 | -0.097328 | 5.3354
| -1.6588 | 1.006 | 1.0622 | 4.8917
| 0.23941 | 1.909 | -3.0539 | 13.028
| -0.084833 | -1.2099 | 0.2366 | 1.527
| -0.35524 | 0.56725 | 1.0834 | 1.6217
| 1.4682 | 0.98793 | -0.20067 | 3.172
| 0.29112 | 1.2267 | 0.49707 | 1.8367
| 0.83272 | 1.4821 | 0.76998 | 3.4829
| 1.7615 | -1.8927 | 0.79768 | 7.3214
| 1.1739 | 0.28858 | -1.8138 | 4.7513
| -0.91308 | -0.80012 | -0.64064 | 1.8843
| 1.8135 | -0.025205 | 1.4444 | 5.3757
| 0.0023977 | 0.083279 | -0.3744 | 0.14712
| 0.4328 | -0.45152 | -0.035666 | 0.39245
| -0.45308 | 0.16444 | -0.92937 | 1.0961
| -0.49258 | 1.1671 | 1.0932 | 2.7999
| 0.70239 | -0.30107 | -1.8323 | 3.9413
| -0.11664 | 0.11507 | 0.35733 | 0.15453
| 0.32461 | -1.2741 | -0.010333 | 1.7287
| 0.57969 | -0.17532 | -0.96913 | 1.306
| -1.1581 | -0.11901 | 0.75886 | 1.9313
| 0.21461 | -0.78435 | -1.2323 | 2.1799
| -0.70554 | -0.48165 | 1.0087 | 1.7472
| 1.479 | 0.26961 | 0.63383 | 2.6617
| 1.5768 | 1.1988 | -0.20987 | 3.9675
| 0.14552 | 0.41585 | 0.19312 | 0.2314
| -0.79919 | -0.25155 | 0.38238 | 0.8482
| 1.6732 | 1.6006 | 0.3044 | 5.4541
| 0.79405 | 0.76998 | -0.73302 | 1.7607
| 0.52753 | 2.1191 | -1.0805 | 5.9361
| 1.5095 | -1.1581 | 0.93678 | 4.4975
| -1.828 | 0.71788 | 1.3182 | 5.5944
| -0.58278 | 1.0595 | 1.012 | 2.4862
| -1.8753 | -1.7818 | 0.40442 | 6.8551
| 0.71604 | 2.072 | -1.484 | 7.0082
| -1.4539 | -0.16148 | 0.64548 | 2.5565
| 0.866 | -0.37527 | -1.0035 | 1.8978
| 3.1835 | 0.43489 | -0.083233 | 10.3306
| 0.34699 | -1.1807 | 1.4675 | 3.668
| -1.0533 | -1.6655 | -0.11593 | 3.8966
| -0.090693 | -0.20651 | 0.10625 | 0.06216
| 0.12371 | -0.75377 | 0.015361 | 0.58371
| -0.12836 | 0.14245 | 1.1512 | 1.3621
| 0.64149 | -1.1412 | 0.2422 | 1.7724
| -1.1119 | -0.44048 | 1.8061 | 4.6924
| -0.4588 | -0.20424 | -0.35737 | 0.37992
| -0.85279 | 0.44719 | 1.0553 | 2.041
| 0.34544 | -1.8978 | 0.26386 | 3.7904
| -0.90227 | 0.046831 | 1.6025 | 3.3845
| 0.90372 | 0.21859 | -0.39969 | 1.0242
| -0.24969 | -0.51988 | -0.8257 | 1.0144
| -0.12697 | -0.71455 | -0.41541 | 0.69926
| 0.49338 | -0.88711 | -0.7915 | 1.6569
| -0.72179 | 0.63135 | -0.17049 | 0.94864
| -0.57452 | -1.0956 | -1.1667 | 2.8915
| 0.25647 | -0.77878 | -0.87546 | 1.4387
| -0.19005 | 0.078208 | -1.9805 | 3.9645
| 0.34691 | 0.2595 | 0.8129 | 0.84849
| 0.27624 | 0.8734 | -0.2871 | 0.92157
| 1.1923 | -0.31212 | -0.16684 | 1.5469
| -0.063237 | -0.81277 | 0.24205 | 0.72319
| -0.89232 | -1.369 | 1.402 | 4.6359
| -1.6376 | -0.88405 | 0.11957 | 3.4776
| -0.52646 | -0.93127 | -0.43126 | 1.3304
| -0.8462 | -0.64049 | -1.1564 | 2.4635
| -0.43693 | 0.51649 | 1.0631 | 1.5879
| 0.67217 | -0.8317 | 1.2495 | 2.7048
| -1.2132 | -0.35638 | -0.76585 | 2.1853
| -0.2509 | 0.085595 | -2.3668 | 5.6722
| -0.24015 | -0.3725 | -1.8035 | 3.449
| -1.1215 | 1.2048 | -1.6168 | 5.3233
| -0.78877 | -0.72794 | -0.88089 | 1.928
| 0.86545 | -1.5913 | -2.2623 | 8.399
| -0.2113 | -0.36751 | -0.23503 | 0.23496
| 2.8787 | -2.5251 | -0.5943 | 15.0162
| 0.97988 | 0.18774 | -1.1406 | 2.2963
| -0.30721 | 0.18709 | 1.0671 | 1.2681
| -1.1474 | -0.28702 | -0.64925 | 1.8205
| 1.5478 | -1.436 | 0.38172 | 4.6034
| -0.36725 | -0.10239 | 2.7802 | 7.8748
| 0.77521 | -0.13444 | 1.8887 | 4.1861
| -1.1757 | -1.7966 | 2.1544 | 9.2513
| 0.079475 | 0.84706 | -2.1457 | 5.3279
| 2.2773 | -1.2605 | -1.233 | 8.2953
| 0.42621 | -0.26843 | -0.79995 | 0.89362
| 0.5344 | -0.54567 | -0.93569 | 1.4588
| 0.51841 | -0.74295 | 0.15897 | 0.84599
| 0.59174 | -1.4857 | 2.802 | 10.4085
| -0.056743 | 1.647 | 1.5306 | 5.0588
| 0.97666 | 1.4958 | -0.085344 | 3.1985
| 1.049 | 0.22166 | -0.75738 | 1.7232
| 1.7606 | -1.4787 | 0.4172 | 5.4603
| -1.1407 | -0.49596 | -1.515 | 3.8424
| 0.51975 | 0.61242 | -0.25306 | 0.70925
| 0.93791 | -0.85515 | -0.4075 | 1.777
| -0.10636 | -0.25026 | 0.019011 | 0.074306
| -1.2356 | -0.70641 | 0.24378 | 2.0852
| 0.59952 | 0.10069 | -0.79392 | 0.99987
| -0.7235 | -0.2848 | -0.4535 | 0.81022
| -0.38126 | -0.35964 | 1.7399 | 3.302
| -0.469 | -0.054857 | -0.99875 | 1.2205
| -0.40605 | -0.030198 | 2.2024 | 5.0164
| -1.9147 | 1.2945 | 0.011812 | 5.3421
| -1.7724 | -0.48105 | -1.1573 | 4.7122
| 1.6844 | -0.14897 | 0.18337 | 2.8929
| -2.0261 | 1.3702 | 0.40369 | 6.1455
| 0.34527 | -1.2099 | 1.6679 | 4.3651
| -0.42155 | -0.27476 | 1.4255 | 2.2852
| 1.1578 | 0.5153 | -0.97172 | 2.5502
| -0.51545 | 0.33274 | 0.19759 | 0.41545
| 1.4779 | -0.43072 | -0.38149 | 2.5154
| -0.32628 | 0.34431 | -0.28485 | 0.30614
| -0.55009 | -0.057644 | -1.0511 | 1.4107
| -0.54624 | -0.2398 | 0.50637 | 0.61229
| 0.74604 | -0.46221 | 0.12723 | 0.7864
| -0.083496 | -3.5622 | -0.94723 | 13.5932
| -0.70362 | 1.5183 | -0.72877 | 3.3315
| -2.5931 | -0.048088 | -0.225 | 6.777
| 1.6908 | -1.7659 | 1.0976 | 7.1817
| -1.4092 | -1.9842 | -1.5181 | 8.2279
| -0.2242 | -2.084 | -1.5519 | 6.8018
| 0.23083 | -0.37333 | 1.725 | 3.1684
| -1.1984 | 1.2878 | -0.81645 | 3.7612
| 0.20144 | 1.0991 | -1.3316 | 3.0217
| 0.84873 | 0.3079 | -1.7413 | 3.8471
| 0.060327 | 0.25276 | -1.2161 | 1.5464
| 1.1771 | -1.5981 | -0.87378 | 4.7032
| 1.2005 | 1.1185 | -0.83981 | 3.3976
| -0.89399 | 0.78606 | 0.1728 | 1.447
| -1.1195 | -1.3633 | -0.46587 | 3.3289
| 0.55442 | -0.61893 | -0.11963 | 0.70476
| -0.082932 | -0.025631 | -1.8924 | 3.5888
| 1.5343 | 0.23395 | 1.1041 | 3.6281
| 0.10364 | 0.50841 | -1.3444 | 2.0767
| 0.45559 | 0.21239 | -0.39343 | 0.40746
| -0.42059 | 0.83706 | 2.0611 | 5.1255
| 0.41314 | 0.43538 | -0.002795 | 0.36025
| 0.80551 | -0.3483 | 0.1551 | 0.79422
| -0.015923 | 0.091393 | 1.8051 | 3.267
| -0.59211 | 1.7735 | -0.072277 | 3.5012
| -1.0861 | -0.97863 | -1.3651 | 4.0008
| -0.39157 | 1.8488 | 0.91672 | 4.4117
| 0.75874 | -0.1643 | -0.12722 | 0.61887
| 1.2284 | -0.21081 | -0.53791 | 1.8427
| 0.30705 | -0.8524 | -0.84799 | 1.5399
| 0.65634 | 0.41799 | -0.30587 | 0.69906
| -0.29649 | -0.61644 | -0.20911 | 0.51163
| 0.32242 | 2.0891 | -0.45698 | 4.6773
| -0.69978 | 0.66582 | -0.78076 | 1.5426
| 1.9816 | -0.23499 | -0.92445 | 4.8365
| 1.1003 | 2.3894 | -0.6875 | 7.3928
| 1.7255 | -0.25659 | 0.84536 | 3.7578
| 0.24285 | -0.92799 | 0.58632 | 1.2639
| -1.8546 | -1.2007 | -0.36704 | 5.0158
| -0.72901 | -0.45829 | 1.3081 | 2.4525
| -1.0795 | -1.1361 | 0.16911 | 2.4847
| 0.86774 | -0.48414 | 0.35092 | 1.1105
| 0.27951 | -0.38851 | -1.0657 | 1.3648
| 0.31332 | -0.54906 | 0.049127 | 0.40205
| 2.0078 | 0.1029 | -0.15379 | 4.0653
| 1.2277 | 0.23157 | -0.031209 | 1.5618
| -0.49849 | -1.3208 | -0.10087 | 2.0033
| 1.4371 | 0.84011 | -0.69906 | 3.2597
| -0.45617 | -0.15207 | -1.1289 | 1.5056
| -1.1241 | 0.91472 | 1.2266 | 3.6047
| 1.0549 | 0.90797 | -0.90999 | 2.7652
| -0.45838 | -1.261 | -2.6715 | 8.9371
| 0.20048 | -0.86878 | 0.4551 | 1.0021
| 2.4278 | 0.083628 | 0.088162 | 5.9089
| 0.13656 | 1.1063 | 0.29386 | 1.329
| -2.6198 | -1.4547 | 0.14323 | 9.0001
| 0.36268 | 1.836 | -0.71573 | 4.0146
| -0.78773 | -0.81269 | -1.4078 | 3.2629
| 0.32817 | -0.98888 | 1.6411 | 3.7787
| -0.71223 | 0.95179 | -0.02321 | 1.4137
| -1.1554 | -0.24608 | 0.30497 | 1.4885
| 1.0623 | -0.36055 | 1.0974 | 2.4627
| -0.40246 | -0.63029 | -1.1624 | 1.9104
| -0.22817 | -0.76177 | -1.6429 | 3.3316
| -0.50223 | 0.78502 | -1.2333 | 2.3896
| 1.528 | -0.28391 | 0.21755 | 2.4626
| 0.57799 | 0.11605 | 1.7814 | 3.5211
| -0.75445 | -0.31511 | 0.85076 | 1.3923
| -0.68931 | -0.4977 | 0.019939 | 0.72324
| -0.44538 | -0.40238 | -0.44386 | 0.55728
| -0.27251 | 0.51632 | -0.32952 | 0.44943
| -1.4208 | -0.29537 | 0.4086 | 2.2728
| -0.23537 | 0.5104 | -0.3746 | 0.45623
| -0.64161 | 0.18388 | 0.79958 | 1.0848
| 0.3954 | -0.44467 | 0.59594 | 0.70922
| -1.4983 | -0.25141 | 1.0736 | 3.4608
| -0.12085 | 0.048924 | -0.97006 | 0.95802
| -1.2867 | -1.6252 | 0.3023 | 4.3882
| -0.76608 | -0.054866 | 1.2692 | 2.2006
| -1.2514 | 1.2083 | -1.3746 | 4.9155
| -0.55135 | 0.48736 | -0.625 | 0.93213
| -0.075429 | 0.89898 | -0.094839 | 0.82285
| -1.5056 | 1.2532 | -0.51099 | 4.0984
| -0.50207 | -0.4586 | -1.0967 | 1.6651
| 0.91295 | 0.75419 | 0.24892 | 1.4643
| -0.21876 | -0.37781 | -0.94832 | 1.0899
| 0.5647 | 1.3998 | -2.0921 | 6.6553
| -0.07184 | -0.28314 | 1.0377 | 1.1621
| 0.43923 | 0.67596 | -0.35334 | 0.77469
| -0.16302 | 0.12478 | -0.065332 | 0.046415
| -0.41284 | 1.5764 | 0.83608 | 3.3546
| -0.012892 | -0.98742 | 1.2856 | 2.6278
| 0.84749 | -0.15991 | 0.3787 | 0.88723
| -0.30295 | 0.06831 | 0.0557 | 0.099548
| -0.78914 | 0.6218 | 0.32364 | 1.1141
| 0.27733 | 1.3539 | -0.54489 | 2.2069
| -1.1115 | -0.10145 | 0.91417 | 2.0814
| -2.5863 | -0.77497 | 1.074 | 8.443
| -0.68578 | 0.4717 | -1.3875 | 2.618
| -0.24465 | -0.003818 | 2.0015 | 4.0659
| 0.40368 | 0.80325 | -0.69018 | 1.2845
| -0.95047 | -0.71308 | 0.61302 | 1.7877
| 1.0328 | -0.14502 | -0.62141 | 1.4739
| 1.0546 | -0.037075 | 0.57265 | 1.4416
| 0.28556 | -1.0577 | 0.82711 | 1.8843
| 0.86088 | -1.848 | 0.20443 | 4.1981
| -0.44825 | -0.24189 | -1.3672 | 2.1286
| 0.49986 | -1.807 | 0.79125 | 4.1411
| -0.65641 | -0.47211 | 0.43438 | 0.84245
| 0.59639 | 1.2298 | 0.068461 | 1.8727
| -0.091804 | -1.2895 | 1.135 | 2.9595
| -1.1238 | -1.4353 | -0.21225 | 3.368
| -2.2942 | -0.6602 | 1.8256 | 9.0321
| -1.1756 | -1.092 | 0.13733 | 2.5935
| 0.78145 | 0.47394 | 0.48063 | 1.0663
| -0.19692 | 0.23417 | -0.18961 | 0.12956
| -0.12465 | -0.64992 | 1.0953 | 1.6376
| 1.4639 | -2.4615 | 0.88214 | 8.98
| 1.9897 | -1.5559 | -1.8984 | 9.9837
| 0.76884 | 1.7273 | 0.15041 | 3.5972
| 0.32013 | -0.73419 | -0.66899 | 1.0891
| -0.13529 | 0.70115 | 0.20596 | 0.55233
| -0.10271 | 1.0071 | -2.3116 | 6.3684
| -0.59131 | -0.13624 | -1.3411 | 2.1669
| -0.16548 | -0.99802 | -0.31975 | 1.1257
| -0.43491 | 0.553 | 0.77285 | 1.0922
| -1.4101 | -1.2936 | -0.10436 | 3.6726
| -0.61163 | -1.3652 | -1.2289 | 3.748
| -1.4837 | 1.0528 | 1.3007 | 5.0015
| -0.74323 | -0.025904 | -0.36203 | 0.68412
| 0.21096 | -2.9495 | -0.56859 | 9.0675
| -0.40034 | 1.1974 | -0.61559 | 1.9729
| 1.9205 | 0.85952 | 0.044156 | 4.4291
| -0.28359 | -1.404 | 1.4769 | 4.233
| -0.034099 | 0.33014 | -1.1091 | 1.3402
| 0.47808 | 0.062163 | 0.60511 | 0.59859
| -0.91225 | -1.2248 | 1.3869 | 4.2557
| 2.3198 | 0.23417 | -0.83318 | 6.1303
| 1.2091 | -0.75671 | -0.65215 | 2.4598
| -0.78276 | 0.18472 | -0.72234 | 1.1686
| -1.0934 | 0.32784 | 0.80961 | 1.9584
| -0.79783 | -0.51954 | 0.84421 | 1.6191
| -0.19624 | -1.3054 | -0.29411 | 1.8292
| -0.03723 | 0.8947 | -0.036202 | 0.80319
| 1.6945 | 0.25067 | 1.2047 | 4.3856
| 1.0089 | -0.52994 | -1.0876 | 2.4815
| 0.01156 | -0.44309 | -0.58165 | 0.53478
| -0.64434 | -1.3105 | 0.037863 | 2.134
| 0.42113 | -0.86709 | 1.6776 | 3.7434
| 0.4457 | 1.0361 | -0.69296 | 1.7523
| 0.51356 | 1.3693 | 0.63316 | 2.5396
| -0.088927 | 0.73671 | 1.1859 | 1.9571
| -1.7954 | 0.67746 | 1.1934 | 5.1068
| -0.60705 | 0.034817 | 0.51237 | 0.63225
| 0.75823 | 2.3565 | -0.98131 | 7.0909
| -0.45858 | 0.57092 | -0.78712 | 1.1558
| -1.3193 | 0.55881 | 1.9182 | 5.7323
| 1.4287 | 1.0076 | -0.51935 | 3.3261
| 0.23512 | -1.2737 | 0.23709 | 1.7337
| -2.1527 | -0.30281 | -0.29977 | 4.8156
| -0.38104 | -2.1588 | -1.3278 | 6.5687
| -1.5193 | 0.80855 | 1.1298 | 4.2383
| -0.34328 | 0.17996 | -0.17818 | 0.18197
| -2.4087 | -0.55467 | -1.8009 | 9.3528
| -0.97934 | 1.723 | -0.72787 | 4.4577

## Release Notes
Version | Changes | Date
0.1 | Initial Release | 2026-01-06
0.2 | • Added basic units conversion • Implemented Autocorrelation uncertainty calculator • Added Student-T calculation using Algorithm 396 because Excel's built-in T.INV does not handle fractional degrees of freedom. | 2026-05-21

## Formula Reference (hidden)
| These are the same formulas entered in the Name Manager, listed here for readability. To modify, copy them to a real code editor then paste in to the name manager. |  | Comments
General Equations
| __MEASUREMENT_MODEL | =LAMBDA(delta,t,rs,p,SQRT(2*delta*t*rs/p)) | The function the user should update to define a measurement equation. Given here is the Pitot Tube Airspeed Calculation example from Kline and McClintock's paper that originated the GUM equations.
| K_LOOKUP | =LAMBDA(x, IFS( x="Normal k=1", 1, x="Normal k=2", 2, x="Normal k=3", 3, x="Triangular k=2.45", SQRT(6), TRUE, SQRT(3) ) ) | Look up the coverage factor k from distribution dropdown options, which should match the Validation Options in the Lookups sheet.
| HAS_CORRELATION | =LAMBDA(cor, SUM(ABS(cor))>8) | Determine whether there are any non-zero correlation coefficients
| EVAL_MODEL | =LAMBDA(values,function, IFERROR(function(INDEX(values,1),INDEX(values,2),INDEX(values,3),INDEX(values,4),INDEX(values,5),INDEX(values,6),INDEX(values,7),INDEX(values,8)), IFERROR(function(INDEX(values,1),INDEX(values,2),INDEX(values,3),INDEX(values,4),INDEX(values,5),INDEX(values,6),INDEX(values,7)), IFERROR(function(INDEX(values,1),INDEX(values,2),INDEX(values,3),INDEX(values,4),INDEX(values,5),INDEX(values,6)), IFERROR(function(INDEX(values,1),INDEX(values,2),INDEX(values,3),INDEX(values,4),INDEX(values,5)), IFERROR(function(INDEX(values,1),INDEX(values,2),INDEX(values,3),INDEX(values,4)), IFERROR(function(INDEX(values,1),INDEX(values,2),INDEX(values,3)), IFERROR(function(INDEX(values,1),INDEX(values,2)), function(INDEX(values,1))))))))) ) | Evaluates the measurement model function using a range of values, limiting the number of arguments to the number in the array. This is ugly, but there's no introspection or alternative to provide aribtrary-length lambda input. Extend if adding more than 8 variables.
| STUDENTT | =LAMBDA(conf,degf, LET(p, 1-conf, halfpi, PI()/2, IFS(degf=2, SQRT(2/(p*(2-p))-2), degf=1, COS(p*halfpi)/SIN(p*halfpi), TRUE(), LET( a, 1/(degf-0.5), b, 48/a^2, c, ((20700*a/b-98)*a-16)*a + 96.36, d, ((94.5/(b+c)-3)/b+1) * SQRT(a*halfpi) * degf, x, d*p, y, x^(2/degf), ytwo, IF(y>0.05+a, LET( xa, NORM.S.INV(p*0.5), ya, xa^2, ca, IF(degf<5, c+0.3*(degf-4.5)*(xa+0.6), c), cb, (((0.05*d*xa-5)*xa-7)*xa-2)*xa+b+ca, yb, (((((0.4*ya+6.3)*ya+36)*ya+94.5)/cb-ya-3)/b+1)*xa, yc, a*yb^2, IF(yc>0.1, EXP(yc)-1, ((yc+4)*yc+12)*yc*yc/24 + yc) ), ((1/(((degf+6)/(degf*y) - 0.089*d - 0.822) * (degf+2)*3)+0.5/(degf+4))*y-1) * (degf+1)/(degf+2) + 1/y ), SQRT(degf*ytwo) ) ) ) ) | Calculate Student T coverage factor using "Algorithm 396" by G.W. Hill. Excel's built-in T.INV.2T doesn't handle fractional degrees of freedom.
| STUDENTT_CONF | =LAMBDA(degf,kfactor, BISECT(LAMBDA(conf, STUDENTT(conf, degf)-kfactor),0.00001,0.99999) ) | Find level of confidence for student t distribution
| STUDENTT_DEGF | =LAMBDA(conf,kfactor, BISECT(LAMBDA(df, STUDENTT(conf, df)-kfactor),1,1000) ) | Find degrees of freedom for student t distribution
GUM Equation Functions
| EPSILON | =LAMBDA(uncs, i, MAP(SEQUENCE(8), LAMBDA(k, IF(k=i, INDEX(uncs,k)/1000, 0)) ) ) | Calculate the delta used in numerical differentiation, based on magnitude of uncertainty
| DFDX | =LAMBDA(values,uncs,function, MAP(SEQUENCE(8), LAMBDA(k, LET( eps, EPSILON(uncs, k), plus, EVAL_MODEL(values+eps, function), minus, EVAL_MODEL(values-eps, function), dfdx, (plus-minus)/(INDEX(eps,k)*2), IF(INDEX(eps,k)=0, 0, IFERROR(dfdx, "")) ) ) ) ) | Calculate sensitivity coefficients of the measurement model - GUM Equation 11b.
| CORRELATION_SUM | =LAMBDA(uncerts,dfdx,coeffs, 2*SUM( MAP(SEQUENCE(8), LAMBDA(i, SUM(MAP(SEQUENCE(8), LAMBDA(k, IF(i<k, INDEX(uncerts,i)*INDEX(uncerts,k)*INDEX(dfdx,i)*INDEX(dfdx,k)*INDEX(coeffs,i,k), 0 ) ) )) ) ) ) ) | Calculates the contribution to the combined variance due to correlation terms - GUM Equation 15.
| UNCERTAINTY_VARIANCE | =LAMBDA(uncerts,dfdx,[correlations], LET(x, uncerts*dfdx, combined, SUMSQ(FILTER(x, NOT(ISERROR(x)))), IF(ISOMITTED(correlations), combined, combined+CORRELATION_SUM(uncerts, dfdx, correlations) ) ) ) | Calculate variance of the measurement model using GUM approach - GUM Equation 13
| DEGF_EFFECTIVE | =LAMBDA(ucombined,uncerts,dfdx,degfs, LET( s_over_d, uncerts^4*dfdx^4/degfs, ucombined^4 / SUM(FILTER(s_over_d, NOT(ISERROR(s_over_d)))) ) ) | Effective degrees of freedom using Welch-Satterthwaite approximation - GUM Equation G.2b.
Monte Carlo Functions
| RAND_NORM | =LAMBDA(mu, a, NORM.INV(RAND(), mu, a)) | Generate a random number from normal distribution with mean mu and standard devaition a.
| RAND_UNIFORM | =LAMBDA(mu, a, (RAND()-0.5)*a*2 + mu) | Generate a random number from uniform distribution with mean mu and half-width a.
| RAND_TRIANG | =LAMBDA(mu, a, LET(b, mu-a, c, mu+a, b + (c-b)/2*(RAND()+RAND()) ) ) | Generate a random number from triangular distribution with mean mu and half-width a. See JCGM 101:2008, 6.4.5.
| RAND_SAMPLES | =LAMBDA(shape,mu,a,nsamples, MAP(SEQUENCE(nsamples), LAMBDA(k, IF(shape="Uniform", RAND_UNIFORM(mu, a), IF(shape="Triangular", RAND_TRIANG(mu, a), RAND_NORM(mu, a)) ))) ) | Generate random samples using the given distribution shape
| INVCDF_UNIFORM | =LAMBDA(mu, a, p, p*(2*a)+mu-a) | Inverse CDF of uniform distribution
| INVCDF_TRIANG | =LAMBDA(mu, alpha, p, LET(a, mu-alpha, b, mu+alpha, IF(p<0.5, a+SQRT(2)*SQRT(p)*(b-a)/2, b+SQRT(2-2*p)*(a-b)/2 ) ) | Inverse CDF of triangular distribution
| MC_SAMPLES | =LAMBDA(function,nominals,uncerts,distributions,corr,nsamples, IF(HAS_CORRELATION(corr), MC_SAMPLE_CORRELATED(function,nominals,uncerts,distributions,corr,nsamples), MC_SAMPLE_UNCOR(function,nominals,uncerts,distributions,nsamples) ) ) | Generate Monte Carlo samples of the output value. This function just checks for correlations then calls the appropriate sub-function below.
| MC_SAMPLE_UNCOR | =LAMBDA(function,nominals,uncerts,distributions,nsamples, LET(pos_uncerts, IF(uncerts<=0, 1E-99, uncerts), MAP(SEQUENCE(nsamples), LAMBDA(i, EVAL_MODEL( MAP(SEQUENCE(8), LAMBDA(k, LET(shape, LEFT(INDEX(distributions,k),1), mu, INDEX(nominals, k), alpha, INDEX(pos_uncerts, k), IF(shape="U", RAND_UNIFORM(mu, alpha*SQRT(3)), IF(shape="T", RAND_TRIANG(mu, alpha*SQRT(6)), RAND_NORM(mu, alpha)) ) ) )), function ) ) ) ) ) | Generate uncorrelated random Monte Carlo samples and evaluate measurement model. Samples may be normal, uniform, or triangular. Returns a sequence of length nsamples.
| MC_SAMPLE_CORRELATED | =LAMBDA(function,nominals,uncerts,distributions,corr,nsamples, LET(nvars,ROWS(uncerts), pos_unc,IF(uncerts<=0,1E-20,uncerts), shape, IFS(LEFT(distributions,1)="U", 2, LEFT(distributions,1)="T", 3, TRUE, 1), rtranspose,V_CHOLESKY(corr), z,NORM.S.INV(RANDARRAY(nvars,nsamples)), normcdf, NORM.S.DIST(MMULT(rtranspose,z), TRUE), uniformsamples, INVCDF_UNIFORM(nominals, pos_unc*SQRT(3), normcdf), trisamples, INVCDF_TRIANG(nominals, pos_unc*SQRT(6), normcdf), normsamples, NORM.INV(normcdf, nominals, pos_unc), samples, CHOOSE(shape, normsamples, uniformsamples, trisamples), BYCOL(samples, LAMBDA(i, EVAL_MODEL(i, function))))) | Generate correlated Monte Carlo samples and evaluate measurement model. Samples may be normal, uniform, or triangular. Returns a sequence of length nsamples. See JCGM 101:2008 Appendix C.5.
| MONTE_CARLO | =LAMBDA(function,values,uncerts,distributions,corr,nsamples,confidence, LET(samples, MC_SAMPLES(function, values, uncerts, distributions, corr, nsamples), hbins, HISTBINS(samples, 40), bincounts, HIST(samples, hbins), VSTACK( AVERAGE(samples), STDEV.S(samples), MIN(samples), QUARTILE.INC(samples, 1), MEDIAN(samples), QUARTILE.INC(samples, 3), MAX(samples), PERCENTILE.INC(samples, (1-confidence)/2), PERCENTILE.INC(samples, 1-(1-confidence)/2), hbins, bincounts ) ) ) | Calculate Monte Carlo samples and evaluate statistics, including max, min, quartiles, and 95% coverage region. Also generates histogram bins and density for plotting.
| HISTBINS | =LAMBDA(samples,nbins, LET( xmin, MIN(samples), xmax, MAX(samples), width, (xmax-xmin)/nbins, MAP(SEQUENCE(nbins+1), LAMBDA(x, xmin+width*(x-1))) ) ) | Divide the range into suitable histogram bins
| HIST | =LAMBDA(samples, bins, LET( nsamples, COUNT(samples), bwidth, INDEX(bins, 2)-INDEX(bins, 1), MAP(SEQUENCE(COUNT(bins)-1), LAMBDA(i, COUNT( LET(gg, FILTER(samples, samples>=INDEX(bins, i)), FILTER(gg, gg<INDEX(bins, i+1)) ) )/nsamples/bwidth ) ) ) ) | Generate histogram densities for the samples with the bins calculated using HISTBINS. NOTE: COUTIFS doesn't work here - Excel will throw Nested Array errors - so we fall back on this COUNT and FILTER implementation.
Autocorrelation
| AUTOCORRELATION | =LAMBDA(values, [kfactor], LET(kfactor, IF(ISOMITTED(kfactor), 2, kfactor), n, COUNTA(values), mean, AVERAGE(values), denom, SUM((INDEX(values, SEQUENCE(n))-mean)^2), rho, MAP(SEQUENCE(n,1,0), LAMBDA(k, LET(a, INDEX(values, SEQUENCE(n-k)), b, INDEX(values, SEQUENCE(n-k,1,k+1,1)), SUM((b-mean) * (a-mean))/denom ) )), sigrho, VSTACK( SQRT(1/n), MAP(SEQUENCE(n-1), LAMBDA(k, SQRT((1 + 2*SUM(INDEX(rho,SEQUENCE(k,1,2))^2))/n) ))), HSTACK(rho, kfactor*sigrho) ) ) | Calculate autocorrelation vs. lag array
| AUTOCORR_UNCERT | =LAMBDA(values, rho, LET(n, COUNTA(values), sigrho, INDEX(rho, SEQUENCE(n), 2), rho, INDEX(rho, SEQUENCE(n), 1), nc, XMATCH(TRUE(), ABS(rho)>sigrho,0,-1), r, 1 + 2/n*SUM((n-SEQUENCE(nc))*INDEX(rho,SEQUENCE(nc,1,2))), variance, VAR.S(values)/n*r, VSTACK( SQRT(variance), variance, r, SQRT(r), nc ) ) ) | Calculate uncertainty due to autocorrelation
Units Conversion
| ISUNIT | =LAMBDA(unit, OR(unit="", IF( ISNUMBER(MATCH(unit,BaseUnitsTable[Unit],0)), TRUE, ISNUMBER(MATCH(unit,CustomUnitsTable[Unit],0))) ) ) | TRUE if input defined in the units tables
| BASEUNIT | =LAMBDA(unit, BYROW(unit, LAMBDA(u, IF(u="", "", XLOOKUP(u, BaseUnitsTable[Unit], BaseUnitsTable[Base], XLOOKUP(u, CustomUnitsTable[Unit], CustomUnitsTable[Base]))) )) ) | Get name of base units
| FROM_BASEUNIT | =LAMBDA(values,units, BYROW(HSTACK(values, units), LAMBDA(v, LET(val, INDEX(v,1), unit, INDEX(v,2), baseunit, XLOOKUP(unit, BaseUnitsTable[Unit],BaseUnitsTable[Base]), IF(unit="", val, IFERROR(CONVERT(val, baseunit, unit), val/XLOOKUP(unit,CustomUnitsTable[Unit],CustomUnitsTable[Factor])) ) ) )) ) | Convert the value from base units to the provided unit
| TO_BASEUNIT | =LAMBDA(values,units,[diff], LET(diff, IF(ISOMITTED(diff), FALSE, diff), BYROW(HSTACK(values, units), LAMBDA(v, LET(val, INDEX(v,1), unitp, INDEX(v,2), unit, IF(diff, IF(unitp="F", "deltaF", IF(unitp="C", "deltaC", unitp)), unitp), IF(unit="", val, IFERROR(CONVERT(val, unit, XLOOKUP(unit,BaseUnitsTable[Unit],BaseUnitsTable[Base])), val*XLOOKUP(unit,CustomUnitsTable[Unit],CustomUnitsTable[Factor]) )) ) )) ) ) | Convert the value to SI base units
Utility Functions
| BISECT | =LAMBDA(f,a,b,[tol],[iters], LET(tol, IF(ISOMITTED(tol), 0.000000001, tol), iters, IF(ISOMITTED(iters), 100, iters), fa, f(a), fb, f(b), IF(SIGN(fa)=SIGN(fb), 0, LET(mid, (a+b)/2, fmid, f(mid), IF(OR(iters<=0, ABS(fa-fmid) <= tol), mid, IF(fa*fmid < 0, BISECT(f, a, mid, tol, iters-1), BISECT(f, mid, b, tol, iters-1)) ) ) ) ) ) | Find zeros of the function f between a and b using bisection
| V_CHOLESKY | =LAMBDA(matrix, LET(doc,"https://www.vertex42.com/lambda/cholesky.html", n,ROWS(matrix), IF(ROWS(matrix)<>COLUMNS(matrix), "Error: matrix not square", IF(SUM(--(matrix<>TRANSPOSE(matrix)))<>0, "Error: not symmetric", IF(MDETERM(matrix)<=0, "Error: not positive definite", REDUCE(0,SEQUENCE(n),LAMBDA(Lmat,j, IF(j=1, LET(Ljj,SQRT(INDEX(matrix,j,j)), Lij,CHOOSEROWS(INDEX(matrix,0,1),SEQUENCE(n-1,1,2,1))/Ljj, VSTACK(Ljj,Lij)), IF(j=n, LET(Ljj,SQRT(INDEX(matrix,j,j)-SUM(CHOOSEROWS(Lmat,j)^2)), HSTACK(Lmat,VSTACK(SEQUENCE(n-1,1,0,0),Ljj))), LET(Ljj,SQRT(INDEX(matrix,j,j)-SUM(CHOOSEROWS(Lmat,j)^2)), Aij,CHOOSEROWS(INDEX(matrix,0,j),SEQUENCE(n-j,1,j+1,1)), Lik,CHOOSEROWS(Lmat,SEQUENCE(n-j,1,j+1,1)), Ljk,CHOOSEROWS(Lmat,j), Lij,1/Ljj*(Aij-MMULT(Lik,TRANSPOSE(Ljk))), HSTACK(Lmat,VSTACK(SEQUENCE(j-1,1,0,0),Ljj,Lij))))))) ))))) | Cholseky decomposition (Source: https://www.vertex42.com/lambda/cholesky.html, MIT Licensed)

## Lookups (hidden)
Distribution Dropdown:
Normal k=1
Normal k=2
Normal k=3
Uniform k=1.73
Triangular k=2.45