## Instructions
Sandia Primary Standards Lab - Smart Measurement Assurance Plan
v0.2
uncertainty@sandia.gov
The Measurement Assurance Plan answers the question "Is my measuring equipment good enough to inspect a UUT against a tolerance?" It computes Test Uncertainty Ratio and Global Probablity of False Accept as metrics of measurement decision risk. It also computes Guardbanded Acceptance Limits, using the Root-Difference-Square (RDS) or Target-PFA methods as a means to mitigate this risk. Other tools are included for monitoring a measurement process over time. No Excel Macros are used, however the calculations make use of Excel's LAMBDA function which is only available in Office 365 or 2024.
Spreadsheet Tabs
MAP-TUR | This tab provides a Measurement Assurance Plan (MAP) that calculates Test Uncertainty Ratio and guardbands for each testpoint.
MAP-PFA | Use this tab to calculate Global (Average) Probability of False Accept (PFA) and guardbands that target a desired PFA.
Equipment | Table of equipment specifications that can be looked up from MAP tabs.
Type A - Repeatability | Calculate uncertainty in N repeated measurements under repeatability conditions.
Type A - Reproducibility | Calculate uncertainty in N repeated conditions x M reproducibility groups (such as days, operators, etc.)
GUM Uncertainty | Calculate uncertainty in an indirect measurement model using the GUM method.
Intervals | Calculate optimal calibration intervals to achieve desired reliability using method A3 from NCSLI RP-1.
SPC | Statistical Process Control for control-chart monitoring of a measurement process over time
Prof.Test | Evaluate results of proficiency testing for monitoring a measurement process
Risk Calculator | Calculates global and specific risk for a single measurement, assuming normal distributions. Includes cost model to optimize tradeoff between false accept and false reject.
Criticality Levels | This tab is used to define Criticality Levels that determine the guardbanding policy, TUR thresholds, and PFA targets.
Formula Reference | Lists the formulas defined in Excel's Name Manager for reference. Hidden by default.
Lookup | Provides allowable options for dropdown menus. Hidden by default.
MAP Input Columns
Quantity Measured | The property of interest being measured on the device under test. Some examples include voltage, current, mass, humidity, length, neutron flux, pressure, pulse time, shock, strain, temperature, torque, vibration, etc.
Requirement Number | Provides traceability to product requirements being tested
Criticality Level | Used to determine guardbanding policy for the item in terms of a minimum TUR or target PFA. See the Criticality Levels tab for definitions.
Expected Value | The nominal or expected value being measured, as a numeric value (do not include units or non-numeric text here).
Tolerance | Enter the product requirement tolerance being tested. Tolerance has 3 spreadsheet columns. The center column provides a dropdown to select the tolerance type. Use "±" or "to" for two-sided tolerances, and ">" or "<" for one-sided tolerances.
Units | Enter measurement units here for recordkeeping. Units conversion is not currently done, but there is a check to ensure the Equipment units match the Tolerance units.
In-Tolerance Probability | Enter the (true) in-tolerance probability as a percent. This is the percent of all similar devices expected to be in tolerance, based on historical calibration/measurement data or other information.
Equipment | Enter the make and model of the equipment used to make the measurement.
Equipment Range | For equipment lookup from the Equipment tab, enter the equipment range.
Type A Uncertainty | Enter any Type A uncertainty associated with this measurement. This may be linked to a Type A worksheet if needed. An expanded uncertainty at a 95% level of confidence (approximately k=2) should be entered.
Other Uncertainty | Enter any other uncertainty associated with the measurement. An expanded uncertainty at a 95% level of confidence (approximately k=2) should be entered.
Equipment Units | Enter the units of measure for the equipment. No units conversion is currently done, but the equipment units will be checked against the Tolerance units.
Guardband Method | (MAP-PFA Tab): Select between Root-Difference-Square guardbanding and PFA guardbanding. RDS guardbanding calculates the guardband factor as (SQRT(1 - 1/TUR^2)). PFA guardbanding iteratively solves for the guardband that acheives the target PFA defined by the Criticality Level table
Guardband Override | Check this box to override the built-in guardband calculation.
Calibrate? | Indicate whether calibration is required. Optionally link to an "Intervals" tab to monitor calibration interval.
SPC? | Indicate whether Statistical Process Control is used to monitor the measurement. Optionally link to an SPC tab.
Proficiency Test? | Indicate whether proficiency testing is performed on the measurement. Optionally link to a Prof.Test tab.
Comments | Enter any additional notes about the measurement, such as limitations, reasoning for not calibrating, etc.
MAP Calculated Columns
Equipment Accuracy Lookup | Compute equipment accuracy using data from Equipment tab.
Equipment Combined Uncertainty | Combines (RSS) uncertainty components of the measurement
TUR | Test Unceratinty Ratio. For two-sided limits, this is the tolerance divided by the equipment uncertainty. For single-sided limits, an "effective TUR" is calculated as ¦Nominal - Tolerance¦ / Uncertainty.
PFA (No Guardband) | The Global Probability of False Accept with no guardbanding applied.
Acceptance Limits | Limits at which to accept the product, with any guardbanding applied.
PFA | The Global PFA with guardbanding applied
PFR | The Global Probability of False Reject (PFR) with guardbanding applied.
About
Sandia National Laboratories is a multimission laboratory managed and operated by National Technology & Engineering Solutions of Sandia, LLC, a wholly owned subsidiary of Honeywell International Inc., for the U.S. Department of Energy’s National Nuclear Security Administration under contract DE-NA0003525.
SAND2025-10736O

## MAP-TUR
Product #: | Examples | [30 empty] | ← Intermediate guardband calculations in hidden columns here.
Organization:
Preparer Name:
Date:
Measurand and Tolerance Under Test | [7 empty] | Measurement Equipment and Uncertainty (k=2) | [8 empty] | False Accept Risk and Guardband |  |  |  | Measurement Assurance |  |  | Comments
Quantity Measured | Requirement Number | Criticality Level | Expected Value | Tolerance Limits |  |  | Units | Equipment | Equipment Range | ± | Accuracy Lookup | Accuracy Reliability | Type A Uncertainty | Other Uncertainty | Combined Uncertainty | Units | TUR | Acceptance Limits |  |  | Calibrate? | SPC? | Proficiency Test? |  |  | Lower Tolerance | Upper Tolerance | Guardband? | Guardband Width | Lower Acceptance | Upper Acceptance
Name | Requirement | Criticality | Expected | Tolerance | TolType | ToleranceB | Units | Equipment | Range | ± | Accuracy | MTEReliability | TypeA | OtherUnc | Uncertainty | MeasUnits | TUR | Accept1 | AcceptType | Accept2 | Accept3 | SPC | ProfTest | Comments | …. | TL | TU | Guardband | GBW | AL | AU
Voltage | N/A | 1 | 60 | 60 | ± | 0.05 | mV | Fluke 5522A (DCV) | 330mV | ± | 0.0038 | 95.45% | 0 | 0 | 0.0038 | mV | 13.1579 | 60 | ± | 0.05 | Y |  |  | Equipment accuracy read from Equipment tab |  | 59.95 | 60.05 | FALSE | 0 | 59.95 | 60.05
Pool Length | FR 2.1.2.1 | 1 | 50 | 49.995 | to | 50.005 | m | Total Station |  | ± | 0 | 95.45% | 0 | 0.001 | 0.001 | m | 5 | 49.995 | to | 50.005 | Y |  |  | Equipment not found in Equipment tab. Enter manually in "Other Uncertainty" column. |  | 49.995 | 50.005 | FALSE | 0 | 49.995 | 50.005
Ping pong ball weight |  | 1 | 2.7 | 2.67 | to | 2.77 | g | Balance |  | ± | 0 | 95.45% | 0 | 0.01 | 0.01 | g | 5 | 2.67 | to | 2.77 | Y |  |  | Asymmetric tolerance example |  | 2.67 | 2.77 | FALSE | 0 | 2.67 | 2.77
Football Pressure | Section 1, Rule 2 | 1 | 13 | 13 | ± | 0.5 | psig | Wilson (?) Pressure Gage |  | ± | 0 | 95.45% | 0 | 0.8 | 0.8 | psig | 0.625 | 13 | ± | 0 | Y |  |  | Example from ENGR224. Notice guardbanded tolerance is ±0. |  | 12.5 | 13.5 | TRUE | 0.5 | 13 | 13
One-sided Length |  | 1 | 10 |  | max | 12 | cm | Yardstick |  | ± | 0 | 95.45% | 0 | 0.5 | 0.5 | cm | 4 |  | max | 12 | Y |  |  | Example maximum limit |  | #N/A | 12 | FALSE | 0 | #N/A | 12
One-sided Length |  | 1 | 13 |  | min | 12 | cm | Yardstick |  | ± | 0 | 95.45% | 0 | 0.5 | 0.5 | cm | 2 |  | min | 12.134 | Y |  |  | Example minimum limit |  | 12 | #N/A | TRUE | 0.13397 | 12.134 | #N/A
⬑Equipment Accuracy is extracted from Equipment tab using name in "Equipment Used" and "Range" column, or "Quantity" column to chain calibrations
⬑ Use 95.45% reliability unless there is data suggesting a different value
⬑ Type A may be linked to a Type A Uncertainty tab
⬑ Other Uncertainty may be linked to a GUM Uncertainty tab for indirect measurement models.
⬑ Combined uncertainty is the RSS of all the uncertainty components.
⬑Link the Measurement Assurance columns to interval, SPC, or Prof.Test tabs as appropriate.

## MAP-PFA
Product #: | Examples | [37 empty] | ← Intermediate guardband calculations in hidden columns here.
Organization:
Preparer Name:
Date:
Measurand and Tolerance | [8 empty] | Measurement Equipment and Accuracy | [8 empty] | False Accept Risk and Guardband | [8 empty] | Measurement Assurance |  |  | Comments
Quantity Measured | Requirement Number | Criticality Level | Expected Value | Tolerance |  |  | Units | In Tolerance Probability | Equipment | Equipment Range |  | Equipment Accuracy Lookup | Equipment Reliability | Type A Uncertainty (k=2) | Other Uncertainty (k=2) | Combined Uncertainty (k=2) | Units | TUR | PFA (No Guardband) | Guardband Method | Acceptance Limits |  |  | Guardband Override? | PFA (Guardbanded) | PFR (Guardbanded) | Calibrate? | SPC? | Proficiency Test? |  |  | Lower Tolerance | Upper Tolerance | Process Stdev | Guardband Needed? | Guardband Width | Lower Acceptance Limit | Upper Acceptance Limit
Quantity | Requirement | Criticality | Expected | TolA | TolType | TolB | Units | ITP | Equipment | Range | ± | Accuracy | MTEReliability | TypeA | OtherUnc | Uncertainty | MeasUnits | TUR | PFA | Gbmethod | AcceptA | AcceptType | AcceptB | GBOverride | PFAgb | PFR | Calibrate | SPC | ProfTest | Comments | … | TL | TU | stdev | Guardband | GBW | AL | AU
Voltage Calibrator | N/A | 1 | 60.01 | 60 | ± | 0.1 | mV | 95.45% | Fluke 5522A (DCV) | 330mV | ± | 0.0038003 | 95.45% |  | 0.05 | 0.050144 | mV | 1.9942 | 1.243% | PFA | 60 | ± | 0.1 |  | 1.243% | 4.097% | Y |  |  | Equipment accuracy and reliability is read from the Equipment tab. |  | 59.9 | 60.1 | 0.048997 | FALSE | 0 | 59.9 | 60.1
⬑Battery Voltage | N/A | 1 | 60 | 60 | ± | 0.25 | mV | 80% | Voltage Calibrator |  | ± | 0.097994 | 95.45% |  | 0 | 0.097994 | mV | 2.5512 | 2.87% | PFA | 60 | ± | 0.23608 |  | 2% | 6.06% | Y |  |  | Example of chaining measurement traceability. The accuracy/reliability of the Voltage Calibrator is pulled from the previous row. |  | 59.75 | 60.25 | 0.19508 | TRUE | 0.013925 | 59.7639 | 60.2361
Pool Length | FR 2.1.2.1 | 1 | 50 | 49.995 | to | 50.005 | m | 95.45% | Total Station |  | ± | 0 | 95.45% |  | 0.001 | 0.001 | m | 5 | 0.6784% | PFA | 49.995 | to | 50.005 |  | 0.6784% | 1.118% | N |  |  | Example of a "to" limit. |  | 49.995 | 50.005 | 0.0025 | FALSE | 0 | 49.995 | 50.005
Ping pong ball weight |  | 1 | 2.7 | 2.67 | to | 2.77 | g | 95.45% | Balance |  | ± | 0 | 95.45% |  | 0.01 | 0.01 | g | 5 | 0.807% | PFA | 2.67 | to | 2.77 |  | 0.807% | 1.453% | Y |  |  | Tolerances don't have to be centered about the expected value. |  | 2.67 | 2.77 | 0.017746 | FALSE | 0 | 2.67 | 2.77
Football Pressure | Section 1, Rule 2 | 1 | 13 | 13 | ± | 0.5 | psig | 80% | Wilson |  | ± | 0 | 95.45% |  | 0.8 | 0.8 | psig | 0.625 | 6.689% | RDS | 13 | ± | 0 |  | 0% | 80% | N |  |  | Example from ENGR224 class. Notice acceptance limits are ±0. No false accepts when nothing is accepted! |  | 12.5 | 13.5 | 0.39015 | TRUE | 0.5 | 13 | 13
One-sided Length |  | 1 | 10 |  | max | 12 | cm | 95.45% | Yardstick |  | ± | 0 | 95.45% |  | 0.5 | 0.5 | cm | 4 | 0.6485% | PFA |  | max | 12 |  | 0.6485% | 1.011% | Y |  |  | One-sided maximum limit |  | #N/A | 12 | 1.1833 | FALSE | 0 | #N/A | 12
One-sided Length |  | 1 | 13 |  | min | 12 | cm | 95.45% | Yardstick |  | ± | 0 | 95.45% |  | 0.5 | 0.5 | cm | 2 | 1.059% | PFA |  | min | 12 |  | 1.059% | 2.494% | Y |  |  | One-sided minimum limit |  | 12 | #N/A | 0.59166 | FALSE | 0 | 12 | #N/A
[12 empty] | ⬑Equipment Accuracy is extracted from Equipment tab using name in "Equipment Used" and "Range" column, or "Quantity" column to chain calibrations | [14 empty] | ⬑Link the Measurement Assurance columns to interval, SPC, or Prof.Test tabs as appropriate.
⬑ Use 95.45% reliability unless there is data suggesting a different value
⬑ Type A may be linked to a Type A Uncertainty tab
⬑ Other Uncertainty may be linked to a GUM Uncertainty tab for indirect measurement models.
⬑ Combined uncertainty is the RSS of all the uncertainty components.

## Equipment
[6 empty] | Uniform | Normal k=2 | Normal k=2 | 95.45% for k=2
Equipment/Function | Range Name | Range Max | %Reading | %Range | Constant | Resolution | Other1 | Other2 | Reliability | Combined Constant
Keysight 3458A (DCV) | 100uV | 100 | 0.0009% | 0.0003% | 0 | 1e-08 | 0 | 0 | 95.45% | 0.0003
Keysight 3458A (DCV) | 1V | 1 | 0.0008% | 3e-05% | 0 | 1e-08 | 0 | 0 | 95.45% | 3.0006e-07
Keysight 3458A (DCV) | 10V | 10 | 0.0008% | 5e-06% | 0 | 1e-07 | 0 | 0 | 95.45% | 5.0332e-07
Keysight 3458A (DCV) | 100V | 100 | 0.001% | 3e-05% | 0 | 1e-06 | 0 | 0 | 95.45% | 3.0006e-05
Keysight 3458A (DCV) | 1000V | 1000 | 0.001% | 1e-05% | 0 | 1e-05 | 0 | 0 | 95.45% | 0.00010017
0
Fluke 5522A (DCV) | 330mV | 330 | 0.003% |  | 0.002 |  |  |  | 95.45% | 0.002
Fluke 5522A (DCV) | 3.3V | 3.3 | 0.002% |  | 2e-06 |  |  |  | 95.45% | 2e-06
Fluke 5522A (DCV) | 33V | 33 | 0.002% |  | 2e-05 |  |  |  | 95.45% | 2e-05
Fluke 5522A (DCV) | 330V | 330 | 0.003% |  | 0.00015 |  |  |  | 95.45% | 0.00015
Fluke 5522A (DCV) | 1000V | 1000 | 0.003% |  | 0.0015 |  |  |  | 95.45% | 0.0015
0
0

## Type A - Repeatability
9.9727 | ← Repeatability. Enter Repeat Measurements in Column B
10.2405
10.2258 |  | Uncertainty in determining this mean value |  | ← Use this box, for example, when measuring an attribute of a product to see if it meets specifications
10.0813 |  | Standard Deviation | 0.16821
10.2493 |  | Standard Error of the Mean | 0.037612
10.0303
9.7967 |  | Mean | 10.0517
10.1907 |  | Standard Uncertainty | 0.037612
10.1888 |  | Degrees Freedom | 19
10.1413 |  | Expanded (k=2) Uncertainty | 0.078724
9.7663
10.1827
10.2104 |  | Applied to uncertainty of NEW measurements |  | ← Use this box, for example, when calibrating a measurement instrument to be used to take future measurements
9.8051 |  | Number of new repeat measurements | 1
9.9631 |  | Standard Uncertainty | 0.16821
10.0689 |  | Degrees Freedom | 19
10.0694 |  | Expanded (k=2) Uncertainty | 0.35206
9.8415
9.8147
10.1946

## Type A - Reproducibility
Reproducibility. Enter R&R measurements in table below.
Run ↴ \ Group → | A | B | C | D | E | F | G | H | I | J | Reproducibility |  | Deg. F.
1 | 0.95101 | 1.0089 | 1.0132 | 0.96616 | 1.0071 | 0.96762 | 0.95905 | 1.0167 | 1.0331 | 0.98005 | Grand Mean | 0.9974
2 | 0.95833 | 0.95468 | 0.99785 | 0.98744 | 1.0267 | 1.0039 | 1.0059 | 1.018 | 1.04 | 0.9586 | Repeatability Std. Dev. | 0.030765 | 90
3 | 0.95919 | 0.96017 | 0.99752 | 0.95252 | 1.0392 | 0.97402 | 1.0059 | 1.0222 | 1.0452 | 0.96819 | Reproducibility Std. Dev. | 0.010637 | 9
4 | 1.0496 | 1.0379 | 1.0229 | 0.98583 | 1.018 | 0.95183 | 0.97295 | 0.95036 | 1.0358 | 0.98732
5 | 1.0319 | 0.9628 | 0.99968 | 0.98254 | 1.0365 | 1.0035 | 1.0223 | 0.9701 | 1.0364 | 0.96474 | Uncertainty in this mean value |  |  | ← Use this box, for example, when measuring an attribute of a product to see if it meets specifications
6 | 1.0327 | 0.96672 | 1.0192 | 0.99218 | 0.98317 | 0.97926 | 1.008 | 1.0382 | 0.98304 | 1.0278 | Standard Uncertainty | 0.0031037
7 | 1.0048 | 1.0477 | 0.99156 | 0.99653 | 1.0337 | 1.0236 | 1.0197 | 0.95418 | 1.0414 | 0.97323 | Degrees Freedom | 99
8 | 1.0475 | 0.95285 | 0.97403 | 0.96524 | 0.95035 | 1.0291 | 1.0167 | 0.97867 | 0.9922 | 1.0353 | Expanded (k=2) Uncertainty | 0.0061584
9 | 0.95565 | 1.0293 | 0.98272 | 1.0131 | 0.96504 | 0.96971 | 0.96446 | 0.96961 | 0.96958 | 1.0094
10 | 0.99312 | 0.95638 | 1.0153 | 0.95917 | 0.95145 | 1.0214 | 1.0486 | 1.0413 | 1.0279 | 1.037
[11 empty] | Applied to uncertainty of NEW measurements |  |  | ← Use this box, for example, when calibrating a measurement instrument to be used to take future measurements
[11 empty] | Number of new repeat measurements | 1
[11 empty] | Standard Uncertainty | 0.030765
[11 empty] | Degrees Freedom | 99
Group Averages Vj: | 0.99838 | 0.98774 | 1.0014 | 0.98007 | 1.0011 | 0.9924 | 1.0024 | 0.99593 | 1.0205 | 0.99416 | Expanded (k=2) Uncertainty | 0.061044
Group Variances s^2(Vjk): | 0.001624 | 0.001489 | 0.00025959 | 0.00035653 | 0.0012679 | 0.00074507 | 0.00080825 | 0.0012143 | 0.00076907 | 0.00093093
Group Std. Devs s(Vjk): | 0.040299 | 0.038588 | 0.016112 | 0.018882 | 0.035607 | 0.027296 | 0.02843 | 0.034847 | 0.027732 | 0.030511
# Measurements nj: | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
Deg. Freedom | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9
# Groups J: | A: 10
Total Measurements: | A: 100
|  | ANOVA | Value | Deg. Freedom
|  | F-test Confidence | 95%
|  | Variance of Means s^2(Vj) | 0.00011315 | 9
| Between Group Variance | sa^2 | 0.0011315 | 9
| Repeatability Variance | sb^2 | 0.00094646 | 90
|  | F | 1.1955
|  | F critical | 1.9856
|  | Between-Group Effect? | FALSE
|  | Uncertainty if no between-group effect | 0.0031037 | 99
|  | Uncertainty if between-group effect | 0.0033637 | 9

## GUM Uncertainty
Press CTRL+F3 to open Name Manager, then edit the __GUM_MODEL function to define the measurement equation.
To duplicate this tab, create another copy of __GUM_MODEL and change cells L6 and Q6 (dashed outline) to reference the new function.
Measured Variables and Uncertainties | [5 empty] | Results |  | Value | Units | Variable | Standard Unc. | Sensitivity | Proportion
Variable | Value | Units | Uncertainty | k = | Deg. Freedom | Measurand Value |  | 57.488 | m/s | Δ | 0.001806 | 99.4602 | 27.27%
Δ | 0.289 | psi | 0.003612 | 2 | 1000000 | Combined Standard Uncertainty |  | 0.34398 | m/s | T | 0.056 | 0.098169 | 0.02554%
T | 292.8 | K | 0.112 | 2 | 1000000 |  | Variance | 0.11832 |  | Rs | 0 | 0 | 0%
Rs | 287.06 | J/K/kg | 0 | 2 | 1000000 | Effective Degrees of Freedom |  | 1658450.3939 |  | p | 0.15 | -1.9554 | 72.71%
p | 14.7 | psi | 0.3 | 2 | 1000000 | Level of Confidence |  | 95% |  | - | 0 | 0 | 0%
- | [5 empty] | Coverage Factor (k) |  | 1.96 |  | - | 0 | 0 | 0%
- | [5 empty] | Expanded Uncertainty |  | 0.23191 | m/s | - | 0 | 0 | 0%
- | [5 empty] | Uncertainty Range, Low |  | 57.2561 | m/s | - | 0 | 0 | 0%
- | [5 empty] | Uncertainty Range, High |  | 57.7199 | m/s
Tip: Use k=1.73 for uniform distributions
Examples - Copy/Paste into variables table and change the __GUM_MODEL to the given LAMBDA:
Pitot Tube Airspeed Example |  |  |  |  | Example from Kline and McClintock's 1953 paper originating the GUM method
=LAMBDA(delta,t,rs,p,SQRT(2*delta*t*rs/p))
Variable | Value | Units | Uncertainty | k = | Deg. Freedom
Δ | 0.289 | psi | 0.003612 | 2 | 1000000 | Measured difference in static vs total pressure
T | 292.8 | K | 0.112 | 2 | 1000000 | Ambient Temperature
Rs | 287.06 | J/K/kg | 0 | 2 | 1000000 | Specific gas constant for dry air. Assuming no uncertainty.
p | 14.7 | psi | 0.3 | 2 | 1000000 | Ambient pressure
-
-
-
-
Viscosity of test liquid |  |  |  |  | Example E3 from NIST TN-1900
=LAMBDA(muc,rhob,rhoc,rhom,tc,tm, muc*tm*(rhob-rhom)/(tc*(rhob-rhoc)))
Variable | Value | Units | Uncertainty | k = | Deg. Freedom
mu | 4.63 | mPa.s | 0.0463 | 1 | 1000000 | Viscosity of calibration liquid
rho_b | 2217 | kg/m3 | 0.5 | 1 | 1000000 | Density of ball
rho_c | 810 | kg/m3 | 0.5 | 1 | 1000000 | Density of calibration liquid
rho_m | 1180 | kg/m3 | 0.5 | 1 | 1000000 | Densityof sodium hydroxide
t_c | 36.6 | s | 5.49 | 1 | 1000000 | Travel time in calibration liquid
t_m | 61 | s | 6.1 | 1 | 1000000 | Travel time in sodium hydroxide

## Intervals
May include calibration cycles for similar items (other items of same model number, for example)
Enter "1" for in-tolerance, "0" for out of tolerance. Can use "NA", or other designators to exclude items that were damaged, lost, etc.
Asset | Workorder | Start Date | End Date | Interval Length | IT(1)/OOT(0) | Parameters
|  | 4017900% | 4054400% | 365 | 1 | Interval (days) to Test | 365
|  | 4054400% | 4090900% | 365 | 1 | Tolerance (days) | 30
|  | 4090900% | 4127500% | 366 | 1 | Reliability Target | 95%
|  | 4127500% | 4164300% | 368 | 1 | Change confidence | 50%
|  | 4164000% | 4237000% | 730 | 1 | Maximum change factor | 2
|  | 4237000% | 4310100% | 731 | 1
|  |  |  | 0 | 1
|  |  |  | 0 | 1 | Total calibrations | 4
|  |  |  | 0 | 1 | Passing calibrations | 4
|  |  |  | 0 | 1 | Observed Reliability | 100%
|  |  |  | 0 | 1
|  |  |  | 0 | 1 | True reliability (lower) | 84.09%
|  |  |  | 0 | 1 | True reliability (upper) | 100%
|  |  |  | 0 | 1 | Rejection Confidence | 18.55%
|  |  |  | 0 | 1
|  |  |  | 0 | 1 | Suggested interval | 365
|  |  |  | 0 | 1

## SPC
Date/Time | Measured Value | Average | Lower Limit | Upper Limit | Exceed Limit?
4565800% | 10.2956 | 10.0188 | 9.3637 | 10.674 |  | Item Description
4565900% | 9.932 | 10.0188 | 9.3637 | 10.674 |  | Control Limit σ | 3
4566000% | 10.213 | 10.0188 | 9.3637 | 10.674
4566100% | 9.9751 | 10.0188 | 9.3637 | 10.674
4566200% | 10.1312 | 10.0188 | 9.3637 | 10.674
4566300% | 10.4809 | 10.0188 | 9.3637 | 10.674
4566400% | 10.2027 | 10.0188 | 9.3637 | 10.674
4566500% | 9.8118 | 10.0188 | 9.3637 | 10.674
4566600% | 9.97 | 10.0188 | 9.3637 | 10.674 |  | Control Limit | 0.65515
4566700% | 9.8931 | 10.0188 | 9.3637 | 10.674
4566800% | 10.3663 | 10.0188 | 9.3637 | 10.674
4566900% | 9.4772 | 10.0188 | 9.3637 | 10.674
4567000% | 9.9425 | 10.0188 | 9.3637 | 10.674
4567100% | 10.0361 | 10.0188 | 9.3637 | 10.674
4567200% | 9.7603 | 10.0188 | 9.3637 | 10.674
4567300% | 10.1887 | 10.0188 | 9.3637 | 10.674
4567400% | 10.0131 | 10.0188 | 9.3637 | 10.674
4567500% | 9.9493 | 10.0188 | 9.3637 | 10.674
4567600% | 10.0199 | 10.0188 | 9.3637 | 10.674
4567700% | 9.7699 | 10.0188 | 9.3637 | 10.674
4567800% | 10.0861 | 10.0188 | 9.3637 | 10.674
4567900% | 9.919 | 10.0188 | 9.3637 | 10.674
4568000% | 9.9183 | 10.0188 | 9.3637 | 10.674
4568100% | 10.2972 | 10.0188 | 9.3637 | 10.674
4568200% | 10.0676 | 10.0188 | 9.3637 | 10.674
4568300% | 9.7731 | 10.0188 | 9.3637 | 10.674

## Prof.Test
Proficiency Testing Worksheet
Measurement System:
Test Date | Participating Lab ID | Nominal Value | Units | Measured Value | Measured Uncertainty (k=2) | External Lab's Measured Value | External Lab's Measured Uncertainty (k=2) | Normalized Error En | Pass/Fail | Notes
4599400% | PSL | 10 | g | 9.9986 | 0.0015 | 10.0002 | 0.001 | -0.88752 | 0.88752
4599400% | PSL | 20 | g | 19.9997 | 0.0015 | 20.0001 | 0.001 | -0.20524 | 0.20524

## Risk Calculator
Global (Average) Risk Calculator | [5 empty] | Lower Tolerance | 54 | ← Intermediate calculations in hidden columns here.
Nominal/Expected Value | 60 |  |  |  |  | Upper Tolerance | 66
Tolerance | 60 | ± | 6 |  |  | Process Stdev | 4.168
In-Tolerance Probability | 85% |  |  |  |  | Guardband Width | 0.34315
Measurement Uncertainty (k=2) ± | 2 |  |  |  |  | Lower Acceptance Limit | 54.3431
Guardband Method | RDS |  |  |  |  | Upper Acceptance Limit | 65.6569
Criticality Level | 1
TUR | 3
PFA (No Guardband) | 2.187%
Acceptance Limit | 60 | ± | 5.6569
PFA(Guardbanded) | 1.406%
Specific Risk Calculator
Measured Value | 62.6 |  |  |  |  | Lower Tolerance | 54
Tolerance | 60 | ± | 6 |  |  | Upper Tolerance | 66
Measurement Uncertainty (k=2) ± | 2
Specific Risk, Upper Limit | 0.03369%
Specific Risk, Lower Limit | 3.986e-16%
Specific Risk | 0.03369%
Probability of Conformance | 99.97%
Cost Model - compare financial tradeoff between false accept and false reject
Global (Average) Risk Calculator |  |  |  | Cost Model |  |  |  | ← Intermediate calculations in hidden columns here.
PFA Required | 0.1% |  |  | Cost of False Accept | 200
Nominal/Expected Value | 0 |  |  | Cost of False Reject | 30 | Lower Tolerance | -1
Tolerance | 0 | ± | 1 | Number of units produced | 1000000 | Upper Tolerance | 1
In-Tolerance Probability | 90% |  |  |  |  | Process Stdev | 0.60796
Measurement Uncertainty (k=2) ± | 0.04
[6 empty] | Guardband Width | 0.01334
Risk Results |  |  |  | Cost Results |  | Lower Acceptance Limit | -0.98666
TUR | 25 |  |  | Number of Falsely Accepted Items | 999.9888 | Upper Acceptance Limit | 0.98666
PFA (No Guardband) | 0.2622% |  |  | Number of Falsely Rejected Items | 5797.4885
PFR (No Guardband) | 0.2807% |  |  | Cost of Falsely Accepted Items | 199997.7619
Acceptance Limit | 0 | ± | 0.98666 | Cost of Falsely Rejected Items | 173924.6559
PFA (Guardbanded) | 0.1% |  |  | Total cost of false decisions | 373922.4178
PFR (Guardbanded) | 0.5797%

## Criticality Levels
Level | MaxPfa | MinTur | Define thresholds at which to apply guardbanding for different criticality levels.
0 | 5% | 1
1 | 2% | 4
2 | 1.5% | 6
3 | 1% | 8
4 | 0.5% | 10
5 | 0.25% | 12

## Formula Reference (hidden)
These are the same formulas entered in the Name Manager, listed here for readability and additional comments. To modify, copy them to a real code editor then paste in to the name manager. |  | Comments
General Equations
BIVARIATE_NORMCDF | =LAMBDA(p,q,rho, LET(cx, -1.0950081470333, cy, -0.75651138383854, d, SQRT(1-rho^2), a, -rho/d, b, p/d, aqplusb, a*q + b, z, 1-a^2*cy, sqrtz, SQRT(z), sqrt2, SQRT(2), IF(a>=0, NA(), IF(aqplusb>=0, 0.5 + 0.5 * ERF(q/sqrt2) - 1/(4*sqrtz) * EXP((a^2*cx^2+2*sqrt2*b*cx+2*b^2*cy)/(4*z)) * (1 + ERF((sqrt2*q - sqrt2*a^2*cy*q - sqrt2*a*b*cy - a*cx)/(2*sqrtz))), 0.5 - 0.5 * ERF(b/(sqrt2*a)) - 1/(4*sqrtz) * EXP((a^2*cx^2+2*sqrt2*b*cx+2*b^2*cy)/(4*z)) * (1 - ERF((sqrt2*b + a^2*cx)/(2*a*sqrtz))) + 1/(4*sqrtz) * EXP((a^2*cx^2-2*sqrt2*b*cx+2*b^2*cy)/(4*z)) * (ERF((sqrt2*q-sqrt2*a^2*cy*q - sqrt2*a*b*cy + a*cx)/(2*sqrtz)) + ERF((-(a^2)*cx + sqrt2*b)/(2*a*sqrtz))) ) ) ) ) | Compute integral of standard bivariate normal distribution with correlation rho from (-infty..p), (-infty..q). Uses the algorithm given by Tsay and Ke, "A simple approximation for the bivariate normal integral", 2021. https://doi.org/10.1080/03610918.2021.1884718 In the PFA and PFR calculations, rho is always positive, so only cases 4 and 5 of the Tsay-Ke algorithm are needed. Other cases (a>0) will return #NA, but could be added easily. This approximation is much faster than using numerical (Trapezoidal or Simpson) integration methods, and results in similar or better accuracy. The distributions used in the standard PFA equation must be scaled appropriately before input to the standard bivariate normal distribution. PFA_SYMMETRIC and PFA_ASYMMETRIC functions apply this scaling.
EQUIPMENT_LOOKUP | =LAMBDA(name,rangename,reading, LET(pctreading, XLOOKUP(1, (EquipmentSpecs[Equipment/Function]=name)*(EquipmentSpecs[Range Name]=rangename), EquipmentSpecs[%Reading], 0 ), const, XLOOKUP(1, (EquipmentSpecs[Equipment/Function]=name)*(EquipmentSpecs[Range Name]=rangename), EquipmentSpecs[Combined Constant], 0), pctreading*reading + const ) ) | Look up Equipment Accuracy from the EquipmentSpecs table. Adds %reading to %range and other uncertainties in the table.
Tolerance Functions |  | For manipulating tolerance defined by three cells.
IS_SYMMETRIC | =LAMBDA(tl,tu,al,au,nom, IFS(ISERROR(tu-nom), FALSE, ISERROR(nom-tl), FALSE, ISERROR(au-nom), FALSE, ISERROR(al-nom), FALSE, ABS((tu-nom)-(nom-tl))>1E-9, FALSE, ABS((al-nom)-(nom-al))>1E-9, FALSE, TRUE, TRUE ) ) | Determine whether the tolerance is symmetric about the nominal.
TOL_LOWER_LIMIT | =LAMBDA(tola,toltype,tolb, IFS(toltype="±", tola-tolb, toltype="to", tola, toltype="min", MAX(tolb, tola))) | Returns the lower limit of a tolerance cell set, or NA for maximum tolerances
TOL_UPPER_LIMIT | =LAMBDA(tola,toltype,tolb, IFS(toltype="±", tola+tolb, toltype="to", tolb, toltype="max", MAX(tolb, tola))) | Returns the upper limit of a tolerance cell set, or NA for minimum tolerances
TOL_DISPLAY_1 | =LAMBDA(tl,tu,toltype, IFS(toltype="±", (tl+tu)/2, toltype="to", tl, 1, " " ) ) | Determine the first value (before the ± indicator) to display in the acceptance limit column.
TOL_DISPLAY_2 | =LAMBDA(tl,tu,toltype, LET(center, (tl+tu)/2, pm, ABS(tl-center), IFS(toltype="±", pm, toltype="to", tu, toltype="max", tu, toltype="min", tl ) ) ) | Determine the second value (after the ± indicator) to display in the acceptance limit column.
SIGMA_FROM_ITP | =LAMBDA(nom,tl,tu,itp, IFS(AND(ISNUMBER(tl), ISNUMBER(tu)), LET( center, (tl+tu)/2, width, (tu-tl)/2, bias, (nom-center)/width, IF(ABS(bias)<0.000000001, width/NORM.S.INV((1+itp)/2), LET(func, LAMBDA(sigma, NORM.DIST(1, bias, sigma, TRUE)-NORM.DIST(-1,bias,sigma,TRUE)), width*L_SOLVE_NR(func, 0.5, itp, 0.00001) ) ) ), ISNUMBER(tl), IF(tl=nom, NA(), ABS(nom-tl)/NORM.S.INV(itp)), ISNUMBER(tu), IF(tu=nom, NA(), ABS(tu-nom)/NORM.S.INV(itp)) ) ) | Convert limits and ITP into standard deviation
PFA Functions
TUR | =LAMBDA(tola,toltype,tolb,unc,nom, LET(t, MAX(tola,tolb), IFS(toltype="±", tolb/unc, toltype="to", ABS(tolb-tola)/2/unc, toltype="max", (t-nom)/unc, toltype="min", (nom-t)/unc) ) ) | Calculate test uncertainty ratio as tolerance/uncertainty. For one-sided limits, the "modified TUR" [abs(limit-tolerance)/uncertainty] is calculated if the nominal is not equal the tolerance. `unc` is k=2 uncertainty.
PFA | =LAMBDA(tl,tu,al,au,nom,sigt,sigp, IF(IS_SYMMETRIC(tl,tu,al,au,nom), PFA_SYMMETRIC((tu-tl)/2, (au-al)/2, sigt, sigp), PFA_ASYMMETRIC(tl, tu, al, au, nom, sigt, sigp) ) ) | Calculate global/average probability of false accept given tolerance, guardband, measurement uncertainty, and product standard deviation. Assumes normal distributions, but works for one-sided or two-sided tolerances, and nominal value does not have to be centered between the limits.
PFA_SYMMETRIC | =LAMBDA(t,a,sigt,sigp, LET(sp, sigp, st, SQRT(sigt^2 + sigp^2), rho, SQRT(sigp^2 / (sigp^2+sigt^2)), 2*(BIVARIATE_NORMCDF(a/st, -t/sp, rho) - BIVARIATE_NORMCDF(-a/st, -t/sp, rho)) ) ) | Compute global PFA for symmetric tolerances (two-sided with nominal centered between the tolerances). sigmat and sigmap defined by the test measurement and product distribution must be converted to sp, st, and rho as used by the bivariate normal distribuiton. Then the tolerances are scaled by these sigmas for use in the standard bivariate integral.
PFA_ASYMMETRIC | =LAMBDA(tl,tu,al,au,nom,sigt,sigp, LET(sp, sigp, st, SQRT(sigt^2 + sigp^2), rho, SQRT(sigp^2 / (sigp^2+sigt^2)), ttl, IF(ISNUMBER(tl), tl, -1E+99), ttu, IF(ISNUMBER(tu), tu, 1E+99), aal, IF(ISNUMBER(al), al, ttl), aau, IF(ISNUMBER(au), au, ttu), pfaleft, IF(ISNUMBER(ttl), BIVARIATE_NORMCDF((aau-nom)/st, (ttl-nom)/sp, rho) - BIVARIATE_NORMCDF((aal-nom)/st, (ttl-nom)/sp, rho), 0), pfarght, IF(ISNUMBER(ttu), BIVARIATE_NORMCDF((-aal+nom)/st, (-ttu+nom)/sp, rho) - BIVARIATE_NORMCDF((-aau+nom)/st, (-ttu+nom)/sp, rho), 0), pfaleft+pfarght ) ) | Compute global PFA for asymmetric tolerances (nominal not centered, or one-sided limits). sigmat and sigmap defined by the test measurement and product distribution must be converted to sp, st, and rho as used by the bivariate normal distribuiton. Then the tolerances are scaled by these sigmas for use in the standard bivariate integral.
PFR | =LAMBDA(tl,tu,al,au,nom,sigt,sigp, IF(IS_SYMMETRIC(tl,tu,al,au,nom), PFR_SYMMETRIC((tu-tl)/2, (au-al)/2, sigt, sigp), PFR_ASYMMETRIC(tl, tu, al, au, nom, sigt, sigp) ) ) | Compute global probability of false reject tolerance, guardband, measurement uncertainty, and product standard deviation. Assumes normal distributions, but works for one-sided or two-sided tolerances, and nominal value does not have to be centered between the limits.
PFR_SYMMETRIC | =LAMBDA(t,a,sigt,sigp, LET(sp, sigp, st, SQRT(sigt^2 + sigp^2), rho, SQRT(sigp^2 / (sigp^2+sigt^2)), 2*(BIVARIATE_NORMCDF(-a/st, t/sp, rho) - BIVARIATE_NORMCDF(-a/st, -t/sp, rho)) ) ) | Compute global PFR for symmetric tolerances (two-sided with nominal centered between the tolerances). sigmat and sigmap defined by the test measurement and product distribution must be converted to sp, st, and rho as used by the bivariate normal distribuiton. Then the tolerances are scaled by these sigmas for use in the standard bivariate integral.
PFR_ASYMMETRIC | =LAMBDA(tl,tu,al,au,nom,sigt,sigp, LET(sp, sigp, st, SQRT(sigt^2 + sigp^2), rho, SQRT(sigp^2 / (sigp^2+sigt^2)), ttl, IF(ISNUMBER(tl), tl, -1E+99), ttu, IF(ISNUMBER(tu), tu, 1E+99), aal, IF(ISNUMBER(al), al, ttl), aau, IF(ISNUMBER(au), au, ttu), pfrbot, BIVARIATE_NORMCDF((aal-nom)/st, (ttu-nom)/sp, rho) - BIVARIATE_NORMCDF((aal-nom)/st, (ttl-nom)/sp, rho), pfrtop, BIVARIATE_NORMCDF((-aau+nom)/st, (-ttl+nom)/sp, rho) - BIVARIATE_NORMCDF((-aau+nom)/st, (-ttu+nom)/sp, rho), pfrtop+pfrbot ) ) | Compute global PFR for asymmetric tolerances (nominal not centered, or one-sided limits). sigmat and sigmap defined by the test measurement and product distribution must be converted to sp, st, and rho as used by the bivariate normal distribuiton. Then the tolerances are scaled by these sigmas for use in the standard bivariate integral.
Guardbanding Functions |  | RDS and PFA-Target guardband methods are supported.
GB_NEEDED | =LAMBDA(gbmethod,turval,pfaval,criticality, IFS(gbmethod="RDS", turval<GB_MINTUR(criticality), gbmethod="PFA", pfaval>GB_MAXPFA(criticality), TRUE, FALSE ) ) | Determine whether guardbanding is needed per the criticality level and criticality table.
GB_MAXPFA | =LAMBDA(criticality, INDEX(CriticalityTable[MAXPFA], MATCH(criticality,CriticalityTable[Criticality]), 0)) | Get maximum allowable PFA from criticality table
GB_MINTUR | =LAMBDA(criticality, INDEX(CriticalityTable[MINTUR], MATCH(criticality,CriticalityTable[Criticality]), 0)) | Get minimum allowable TUR from criticality table
GB_WIDTH | =LAMBDA(tur,unc,tl,tu, IF(ISNUMBER(tur), IF(tur>=1, unc*(tur*(1-(SQRT(1-1/tur^2)))), (tu-tl)/2 ), unc ) ) | Calculate guardbanded width, based on the guardbanding method (either RDS or PFA) and criticality level.
GBW_RDS | =LAMBDA(tur,unc,tl,tu, IF(AND(ISNUMBER(tur), tur>=1), unc*(tur*(1-(SQRT(1-1/tur^2)))), unc ) ) | Calculate guardbanded width using Root-Difference-Squares method. If TUR is not defined (one-sided limit), the width of the guardband is 1*uncertainty. `unc` is k=2.
GBW_TARGET | =LAMBDA(tl,tu,nom,sigt,sigp,pfaval, IFS(AND(ISNUMBER(tl), ISNUMBER(tu)), LET( func, LAMBDA(x, PFA(tl,tu,tl+x,tu-x,nom,sigt,sigp)), L_SOLVE_NR(func, sigt/10, pfaval, sigt/100, sigt/5000, 25) ), ISNUMBER(tl), LET( func, LAMBDA(x, PFA_ASYMMETRIC(NA(),tu,NA(),tu-x,nom,sigt,sigp)), L_SOLVE_NR(func, sigt/10, pfaval, sigt/100, sigt/5000, 25) ), ISNUMBER(tu), LET( func, LAMBDA(x, PFA_ASYMMETRIC(tl,NA(),tl+x,NA(),nom,sigt,sigp)), L_SOLVE_NR(func, sigt/10, pfaval, sigt/100, sigt/5000, 25) ) ) ) | Calculate guardband width to target the desired PFA. Results in symmetric guardband widths. Solved numerically.
GUM Uncertainty Functions
__GUM_MODEL | =LAMBDA(delta,t,rs,p,SQRT(2*delta*t*rs/p)) | User-defined measurement equation. Defaults to the Pitot-Tube example from Kline-McClintock paper. Edit this in Name Manager for different measurement equations.
GUM_EVAL | =LAMBDA(values,function, IFERROR(function(INDEX(values,1),INDEX(values,2),INDEX(values,3),INDEX(values,4),INDEX(values,5),INDEX(values,6),INDEX(values,7),INDEX(values,8)), IFERROR(function(INDEX(values,1),INDEX(values,2),INDEX(values,3),INDEX(values,4),INDEX(values,5),INDEX(values,6),INDEX(values,7)), IFERROR(function(INDEX(values,1),INDEX(values,2),INDEX(values,3),INDEX(values,4),INDEX(values,5),INDEX(values,6)), IFERROR(function(INDEX(values,1),INDEX(values,2),INDEX(values,3),INDEX(values,4),INDEX(values,5)), IFERROR(function(INDEX(values,1),INDEX(values,2),INDEX(values,3),INDEX(values,4)), IFERROR(function(INDEX(values,1),INDEX(values,2),INDEX(values,3)), IFERROR(function(INDEX(values,1),INDEX(values,2)), function(INDEX(values,1))))))))) ) | Evaluates the measurement model function using a range of values, limiting the number of arguments to the number in the array. This is ugly, but there's no introspection or alternative to provide aribtrary-length lambda input. Extend if adding more than 8 variables.
GUM_EPSILON | =LAMBDA(uncs,i, MAP(SEQUENCE(8), LAMBDA(k, IF(k=i, INDEX(uncs,k)/1000, 0)) ) ) | Compute epsilon used to compute numerical derivatives in GUM sensitivity coefficients
GUM_DFDX | =LAMBDA(values,uncs,function, MAP(SEQUENCE(8), LAMBDA(k, LET(eps, GUM_EPSILON(uncs, k), plus, GUM_EVAL(values+eps, function), minus, GUM_EVAL(values-eps, function), dfdx, (plus-minus)/(INDEX(eps,k)*2), IF(INDEX(eps,k)=0, 0, IFERROR(dfdx, "")) ) )) ) | Calculate GUM sensitivity coefficients numerically as [f(x+h)-f(x-h)]/2h, where h is computed using GUM_EPSILON.
GUM_VARIANCE | =LAMBDA(uncerts, dfdx, LET(x, uncerts*dfdx, SUMSQ(FILTER(x, NOT(ISERROR(x)))) ) ) | Compute variance in measurement model using GUM equation, given uncertainties and sensitivity coefficients.
GUM_DEGF | =LAMBDA(varcombined,uncerts,dfdx,degfs, LET(s_over_d, uncerts^4*dfdx^4/degfs, varcombined^2 / SUM(FILTER(s_over_d, NOT(ISERROR(s_over_d)))) ) ) | Calculate effective degrees of freedom using Welch-Satterthwaite
Third-Party Functions
L_SOLVE_NR | =LAMBDA(function,xstart,[y],[xstep],[epsilon],[iterations],[info], LET(doc,"https://www.vertex42.com/lambda/newton-raphson.html", version,"1.0.0", n,ROWS(xstart), y,IF(ISOMITTED(y),0,y), xstep,IF(ISOMITTED(xstep),0.1,xstep), epsilon,IF(ISOMITTED(epsilon),1E-14,epsilon), iterations,IF(ISOMITTED(iterations),20,iterations), ret,REDUCE({"x","Fx","Iterations"},SEQUENCE(n),LAMBDA(acc_row,i, VSTACK(acc_row, REDUCE( HSTACK(INDEX(xstart,i),function(INDEX(xstart,i))-y,0), SEQUENCE(iterations), LAMBDA(acc,k, LET(cur_row,TAKE(acc,-1), IF(ABS(INDEX(cur_row,1,2))<ABS(epsilon),acc, LET( xn,INDEX(cur_row,1,1), fn,INDEX(cur_row,1,2), fprimen,(function(xn+xstep)-y-fn)/xstep, xnew,xn-fn/fprimen, fnew,function(xnew)-y, new_row,HSTACK(xnew,fnew,k), output,IF(info="verbose",VSTACK(acc,new_row),new_row), output ) ) )) )) )), IF(OR(info=TRUE,info="verbose"),ret,DROP(ret,1,-2)) )); | Newton-Raphson solver From https://www.vertex42.com/lambda/newton-raphson.html (MIT Licensed)

## Release Notes
Version | Changes | Date
0.1 | Initial Release | 2025-07-10
0.2 | • Convert MAPs to use Excel Table so the hidden columns are automatically extended when adding new rows. • Fix conditional formatting error due to floating-point comparison • Change PFA calculation algorithm to increase performance • Improve GUM Uncertainty calculation to use LAMBDA functions • Added Equipment Lookup table • Added A3 Interval calculation worksheet • Added Proficiency Testing worksheet • Added Statistical Process Control worksheet | 2025-12-17

## Lookup
Tolerance Types | Guardband Methods
± | RDS
to | PFA
max
min